1. [bookmark: _GoBack]Right click on empty space at Lotus Note Home page and select “Open Application” 
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2. Then, select “Zephyr” from the “Look in” drop down list. 
[image: ]
3. Select “Form Database” ,then click “open”.
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the ones you wish to keep.

Available fles

1 Venture-GES QDa DA Tracker 2018
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1 D7075 05.Feb.18 | Simon Hanley W1441001408-R Basedon |Due to ongoing yield issues experienced with the 186008511 - KBD5001 Open  |provide Serial numbers of Instruments affected.
& 5:5:::?;32‘ d‘;::‘;g“::‘f i 13 Units build | standard 410004501 Watlow manufactured |Affected Pumping blcok Assy's may ship as part of a
15041 Morntay, 7oy 2018 [thermacouple, items from an alternate supplier - Provide all affected SN |QDA instrument once tested to doc 46202: TSP - QDA
) ITC Thermocouples are being evaluated. These Manufacturing Test. Affected items should not be
(% Falure reporting 1ovay 5. S T e [Smontaniey | wisootosR Ssag o items have been-tested for insulation resistance 553085 T KBDE0S e [shipped as spares material (Due to isible difference
Version reated st tme the 20 Units build _|and vacuum leak integrity, and have been shown in colours of thermocouple wires from current
10:51Thursday, 10 May, 2018 lto meet the specifications shown on the current Provide all aflected SN revision
— |, \version of drawing 441001408, Deviation raised to
(R SF Upcte - 30 Apr.ésx [Orc
Version cested st s the 3 07163 16.Mar.18 | Simon Hanley |GES DFR (PN 71600625) _ Based o |Release of WPPL / Venture specific DHR (716005625 86008511 - Gpen _[Allow the use of WPPL / Venture specific DR and
06:00 Monday, 1 January, 1601 25 Units build  is required prior to build of pilot batch of 25 KBDS001 to KBD5025 lassociated documentations ahead of their release. This
instruments |deviation only applies to 25 off pilot batch instruments
manufactured by Venture. No requirement to record
s |deviation in the DHR.
2 07158 15.Mar.18 | Simon Forrester 400005478 Based on _|Suppler T Electronics feedback that PN 400005478 86008511 - Gpen _[No further actions on Venture other than cansuming the
Wi Sing Tan | (Feedthrough Support | 30 Units build  [have been producing using local standard material GB | Pending GES RF Coil investigation 130 off parts for this deiation
(RF Coil Box)) 152312012 T2 since year 2013 when project started First 20 - TT China
with Waters Last 5 - Venture
Provide all affected SN
5
5 o771t 27.Mar.18 | Simon Forrester 400005478 Based on___|Suppier Shenzen Baiteng in China proposed local 86008511 - Gpen _|Venture to provide assembly Serial Numbers
Wi Sing Tan | (Feedthrough Support | 200 Units build ~|standard material GB 5231-2012 / T11050 T2 material
(RF Coil Box)) 0.8 THK Copper Provide all affected SN
7
5 7167 03.Apr.18 | Wan Jing Loh 615003626 Basedon | Include 616003826 (ETL Label) and 615005062 86008511 - Gpen _|Change PN 615003826 i Doc 48906 page 3 o correct
(ETL Label) 25 Units build |(Rating Plate) to the BOM KBDS001 to KBD5025 the label released
s
7 7176 06.Apr.18 | Wan Jing Loh 615003649 Basedon | The specs call for Suppher PIN to be "L30" 86008511 - Gpen _[Allow Venture to use Shock Dot Label with supplier PN
(Shock Dot Label) 25 Units build KBDS001 to KBD5025 /SD-100G and instead of a Shock Dot Placement Label
allow use of a Shock Watch Pllacement Label
]
5 o781 T3.Apr.18 | Wan Jing Loh 289011127 Based on units _[The performance of the Intemal Cal_system can b 86008511 - Gpen _[Prowde Serial Numbers of fon Biock Assy and top level
(Calibration Pin Assy) ~|used with modified improved by creating a fough / sharp edge on the Cal instrument assy which are fitted with modified pins
CalPins  [Pin. This allows a method of reworking 289001127
(Cal Pin Assy. Use P600 paper or neede file should be To rework the parts a minimum of 2 times. IFfails on the
visible when viewed an eye piece and must not 12nd attempt it should be failure tagged tested in a 2nd
significantly afer the length of the pin. instrument before s rejected or scrap
10
5 07185 T6.Apr.18 | Wan Jing Loh | Acquity GDa ANITS V4.1 | Four Test Bays in | The four T8 uses ThinkStation P50 PC to run all 86008511 - KBD500T Gpen _|Waters action required are the following.
i software GES |AMITS tests. The increased speed has an impact on
lonly 1 test AMITS sequence "Intemal Calibrant 1. Amend the command value within al installations of
Benchmark”. lAamTs
_ 2. Log which installations of AMITS had been deviated
© Which file do I want to save? By adding 1 second delay between commands, 3. Update Doc 46215 to change any future installations
Gose |AMITS siowed down and allows the unit time to Jof AMITS on P510 PC.
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