Carryover specs improved from 0.01% to 0.002%.

Setup of Alliance 6th Final Tester

- increase the capacity by 20%.

- IQ in progress, tester target to be live in end Q2, 2018.

- Alliance packaging height reduce to 80cm proposal.

Advice as ex-field service engineer

FTN Dynamic Leak due to loose fitting improvements

- Study of wear rate of fitting tubings used for interconnect between UUT and testers was carried out on and it was found that average of 10% of the Dynamic Leak failures was due to leak/loose fitting.

- It was evaluated that the current change of tubing during PM is insufficient and a monthly change of tubing for the tester   was implemented.

- The Dynamic Leak failures due to leak/loose fitting drop to an average of 0.64% during a 6 month period.

Pump Seal Change Frequency  reduce downtime ?%

- Bifrost MR was conducted and completed for the following models:

  - 186015085 ACQUITY UPLC H-Class PLUS SM FTN-H

  - 186015088 ACQUITY UPLC I-Class PLUS SM FTN-I

  - 186015080 ACQUITY UPLC H-Class PLUS QSM

  - 186015082 ACQUITY UPLC I-Class PLUS BSM

As member of Venture Bifrost MR audit team point out the hiport tester safety requirement.

IQ

- Seamless execution on IQ activities for Bifrost preparation

before MR.

  - IQ activities for a total 8 out of 22 test systems was

performed and completed across 2 sites.

  - Tight schedule and planning for IQ so as not

affecting production shipment for April month end.

  - 3 full time headcount including OT together with

Milford and Contract Manufacturer test operators

was involved in completing the IQ activities.

IQ – Detector firmware downloader at Flex in July2018

4 FTN testers, 4 pump testers

1 PDA/eLambda linearity tester IQ

Reducing downtime in Venture

FTN  reduced testers unscheduled downtime 50% by performing weekly BSM pump priming and daily tester rebooting

BSM/QSM dynamic leak failure rate redueced 3% by implement the periodically changing of test tubing

Field Feedback investigation for 	GPV and FTN seal

Complete tester health review