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Setup of Alliance 6th Final Tester 
 
- increase the capacity by 20%.
 - IQ in progress, tester target to be live in end Q2, 2018.



· Alliance packaging height reduce to 80cm proposal.
Advice as ex-field service engineer

FTN Dynamic Leak due to loose fitting improvements
· Study of wear rate of fitting tubings used for interconnect between UUT and testers was carried out on and it was found that average of 10% of the Dynamic Leak failures was due to leak/loose fitting.
· It was evaluated that the current change of tubing during PM is insufficient and a monthly change of tubing for the tester   was implemented. 
·  The Dynamic Leak failures due to leak/loose fitting drop to an average of 0.64% during a 6 month period.
Pump Seal Change Frequency  reduce downtime ?%


· Bifrost MR was conducted and completed for the following models: 
· 186015085 ACQUITY UPLC H-Class PLUS SM FTN-H
·  186015088 ACQUITY UPLC I-Class PLUS SM FTN-I
·  186015080 ACQUITY UPLC H-Class PLUS QSM
·  186015082 ACQUITY UPLC I-Class PLUS BSM


As member of Venture Bifrost MR audit team point out the hiport tester safety requirement. 
IQ
· Seamless execution on IQ activities for Bifrost preparation  
   before MR.
·  IQ activities for a total 8 out of 22 test systems was 
   performed and completed across 2 sites.
·  Tight schedule and planning for IQ so as not 
   affecting production shipment for April month end.
·  3 full time headcount including OT together with 
  Milford and Contract Manufacturer test operators
  was involved in completing the IQ activities.




IQ – Detector firmware downloader at Flex in July2018
4 FTN testers, 4 pump testers
1 PDA/eLambda linearity tester IQ

Reducing downtime in Venture
FTN  reduced testers unscheduled downtime 50% by performing weekly BSM pump priming and daily tester rebooting
BSM/QSM dynamic leak failure rate redueced 3% by implement the periodically changing of test tubing
Field Feedback investigation for 	GPV and FTN seal

Complete tester health review








