| ### Document Title: Detector Shunts Preventive Maintenance Check List |  | ### Engineering Change No.: ### NA | ### Document #: ### WAT-RW-SKI04568 | ### Organization: ### Kallang |
| --- | --- | --- | --- | --- |
| Revision: A |  | Document Owner: Joel Mamalateo Deduyo | Customer: WATERS | Effective Date: FlexQ – On File |
| Applicable To: | Waters Detectors Projects |  |  |  |

| ### APPROVALS |  |  |  |
| --- | --- | --- | --- |
| NAME | TITLE | DEPT NAME | DATE |
| Nallajanula Phani Bhushan | Staff Engineer | Engineering | FlexQ – On File |
| Joel Mamalateo Deduyo | Senior Engineer | Test Engineering | FlexQ – On File |
| Chua Kim Seng | Staff Engineer | Quality Engineering | FlexQ – On File |
| REVIEW / APPROVAL BY CUSTOMER |  |  |  |
| NAME | TITLE | SIGNATURE | DATE |

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| REV. | ECN | DESCRIPTION OF CHANGE | ORIGINATOR | RELEASE DATE |
| A | NA | Initial Release. | Mohammad Shamsul Bin Afandi | FlexQ – On File |

This document is proprietary and confidential property of Flextronics.

PURPOSE:  To perform preventive maintenance for detector shunts to ensure they are in good condition prior to production.

INSTRUCTIONS:  Number below correspond with the numbers in the parentheses ( ) on the form.

| Item | Entitled | To be done by | Enter |
| --- | --- | --- | --- |
| 1 | EQID No | Engineer Assistant | The PM ID (e.g. PM01481). |
| 2 | Results | Engineer Assistant | The value of the measured shunt dimension / the shunt energy. Write NA if not applicable. |
| 3 | Criteria Met? (Yes / No / NA) | Engineer Assistant | Check the box of the applicable acceptance criteria if (Yes / No / NA). |
| 4 | Remarks | Engineer Assistant | If have remarks, must write in details. If no remarks, leave it blank. |
| 5 | Completed by | Engineer Assistant | Employee number, signature and date to indicate PM was performed. |
| 6 | Reviewed by | Engineer Assistant | Employee number, signature and date to indicate PM was performed. |

TRAINING:  Prior to using, users must be trained on this form.

| Preventive Maintenance Checklist for Detector Shunts (Quarterly) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| (1)EQID no : |  |  |  |  |  |  |
| Tick the appropriate box for the choice of shunt to be performed for Preventive Maintenance |  |  |  |  |  |  |
| TUV(p/n:289002450) ☐ |  |  | PDA(p/n:289002149 / 289002226 / 289002227) ☐ |  | UPC2(p/n:289007555) ☐ | Velox(p/n: 289007898) ☐ |
| Dim 1(0.1586 ± 0.0004 ) inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |  |  | Top hole 0.218 ± 0.005 Inches (Drawing Reference: 405001136) Pass ☐ / Fail ☐ Value : |  | 0.1974 ~ 0.1982 inches (Drawing Reference: 405014777) Pass ☐ / Fail ☐ Value : | Dim 1(0.1586 ± 0.0004 ) inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |
| Dim 2 0.1586 ± 0.0004 inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |  |  | Bottom hole 0.218 ± 0.005 Inches (Drawing Reference: 405001136) Pass ☐ / Fail ☐ Value : |  | 0.1974 ~ 0.1982 inches (Drawing Reference: 405014777) Pass ☐ / Fail ☐ Value : | Dim 2 (0.1586 ± 0.0004) inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |
| Dim 3 0.190 ± 0.005 inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |  |  | Shunts energy @ Wavelength = 230nm Shutter position = open in Known Good Check the box to select the size: ☐ For Large shunt @ 5ms > 150,000 counts ☐ For Small shunt @ 8ms > 40,000 counts ☐ For Medium shunt @ 6ms > 90,000 counts Pass ☐ / Fail ☐ Value: |  | Shunts energy @ Wavelength = 230nm Shutter position = open in Known Good 90K ~ 250K counts Pass ☐ / Fail ☐ Value: | Dim 3 (0.190 ± 0.005) inches (Drawing Reference: 405010685) Pass ☐ / Fail ☐ Value : |
| Shunts energy @ Wavelength = 230nm Shutter position = open in Known Good Unit Sample Energy > 150 nA Reference Energy > 200 nA Sample: Ref: Pass ☐ / Fail ☐ |  |  |  |  |  | Shunts energy @ Wavelength = 230nm Shutter position = open in Known Good Unit Sample Energy > 150 nA Sample: Ref: Reference Energy > 200 nA Pass ☐ / Fail ☐ |
| Visual Check (Any physical damage) : Yes / No (If Yes, Please indicate at remarks) |  |  | Visual Check (Any physical damage) : Yes / No (If Yes, Please indicate at remarks) |  | Visual Check (Any physical damage) : Yes / No (If Yes, Please indicate at remarks) | Visual Check (Any physical damage) : Yes / No (If Yes, Please indicate at remarks) |
| Remarks : |  |  |  |  |  |  |
| (5)Completed by (Badge No. / Sign / Date) : |  |  |  |  |  |  |
| (6)Reviewed by (Badge No. / Sign / Date) : |  |  |  |  |  |  |

Preventive Maintenance Checklist for Detector Shunts (Daily)

Please check the following accordingly:

- Any damage found on shunt.

- Any missing component (e. g: screws)

- Pin guide measurement base on type of shunt as per specification.

EQID No: _______________________          Month: _______________________            Year: _______________________

| Date | Pass / Fail | Checking Done By | Remarks |  | Date | Pass / Fail | Checking Done By | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  | 17 |  |  |  |
| 2 |  |  |  |  | 18 |  |  |  |
| 3 |  |  |  |  | 19 |  |  |  |
| 4 |  |  |  |  | 20 |  |  |  |
| 5 |  |  |  |  | 21 |  |  |  |
| 6 |  |  |  |  | 22 |  |  |  |
| 7 |  |  |  |  | 23 |  |  |  |
| 8 |  |  |  |  | 24 |  |  |  |
| 9 |  |  |  |  | 25 |  |  |  |
| 10 |  |  |  |  | 26 |  |  |  |
| 11 |  |  |  |  | 27 |  |  |  |
| 12 |  |  |  |  | 28 |  |  |  |
| 13 |  |  |  |  | 29 |  |  |  |
| 14 |  |  |  |  | 30 |  |  |  |
| 15 |  |  |  |  | 31 |  |  |  |
| 16 |  |  |  |  |  |  |  |  |

Review by (Name/Sign/Date): _______________________