## Sheet1
Novo Nordisk - Alliance Issues reported by Didier Coyman/Mark Wrescahk - 04 September 2019
Comments: | Data provided around bad reproducibility of these systems - the first injection is typically low area count then consistent data obtained after that injection.
Various loop/syirnge and tubing changes performed by customer
Test failures are for custoemr PQ tests - NOT waters SQT tests (However customer has used these proceudres on many systems for a number of years)
No mention of any checks on needle alignment.
No mention of any solvent leaks from base of seal packs.
J18SM4860A:
2019-01-01 | DAEDB5UHL703 | Installed | No issues
2019-02-01 | DAOEB9BKEE12 | Precison test failure | SLT fails. New check valves installed. SLT passes. V1 rebuilt (not necessary). V3 leak observed. Reseated connection at V3
2019-05-01 | 100000947145 | Syringe leak | Syringe leak. New syringe fitted.
2019-05-01 | DALNBC6D7213 | Peak tailing | V2 replaced. Seal pack rebuilt inc needle.
2019-06-01 | DALNBC6D7213 | Screw sheared in injector | Front retained screw on injector assembly sheared when removing injector to rebuild seal pack (above call). Alliance unit RA
J18SM4855A:
2019-04-01 | DAEDB5UHL601 | Installed | No issues
2019-05-01 | DARNBCCCKH58 | Reproducibiltiy issues | CIT has information from multiple systems…..
2019-05-01 | DARNBCCCKH58 | Gradient test fail | GPV replced. (NB Unit is within possible bad GPV serial number range) | GPV issue with valves not opening as expected. Reported to potentially affect B18-K18 serial numbers. Refer to Servcie alert WKB56360
2019-06-01 | DAINBDHCAA14 | Peak tailing | New injector installed. Peak shape now OK.
2019-07-01 | DAOEBDYFM716 | RT shifts | Issue is with CH not alliance. HC engine at fault.
2019-08-01 | DARNBENHQY42 | Bump after gradient. Unexpected baseline shape following gradient change in method | ONGOING - LIKELY INSTRUMENT WILL GET REPLACED. SAME FAULT AS WITH J18SM4852A. Latest update (10Sep19) - Unit inspceted for needle alignment (it was good). However FSE found that screws holding seal pack to injector assmebly were not fully tight, so this has been corrected. Adjust seals values now more typical of what we would expect. Needs to test the system again to see if this makes any differernce to the issues being observed.
J18SM4872A:
2019-04-01 | DAEDB5UHL804 | Installed | No issues
2019-05-01 | DARNBCD9M808 | Gradient test fail | GPV replaced (NB Unit is within possible bad GPV serial number range) | GPV issue with valves not opening as expected. Reported to potentially affect B18-K18 serial numbers. Refer to Servcie alert WKB56360
2019-05-01 | DARNBCD9M808 | Bad reproducibility | V1/V2 repalced. Seal pack replaced.
2019-06-01 | DARNBCD9M808 | Bad reproducibility | V1 rebuilt
2019-05-01 | DARNBCD9NW48 | Bad reproducibility | Alliance Unit RA
J18SM4852A:
2019-04-01 | DAACB5TJFX28 | Installed | No issues
2019-05-01 | DARNBCCCL249 | Gradient test fail | GPV replaced (NB Unit is within possible bad GPV serial number range) | GPV issue with valves not opening as expected. Reported to potentially affect B18-K18 serial numbers. Refer to Servcie alert WKB56360
2019-06-01 | DARNBDCFEZ21 | Peak tailing | Injector to column tubing and injector replaced. No improvement. Switched column. No improvement. V1/V2 replaced. No improvement. Replaced injector again. Peak shape now acceptable
2019-08-01 | DARNBENHQY42 | Bump after gradient. Unexpected baseline shape following gradient change in method | ONGOING - LIKELY INSTRUMENT WILL GET REPLACED. SAME FAULT AS WITH J18SM4855A
J18SM4874A:
2019-04-01 | DAEDB5UHL905 | Installed | No issues
2019-05-01 | DARNBCD9NW48 | Bad reproducibility | V1/V2 and seal pack replaced. Fault not corrected
2019-07-01 | DARNBCD9NW48 | Bad reproducibility | Alliance Unit RA