## Attendance
DATE | FLEX (OPS) |  |  |  | WPPL (OPS) |  |  |  |  | REMARKS
| PHANI | JOEL M.D. | GARY | ALDWIN | EDWIN | DESMOND | YEE HWA | HUA ZHEN | GUAN YEE
2020-01-08 | N/A |  |  |  | N/A | N/A | N/A
2020-01-15 | N/A |  |  |  | N/A | N/A | N/A |  | N/A
2020-01-22 | N/A |  | N/A |  | N/A
2020-01-29
2020-02-05
2020-02-12
2020-02-19
2020-02-26
2020-03-04
2020-03-11
2020-03-18
2020-03-25
2020-04-01
2020-04-08
2020-04-15
2020-04-22
2020-04-29
2020-05-06
2020-05-13
2020-05-20
2020-05-27
2020-06-03
2020-06-10
2020-06-17
2020-06-24
2020-07-01
2020-07-08
2020-07-15
2020-07-22
2020-07-29
2020-08-05
2020-08-12
2020-08-19
2020-08-26
2020-09-02
2020-09-09
2020-09-16
2020-09-23
2020-09-30
2020-10-07
2020-10-14
2020-10-21
2020-10-28
2020-11-04
2020-11-11
2020-11-18
2020-11-25
2020-12-02
2020-12-09
2020-12-16
2020-12-23
2020-12-30

## Minutes Of Meeting-CLOSED (hidden)
ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | ASP/TSP Matrix Development - REFER HISTORY from 20th Jun 2018 email 30 May 2018 - To update and maintain the tracker to be on current state (6 Jun 2018 onwards) 11 Jul 2018 - target ASP BSM\QSM update by 31 July 2018. 15 Aug 2018: Pending ECR/ECO . 17 OCt 2018 : 1 ECO submitted and rest are pending/ to trigger Gary once approve 16 Jan 2019: Details are updated at "ASP TSP Summary" | Gary/Yee Hwa | 2017-07-06 | TBD | Closed | Low
7 | Implentation for Alliance Template 1.9 - Hua Zhen to verify the chemical part number as per WAT271129 Rev ? - after resolving the Carryover Test (High Blank) - 18 Jul 2018 24 Oct 2018 ONHOLD ACTIVITY 16 Jan 2019: No Template 1.9 will be impletemented. New testers are expected to be send to flex - TBA | Hua Zhen | 2018-06-20 | 6/27/2018 7/11/2018 12/30/2018 | Closed | Low
8 | Consigned Equipment Health Review Identify the list of consigned equipment.Total 56 equipment. Identify the downtime data on the consigned equipment Based on Jan-Jun data. Establish a checklist to review the health of the equipment.Avaialable Start the review of the health of the equipment. Review the status of the health with Waters team. Take actions and close on the problematic testers. Continue monitor for effectiveness - Dashbord update in progress for Alliance - 18 Jul 2018 - Alliance Dashboard : completed for submissiont 01 Aug 2018 (time frame CEHR) - Pump Dashboard completed - submitted on 31 Aug 2018 (resend the original file) - 17 Oct 2018 Identify and try to source locally available parts (PUMPS Testers) - 21 Nov 2018 Prioritize PUMP Tester (WPPL to source for PCBA and Pump Control Module). Target 07 Dec 2018 28 Nov 2018: Joel to follow up with a OESR (Pump) so that WPPL can work on the sourcing. 05 Dec 2018: Joel will send out the OESR by today 19 Dec 2019: To send Detector Health review 16 Jan 2019: Will be review through Project meeting | TAN CK/JOEL | 2018-07-07 | 2018-08-31 | Closed | Medium
9 | eDHR into Engg discussion Kick off meeting and align with CFT team on proposal Work with IT on configuration setup Matrix preparation Proposal and change request to customer Verify in test environment Release to production continue monitor and update 16 Jan 2019: Will be review through Project meeting | TAN CK | 2018-07-07 | 2018-11-30 | Closed | Low
10 | Monitor chemical for Carryover Test as per Rev 1.8 - 19 Sep 18 Hua Zhen to follow-up with Milford the implementation of Alliance Template 1.9 - Flex to continue the tracking and record potential issue with new Seal Pack, no details from Eng'g to share but higher suspect. Once design review finish with new seal pack and may have FW upgrade later. If FT performance did well in Flex, Engg will accelerate the upgrade process. - 17 Oct 2018 - To trigger 2 months ahead / During weekly Eng'g meeting and through email - 24 Oct 2018 refer to "CONSUMABLE TAB" - 21 Nov 2018 - To follow up with MCotter for chemical - Yee Hwa - 28 Nov 2018 - Yee Hwa will follow up to get additional 16 boxes that can last till early Jan19 - 05 Dec 2018: Joel to check and update on the qty. - 12 Dec 2018: return 1 box YH 16 Jan 2019: To review based on sheet "Consumables" | JOEL MD | 2018-07-25 | TBD | Closed | Medium
11 | Consigned Health review on Tooling Identify the list of tooling (approx. 275) from Su Fern Priority by family (alliance, pumps, BSM and Detectors) 2 weeks / family - 19 Sep 2018 Alliance looks OK, Test Equipment will be to prioritized - 26 Sep 2018 Onhold until tester - 09 Oct 2018 to discuss further the bandwidth? - 17 Oct 2018 - may need to push thrugh because of Oct'18 Audit - 31 Oct 2018 - resume for tooling with plan and GANTT Chart - 28 Nov 2018 - Discussed on the GANTT Chart - 12 Nov 2018 Item 1 completed (refer to CJFT Tab) - 19 Dec 2018 To provide details CK Tan as per CJFT Tab " Analyze downtime data " 16 Jan 2019: Will be review through Project meeting | Gary /CK | 8/15/2018 01 / 11/ 18 | 2018-12-31 | Closed | Medium
16 | Pump eEvent Log 05 Dec 2018: To generate a configuration file for Pump family 10 Jan 2019 : Awaiting review from YH / file submitted 27 Dec 2018 16 Jan 2019: Will be review through Project meeting | CK/ Joel | 2018-12-05 | 2018-12-31 | Closed | Medium
17 | CPU 2000 – 210000520 - REFER HISTORY from 20th Jun 2018 email - 210000520 to try on other model and pass is shippable. (YH) - to prepare a plan for 7 pcba return to Flex 05 Dec 2018: Yee Hwa to review with Melody to close this action item 19 Dec 2018 To prepare slide comparission for Year to Year CPU 2000 Performance 16 Jan 2019: Joel to consolidate the performance and close the monitoring YYYY - 2016 43/12604 (0.3411%) YYYY - 2017 32/12591 (0.2541%) YYYY - 2018 20/12315 (0.1624%) as of Nov 2018 (Dec 2018 Not Available) | Team | 2017-07-27 | 7/27/2018 8/31/2018 12/31/2018 02/28/2019 | Close | Medium

## Minutes of Meeting-OPEN
ACTION ITEMS DISCUSSED
S/N | DECSRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | Detector PCBA Fixture Qualification Plan:TMD execution, Fixture registration, ASP update 16 Jan 2019: Yee Hwa to trigger TMD for eval run by 25Jan19 27 Feb2019: Yee Hwa to trigger TMD for eval run by 01 Mar 19 05 Apr 2019: TMD 010527v00 start implentation date - Gary to send tracker (excel spread sheet) and schedule 24 Apr 2019: Gary to share the SN List of UUT 21 May 2019 : Serial Number to publish by Gary 12 Jun 019 : TMD 010527v00 evaluation Closed | GLY SFC | 2019-01-16 | 30 Jun 19 | Close | Medium
2 | BSM/QSM PCBA Fixture qualification - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | GLY GYC | 2019-06-12 | TBA | Open | Low
3 | Weekly Downtime Update: Flex to the detail for down equipment every Thursday 5:30PM -Monthly - every end of the month: Dashboard / end of the month -to add summary page on the downtime report / weekly 24 Apr 2019: To include tools/equipment 11Dec2019: Target Closure change from Routine to 24-Apr-19 | JOEL MD | 2018-06-20 | 2019-04-24 | Routine | Low
4 | HDD Back Up/Imaging of all CPU as per PM Checklist: History: 2TB External HDD/Acronis imaging s/w provided by WPPL to backup all consigned testers. Objective :Flex to check with internal IT to HDD Back Up/Imaging of all CPU - In progress - 09 Oct 2018 - Raise Request for 3 Qoutation for 2 HDD - 17 Oct 2018 to discuss further as Flex not allowed to have external HDD inside facility - 24 Oct 2018 WPPL will consign HDD with memo to backup/imaging. Cloud service can also be considered. - 16 Jan 2019: Joel to follow up with Hua Zhen on this. - 14 Feb 2019: Refer to attached excel file) - 27 Feb 2019 To sched one unit a day and plan and create a separate tab for imaging - 12 Jun 2019 After ECO/Validation/Update - shall update the Back-up HDD accordingly - 19 Jun 2019 To backup TE82969-1 and all TE82882 Stations after validation To backup TE82943-11 to 18 after validation - 26 Jun 2019 Perform imaging periodically (6 months as a guideline and if there's ECO \ Firmware update. (JMD) 11Dec2019: Target Closure change from Routine to 26-June-19 and Status change from Onhold to Routine. To submit lastest backup/ imaging list and plan or schedule. David to discuss with Joel. | Joel | 2018-09-19 | 2019-06-26 | Routine | Low
5 | AIR COMPRESSOR requirement: History: Found filter inside Pump tray tester with oil/Water Objective : To provide Oil/Water free compressed air to Waters testers Refer to last Jan 2019 Updates. 12 Feb 2019 - Measurement of Outlet / Inlet of Compress Dry Air at Pump Tray Area (suspected to be the worst condition as last position of the air supply) 27 Feb 2019 To install air filter after recieved the report and monitor the performance, with date of installation and PM. 10 Apr 19 - to identify filter to suit the requirement 21 May 2019 - Equipment Calibration by vendor 29 May 2019 - Target to execute by 03 Jun 19 12 Jun 2019 - Measurement completed 03 Jun 2019 and report to share by 26 Jun 2019 (JMD) 26 Jun 2019 - Tan CK and Vendor in progress of verification of test result (Tan CK) 03 Jul 2019 - Oil and dew pass in requirement but fail in particulate count / get the raw data and recalculate (TCK) 10 Jul 2019 - Flex Facilities will install SMC filter which will control at less than 100 particles of 0.3um at the outlet and re-conduct a third party verification again. 17 Jul 2019 - Third party verfication is completed dated 11Jul19, pending for report. 24 Jul 2019 - As committed by vendor, report will send to flex by Thursday (25Jul19) - CK 31 Jul 2019 - Used filter SMC , Model : AMF350C-N04B-R (cost S$270 each) showing positive result and able to meet Waters specification. 14 Aug 2019 - To verify if filter in place (Joel) 21 Aug 2019 - varified that mainline is using SMC , Model : AMF350C-N04B-R and test result shows all within Waters specification. Refer to attached test result | Joel / Hua Zhen / CK Tan | 2018-09-19 | 31-Dec-2018 26-Jun-2019 03-Jul-2019 26-Jul-2019 31-Jul-2019 21-Aug-2019 | Close | Low
6 | Commissioning of New Seal Pack Jig Plan:TMD execution, Fixture registration, ASP update 16 Jan 2019: Yee Hwa to trigger TMD for eval run 27 Feb 2019 : TMD 010639 has been approved at Waters, release to Flex 28-Feb : Expedite return of fixture from supplier. 01-Mar : DE routed for approval in Flex System. 05-Mar : Received the fixture from supplier. 05-Mar : PM Documentation complete. Training completed. 06-Mar : Implemented TMD010639 v00 P/N: 186269506 : SN C19SM4920A P/N: 186269505 : SN C19SM7166A 17 Apr : WPPL send the drawing to Milfrod for update 24 Apr : Gary to send PACK and Go Package in Solidworks 21 May 2019 : Completed and close | Gary | 2019-01-16 | 2019-04-30 | Close | High
7 | Waters Server installation 28Nov18: Currently 2x servers in rack. Recieving 3rd server. - Collection of 2 server on 03Dec18 from WPPL. (Kallang and Senai) - 1 server to setup at Kallang on 04Dec18. - 1 server to setup at Senai on 05Dec18. - WPPL plan to decommission Kallang old Waters server on Jan19 05Dec18: Will update the status by this week - New Server (ST: FVDCBS2) installed 04 Dec 2018 - utilized existing UPS and to ship new UPS back to WPPL as SPARE - Request to SCRAP the OLD server and defective UPS 09Sep19: JMD send follow-up Whats app message to HZL for update from Waters IT Team 18Sep19: HZL advised to witness the decomissioning of the old sever by 4:00 PM 25Sep19: Pls. refer to previous minutes for the deleted discussions: Server is working and no issue observed for the whole week (Laber printing and test reports) HZL to check on scrapping the Server through local IT Support 27Sep19: HZL Updated that Flex to send the Server HDD to Waters and Scrap Certificate to WPPL. 09Oct19: JMD to get vendor that will scrap and issue certificate to vendor/Server HDD to return to WPPL 16Oct19: JMD to send to scrap today and vendor mentioned that Certificate Of Destruction available after 3 weeks only (to recieve by 30Oct19) 13Nov19: JMD emailed to HZL the Certificate Of Destruction and Scrap Form 3291 | Shamsul HZL | 2018-11-21 | 31-Dec-2018 31-Mar-2019 30-Jun-2019 16-Oct-2019 30-Oct-2019 | Close | Low
8 | Competency Tracking - Alliance 10 Apr 2019 : Need to provide objective evidence of closure, eg photo - Open item to provide target date 24 Apr 2019: to update the progress 21 May 2019 : 8X still need to process, ionized gun and fixturing in-progress target by - to send mark up ASP to WPPL 23 May 2019 (Tan CK) - Fixture ETA 24 May 2019 12 Jun 2019 : CK Tan to update the tracker and send out by end of today - Air-gun PO released to vendor TBD - Torque driver to have a two way driver and RFQ in progress / approx S$800 - 19 Jun 2019 - pending ECR 9229947 Approval stage / Airgun waiting for E.T.A. date by today (19 Jun 2019) tooling PN 999025154 by 21 Jun 2019 (CK Tan) Torque Wrench - use current torque wrench and no issue with operation purchase of new torque wrench for improvement will be under separate item for review - 26 Jun 2019 Next week to have 1x Airgun and 1x tooling PN 999025154 (Tan CK) - 03 Jul 2019 Delivery confirmation to follow-up for Airgun (CK Tan 07 Jul 2019) - 10Jul 2019, Air gun received on 04Jul19 and installed on 05Jul19. Tooling 999025154 received on 08Jul19. | Tan CK | 2019-04-10 | 30-Apr-2019 30-May-2019 30-Jun-2019 05-Jul-2019 | Close | High
9 | 210000415 Shutter Home Failure Investigation - Install to running model 186015032 and observe repeatability of failure by 14 Jun 2019 - Capture and share hyperterminal data to team 19 Jun 2019 No build PDA and eLambda from 14 to 19 Jun 2019, Golden Unit is not able to capture the reject , will check next PDA build (19 Jun 2019 Joel) - will try using elambda and request for build 1 PDA unit to use for verfication (19 Aug 2019) 26 Jun 2019 Verified using eLambda and PDA Good Unit - Found to be NTF / release the board to the line | Joel | 2019-06-12 | 2019-06-26 | Close | Medium
10 | Detector PCBA Fixture Implementation - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08-Jan-2020 : ECR 9233596 waiting for submission 22-Jan-2020 : 2x model affected by TMD010739 will have a separate item to address Gary to prepare a summary slide for 7x model implemented to close the action item | GLY SFC | 2019-06-12 | 30 Sep 19 30 Dec 19 07 Feb 20 | Open | Medium
11 | TE82969 and TE82882 (7x testers) Validation to accomodate new spares (OESR 193421) -12 Jun 2019 : To try the new Linearity Test Software ver 2.2 \ Final Test ver 4.2 -19 Jun 2019: Test in Progress, refer to daily update to team / Get VAF Approval Status (Joel) -26 Jun 2019 HZ preparing IQ for submission / JMD to send test report today -03 Jul 2019 Eric Turner increase empower database to remove loss comm (Joel to try today) - To check with Eric on how to increase empower data base (HZ) 10 Jul 2019 - In procees. HZ have increase database size from 50mb to 1500mb for all PDA tester. Joel to monitor if there is any loss comm issue arise from today onwards. | Joel / Hua Zhen | 2019-06-12 | 6/19/2019 6/26/2019 7/05/2019 | Close | High
12 | TE82943 (8x testers) Qualification to update Firmware & driver Pack - 12 Jun 2019 TMD010661 v01 release and effective 03 Jun 2019 - 19 Jun 2019 TE82943-11 Test Station updated to version 1.8 and ver 2.3, Test Report submitted, pending approval from Milford. (Yee Hwa) - 26 Jun 2019 TE82943-12 Test Station updated to version 1.8 and ver 2.3 last night, Test report submitted to YHC - 03 Jul 2019 TE82943-11 Approved / TE82943-12 IQ Report approval in progress (YH) - 10 Jul 2019: 2 testers are routing for approval - In process - 17Jul 2019: 14 - 18 testers are still in process - 24 Jul 2019: In-process. By this month to complete. (Yee Hwa) - 31 Jul 2019: TE82943-16,17 and 18 on-hold for a week as E. Jackson is onleave. - 07 Aug 2019:TE82943-16,17 in-progress of testing (Joel) - 14 Aug 2019 TE82943-16 and 17 completed and submit to YH (TCK Approved),TE82943-18 Retest is progress (Joel) - 21 Aug 2019 TE82943-18 Test Report, Calibration Report with Evaluation Record for EQID 0497 has been submitted to CYH - 28 Aug 2019 To check with ITS if able to do the PM and verification before 30 Aug 2019 (JMD) Jia Chen Lim (ITS) confirmed schedule balance PM and verification by 02 Sep 2019, due to tight schedule -02 Sep 2019 - ITS Jia Chen verfied balance and found to be within specification -04 Sep 2019 Flex to arrange external calibration for Balance (JMD - arranged 13 Sep 2019) -11 Sep 2019 Ming Deng calibrated AND HR 200 Balance and found to be within spesc (waiting for official report) -24 Sep 2019 EQID 0497 SN 12326186 Calibrated and within specs released to production. VAF not required as per HZ advise (to follow 730001835) | JMD YHC HZL | 2019-06-03 | 15-07-2019 31 Jul 2019 30 Aug 2019 27 Sep 2019 | Close | High
13 | Carryover FPY Tabulation for Alliance - 12 Jun 2019 Continue the tracking and send to WPPL - every Friday Morning (9:30am) - 19 Jun 2019 Routine submission as is. - 26 Jun 2019 Routine submission as is. - 03 Jul 2019 Understand the 5% on WW26 and set trigger point - 17 Jul 2019 Continue monitoring - 24 Jul 2019 Continue monitoring - 11Dec 2019: To close this activity and go by routine monitoring by production line. Any spike to trigger WPPL >95% only for Carryover | Joel / Jing Xuan | 2019-06-12 | 2019-07-31 | Routine | Low
14 | JBS For Alliance Needle Alignment - 12 Jun 2019 CK Tan to start drafting information - 19 Jun 19 : JX will share template to CK - 21 Jun 19 CK Tan to start drafting information and share with WPPL - 03 Jul 19 - 26 Jun 2019 JBS preparation of visual aid (Tan CK) - 03 Jul 2019 CK to send out Draft copy for review (CK) - 10 Jul 2019 Revisit on the Draft JBS and enhance on the KeyPoints and Reason section. - 17 Jul 2019: CK to revised and send out to WPPL for review - 24 Jul 2019: Completed reviewing with Desmond - Send out to Edwin and Desmond by this week. - 31 Jul 2019: Desmond commented the JBS and CK to update accordingly - 07 Aug 2019: JBS is send to Desmond dated 31Jul19. Pending for Desmond's advice on the next course of action. - 14 Aug 2019: Pending for Desmond's advice on the next course of action. 28 Aug 2019 : JBS agreed, Desmond to pass the JBS to CK and insert as per formatting 04 Sep 2019 : CKT to verify ASP. 11 Sep 2019 ECR 9231435 - Review and approval in progress in Waters 18 Sep 2019 Track as per ECO process | CKT JXL | 2019-06-12 | 12-Jul-2019 17-Jul-2019 30-Sep-2019 | Close | Medium
15 | Consign PCBA Debugging (Alliance Instrument) at Senai 26Jun19 to 06 Nov19 refer to previous minutes 13Nov19: Training reschedule due to QBR Preparation. Reference to used ASP 279002695 and TSP WAT271129 and event log recording. 20Nov19: Training completed on the ASP 279002695 and TSP WAT271129 and event log recording.(to send event log sample to WPPL) 27Nov19: Training conducted last 15Nov19 for event log and TSP. | JMD DA | 2019-06-26 | 06 Nov 2019 30 Nov 2019 | Close | Medium
16 | OFF Line Verification 1. Known Good Board Verification must have SN of the UUT as well as failure symptom 2. Action taken to verity the unit 3. Timeline of the verification process workflow by next week (JMD) | Joel | 2019-06-26 | 2019-07-03 | Close | Medium
17 | TUV Flow Cell Low Energy Level. - Need to check the TUV Shunt Tranmittance value. (5 data points each shunt) - review the ASP and TSP - PM to submit (Joel) 21 Aug 2019 - Joel reviewed TSP 735001356 TUV Detector, Final Test Procedure and it does not specified transmittance value TUV TSP's . WPPL to advise after verification from Milford 28 Aug 2019 - Hua Zhen to follow-up transmittance value to Milford (JMD to HZ) 04 Sep 2019 - HZL send email inquiry to MJ Milford (pending advise) 16 Sep 2019 - Mark Jussaume update no limit and sees reading 0f 80, he put Eric in list hoping may have some add'l knowledge of the shunts 18 Sep 2109 WPPL to follow-up in knowledge base info. | JMD HZL | 2019-08-14 | 2019-08-30 | Close | Medium
18 | 186003031 - X-bridge Column Replacement Frequency - to check with HZ the criterial for Column Replacement Frequency 21Aug19 JMD emailed HZ to seek advise for frequency of replacement 28Aug19: Avg per port is 23.7 times (round up to 24) Per test, number of injection is 53 times. Avg per column injection every mth is 24 x 53 = 1272 Pls. advise what is the max. injection before we can replace column (Hua Zhen to follow-up Milford or Mr Wong) 04Sep19: Mark Jussaume advised that 6 months replacement of column is not applicable to Flex at the moment. Its only applicable to new Alliance tester. And the replacement interval hasn’t been finalized yet. So please keep monitoring the usage of column and record the frequency of replacement. 11Sep19: Flex replaced 1X Column Alliance Final Test 3.4 dated 08 Jul 2019 (Chemistry Test consistent fail) 13Nov19: Flex replaced 1X Column Alliance Final Test 4.4 dated 07 Nov 2019 (System Overpressure) 27Nov19: (move to monitoring phase) 11Dec19: Started monitoring for 6 months and result only 2 replacements. To monitor for anotther 6 months. Status change from Monitoring to Routine. | David HZL | 2019-08-14 | 20 Jun 19 31 Dec 19 30 Jun 2020 | Routine | Low
19 | OESR # 193460 Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 2020-01-31 | Close | Medium
20 | OESR # 193461 Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 31-Jan-20 31-Jul-20 | Close | Medium
21 | OESR # 193124 Additional Capacity for Alliance Pump Tray Tester (TE82526) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 2020-01-31 | Close | Medium
22 | OESR 190059 Alliance Final Test 3 Frequently Freeze During Final Test OESR 193244 Milford OESR - 03 Apr 2019 04 Apr 2019 - Matt Brule request Flex to provide detail on Hang issue 05 Apr 2019 - Joel provide detail to Matt Brule for detail of Hang issue 04 Sep 2019 - Joel to monitor and update by 06 Sep 2019 06 Sep 2019 - JMD sent monitoring data to HZL (refer rto attach Alliance FT3 System Freeze Tracking Record) 09 Sep 2019 : HZL updated Matt Brule with attached Alliance FT3 System Freeze Tracking Record 18 Sep 2019: HZL last occurrence is 3 Sep 19, slow build for Alliance follows thus need to extend monitoring / with respective volume (min 100 units), HZL to inform Matt that Flex to further monitor this issue. 25 Sep 2019: JMD issue the tracket for 100 units to Prod'n and monitoring in progress 31 Oct 2019: Monitoring for 100 instrument competed> 23 units freeze / 100 units | JMD HZL | 2019-09-04 | 13 Sep 19 27 Sep 19 30 Oct 19 | Close | High
23 | OESR 190079 TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. 15May2019 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04Sep2019 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11Sep2019 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18Sep2019 : JMD will inform YHC once EQID 0497 is available 19Sep2019: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25Sep2019 VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. | JMD HZL | 2019-09-04 | 2019-09-27 | Close | High
24 | 190102 - TE82988-2 Power Supply 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
25 | 190104 - TE82988-2 Spare Parts holding at Senai - to request return back to WPPL 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10 Sep 2019 - Erzad update ETA back to WPPL by 23 Sep 2019 | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
26 | 190108 - TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep 2019 FLX Kallang JMD to follow-up to Edwin 05 Sep 2019 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
27 | OESR S190109 - TSP having 735002531 with 735001613 does not have consistent content 04 Sep19: reviewed with HZL/EK/DA and JMD and agreed to update the OESR content and its Model PN. 05Sep19: JMD updated the OESR 190109 and send to HZL 18Sep19: HZL to checked for update 24Sep19: HZL to follow-up with E Turner 27Sep19: HZL to follow-up with E Turner reminder due as Waters Annual Audit for Flex 08Oct19 : HZL send reminder to E.T 08Oct19 - HZL ask ET if need to revise procedure/monitoring or uprev the next revision. 13Nov19: HZL to check for update | JMD HZL | 2019-09-04 | 06-Sep-19 10-Oct-19 30-Dec-19 | Monitoring | Medium
28 | 190117 - TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05 Sep 2019 Start 20 Aug 2019 used in 210000411 SN 010919110M 11 Sep 2019 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
29 | OESR #190118/ #190119 Reticule Plate Broken Target Mirror 19Aug19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30Aug19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04Sep19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23Sep19: SFC relelase PO to vendor 10Oct19 PM documentation approval in-progress 16Oct19: Part lead time is 4week from now (16Nov19) 20Oct19: GLY assembled and release to Production. | JMD SFC | 2019-09-04 | 2019-12-30 | Close | Medium
31 | 190124 - 193544 Request for TE# for ICT Fixture 210001344 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05 Sep 2019 Erzad to send the documentation to JMD 06 Sep 2019 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11 Sep 2019 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
32 | OESR S190130 -193550 TE82882 PDA Final Test System Missing Spare OB PN 700011652 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18Sep2019 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25Sep2019 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26Sep2019 - TE82882-10SEP19-VAF_Approved and released tester for usage. | JMD HZL | 2019-09-04 | 06 Sep 19 30 Sep 19 | Close | High
33 | OESR S190132 - TE82943-13 USB Device Error 03Sep19 JMD raised OESR for TE82943-13 USB Device Error 04Sep19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06Sep19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10Sep19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11Sep19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18Sep19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25Sep19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09Oct19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 10Oct19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10Oct19: Informed purchaser to release pay using Paypal. 16Oct19: JMD to follow-up E.T.A from buyer 23Oct19: Bliss target ETA to Flex is on 29Oct19 30Oct19: To install today (Shamsul) and update status 06Nov19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | JMD HZL | 2019-09-04 | 06 Sep 19 18 Sep 19 25 Sep 19 30 Oct 19 | Close | High
34 | 190133 Label Printing 04 Sep 2019 GLY to follow-up with DA | GLY / DA | 2019-09-04 | 2019-09-06 | Close | High
35 | AND HR 200 Balance Misssing Segment Activity 25Sep2019: HZL to discuss with Edwin if scrap or repair .New Model GR 200 S$1765.5 Repair :TE82943-11 AND HR200 09Oct2019: To engage vendor to scrap with certification 16Oct19: JMD to send to scrap today and vendor mentioned that Certificate Of Destruction available after 3 weeks only (to recieve by 30Oct19) 13Nov19: JMD emailel to HZL the Certificate Of Destruction and Scrap Form 3291 | Shamsul HZL | 2019-09-25 | 02 Oct 19 16 Oct 19 30 Oct 19 | Close | Low
36 | Alliance Test Station Study Upgrade to use for common use to other model (Sample Manager) 09Oct19: HZL to study the upgrading for the test station to save space 16Oct19: Milford is studying the proposal 30Oct19: Result from Milford is not feasible to have a common test station. | JMD HZL | 2019-10-10 | 2021-01-31 | Close | Low
37 | OESR S190156 TE82503-5 Injector Seal Pack Ramp and Decay Test Fail 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | JMD HZL | 2019-10-03 | 2019-10-21 | Close | High
38 | Flex-WPPL-CEI Tripartite meeting (David) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08Jan2020: last update: Still under review by Waters Milford on the program proposed by CEI, information shared to WPPL Team Ops, any new information will update WPPL OPs Team and close this action item. | David HZL | 2019-10-30 | 2020-01-08 | Close | Low
39 | Cross Thread Analysis - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: ECO 9233201 to send to Flex today 15 Jan 2020: Start impelementation today | GLY GYC | 2019-11-06 | 2020-01-17 | Close | Low
40 | Hot Part - Return to Flex on 30 Oct 2019 - 700001031 removed from Alliance unit # L18SM4424A - RMA-SKW-00000136 - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: awaiting official instruction from WPPL SQE / Investigation from OPS Team already close as of 09 Dec 2019 | AP JMD GYC | 2019-11-06 | 2019-12-09 | Close | High
41 | Lean Project Update - 196015031 Failed Noise Test - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | AP JMD GYC | 2019-11-06 | 2020-01-03 | Close | High
42 | Additional tacker 11Dec19: Add OESR and TMD in this Engineering Meeting for the purpose of Status update and transparency from both Team (WPPL and Flex) | TEAM | 2019-12-11 | 2020-12-31 | Routine | Low
42 | Asset Manual Update - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: JMD presented to HZL and to email the Asset List by today | JMD HZL | 2019-12-24 | 2020-01-31 | Close | Low
43 | 210000415 Baking and Retest Evalutaion 15-Jan-2020 - Flex Senai to bake today and once finisk, will be cool down and zip lock before sending to Kallang 22-Jan-2020 - Baked 210000415 arrived in sealed ESD Bag with desiccant and good humidity indicator. We retest and the failure still the same : Shutter Fail to Home.(16 Jan 2020) Pending advise from GYC | JMD GYC | 2020-01-14 | 2020-01-22 | Open | Medium
45 | Consumables to include/update SPARE Parts for PM 15-Jan-2020 - Update in the tab for consumable. HZL will share the MRO form to be fill-in by Flex for ordering PM parts as well as the required consumables for Final Testers. Part to be direct ship to Flex. | JMD HZL | 2020-01-15 | 2020-01-29 | Open | Low
46 | QSM Interface Tubing 22-Jan-2020: Must continue the monitoring as per discussion on the MTR | JMD HZL | 2020-01-22 | TBD | Open | Medium
47 | Loss Comm to align TSP to Tester TE82882 PDA Final Tester 22-Jan-2020 TSP 735001613 Step 6.5 To troubleshoot Lost Comm error:, it has observe the empower Nodes haave some or less UUT setup. JMD and HZL to sudy on how to align the TSP and tester | JMD HZL | 2020-01-22 | TBD | Open | Medium
48 | Production utilization on tester 22-Jan-2020 To update the tester utilzation on the dashboard as per forecast | JMD HZL | 2020-01-22 | TBD | Open | Medium

## ASP TSP Summary (hidden)
Pumps | ASP | 15 | 0 | 15 | 4 | 11 | 26.67% | 7x ASP belong to sub assy & Main Line 4x ASP belong to Packing
| TSP | 2 | 0 | 2 | 1 | 1 | 50% | 1x TSP belong to 515 pump
Last update: 26 Sep 2018
Product Family | ASP/ TSP | ASP/TSP Compliance Check by Flex/Waters | [6 empty] | ASP/TSP Compliance Implementation by Flex/Waters |  |  |  | PKI | Arc Bio
|  | Total | Not in scope -Low runners (Velox/5034/5053/Arc Bio) | In scope (High Runners) | Completed Review | Pending Review | % Status | Pending Document Remarks | Completed Implementation | Pending Implementation | % Status | Remarks
eAlliance | ASP | 30 | 4 | 26 | 26 | 0 | 100% | NA | 26 | 0 | 100% | completed
| TSP | 10 | 0 | 10 | 10 | 0 | 100% | NA | 10 | 0 | 100% | completed
BSM/QSM | ASP | 15 | 3 | 12 | 12 | 0 | 100% | NA | 11 | 1 | 91.67% | 7x Red line documents (word format)submitted to WPPL for ECO. 4x competed review. No ECO required. 2x pending ECO. - 19 Sep 18 - 1X ECO release today - 31 Oct 18 - 1X ECO release today 16Jan19 - In process for ECR Target closure: 31Jan19
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed
TUV Detectors | ASP | 26 | 7 | 19 | 19 | 0 | 100% | NA | 11 | 8 | 57.89% | 8x Red line documents (pdf format)submitted to WPPL for ECO. Target 30 Nov 18 | 8 | 2
| TSP | 6 | 0 | 6 | 6 | 0 | 100% | NA | 6 | 0 | 100% | completed | 2
PDA Detectors | ASP | 39 | 13 | 26 | 26 | 0 | 100% | NA | 7 | 19 | 26.92% | 19x Red line documents (pdf format)submitted to WPPL for ECO. Target 31 Oct 18, 31 Dec 2018 | 8 | 2
| TSP | 2 | 1 | 1 | 1 | 0 | 100% | NA | 1 | 0 | 100% | completed
Pump | ASP | 16 | 0 | 16 | 14 | 2 | 87.5% | NA | 6 | 8 | 42.86% | 5x Redline done (PDF format), 3x Redline done (Word format) 7 - In process of ECR 1 - In process of verification with Waters Target closure:31 Jan 2019
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed

## Consumables
CONSUMABLES 22 Jan 2020
P/N | Description | Criteria for replacement | Where Used | Reference Document | Stock Balance (not under use in production) | CATEGORY | UNIT OF MEAS. | LEAD TIME WORKING DAYS | SHELF LIFE (Y/N) | SHELF LIFE (MONTH) | Customer supplier or Flex localize supplier | TOTAL REQUIRED PER MONTH | AVE. PER WK (BASED on 4 WKS) | DATE ARRIVED | Delivery Order # | DATE REQUIRED
WAT 042881 | Detector solution : Linearity Solution | Based on Expiration Date (in a box) 1 month after transfered to cuvette | 186002489 186248900 186002998 186299800 186015032 186015033 186015053 186015034 N2971014 N2971036 | 735003877 735003545 735002531 735003548 | 4 | Chemical | BOX | 44 | Y | 1 month | Waters | 1 | 0.25 | 2018-08-24 | 187212354 | 2020-08-04
WAT 042885 | Detector solution : Erbium Solution | Based on Expiration Date (in a box) 1 month after transfered to cuvette | 186002489 186248900 186002998 186299800 186015032 186015033 186015053 186015034 N2971014 N2971036 | 735003877 735003545 735002531 735003548 | 6 | Chemical | BOX | 44 | Y | 1 month | Waters | 1 | 0.25 | 2019-11-13 | Ship to No: 221751 | 2020-06-12
700004156 | Cuvettes | scratches, broken, wear and tear | 186002489 186248900 186002998 186299800 186015032 186015033 186015053 186015034 N2971014 N2971036 | 735003877 735003545 735002531 735003548 | 6 | Glass | BOX | 44 | N/A | NA | Waters | 1 | 1 | 2019-07-10 |  | Need to have a 1 to 1 exchange | Submitted wear and tear cuvettes to HZ 11Dec19: HZ provided 1 box of 6pcs Cuvettes. David to sent email request. To raise OESR | 24 Dec 2019 OESR S190201 raised 08 Jan 2019 Flex received cuvettes from HZL
700003741 | HPLC System UV Test Solutions (ampules 1, 3, 7, 8 & 9) | 3 box per week | Alliance Final Test | WAT271129 | 4 | Chemical | BOX | 44 | Y | 2 box per week | Waters | 28 | 7 | 2019-11-01 | 13x BOX 187965048 (01 Nov19) | 31-Jan-2020 | 11Dec19; updated usage from 3 to 2
186002127 | Green screw cap | 7/week | Alliance Final Test | WAT271129 | 88 | Cap | PACK | 22 | Y | 7/week | Waters | 28 | 7 | 2019-11-27 | N/A | 2020-04-08 | David to update arrive date
186002128 | Red screw cap | 7/week | Alliance Final Test | WAT271129 | 91 | Cap | PACK | 22 | Y | 7/week | Waters | 28 | 7 | 2019-11-27 | N/A | 2020-04-08 | David to update arrive date
186005827 | Black Screw op cap, Pre-Slit PTFE/Silicone Septa, 9mm | 7/week | Alliance Final Test | WAT271129 | 37 | Cap | PACK | 22 | Y | 7/week | Waters | 28 | 7 | 2019-10-16 | N/A | 2020-01-29
186003031 | X-bridge Column | every 6 months as per TSP 735003902 sec. 5.1 - to do data collection for replacement of column | Alliance Final Tester | WAT271129 | 0 | Column | PC | 22 | Y | NA | Waters | TBD | TBD | 2019-07-10 | NA | Flex column is >15 months old 1. Replaced AFT 3.4 (08 Jul 2019) System Overpressure 2. Repaced AFT 4.4 (07 Nov 2019) System Overpressure | JMD to work with HZ on replacement frequency for Column | 11-Dec-2019: Started monitoring for 6 months and result only 2 replacements. To monitor for anotther 6 months. Status change from Monitoring to Routine. 15-Jan-2020: Replaced Column for Alliance Final Tester 3.3
700005319-1 | GPV Test solution p/n : (Due date on 12 Jun 2020) | Based on Expiration Date | QSM | 735002155 | 4 liter\bottle | Chemical | 4LI /BOTTLE | 44 | Y | 12 | Waters | Based on Expiration Date | N/A | 2018-08-02 | 187136587 | 2020-03-12
402000351 | Lamp (The lamp is considered good if there is less than 2000 hours. However, if the TUV is being PM’ed at it’s yearly interval and if the TUV Lamp Hours are greater than 1000 hours, the Lamp must be replaced. ) | 2000Hr and average per year | Alliance cell:24 BSM/QSM cell:8 | 735003471 735001828 71500121902 pg 3-7 | 0 | UV Source | PC | 7 | Y | 32 per year | Flex | 2000Hr and average per year | N/A |  |  | NA
186002457 | White Screw cap with bonded, pre-slit PTFE/Silicone septa for 12x32 mm vial | 24 pcs per week | Alliance Final Test | WAT271129 | 100 pcs / pack | Cap | PC | 7 | Y | 24/week | Waters | 24 pcs per week | N/A
279001105 | Lamp power supply | wear and tear | adhoc | adhoc | NA | Electrical | PC | 7 | N/A | 1 to 1 replacement | Flex | wear and tear | N/A | N/A | N/A | N/A
279000852 | Power Supply | wear and tear | adhoc | adhoc | NA | Electrical | PC | 7 | N/A | 1 per detector | Flex | wear and tear | N/A | N/A | N/A | N/A
WAT035190 | In line Filter (To be replaced on all Tester using in line filter every 6 months) | wear and tear | adhoc | adhoc | NA | Plug | PC | 7 | N/A | NA | Flex | wear and tear | N/A | N/A | N/A | N/A
WAT025566 | Fitting Plug (To be replaced on wear and tear occasion) | wear and tear | adhoc | adhoc | NA | Plug | PC | 7 | N/A | NA | Flex | wear and tear | N/A | N/A | N/A | N/A
186000307C | - Screw Top Vial, 12x32mm with cap & preslit PFTE Siliconce Septa | Change new vial\cap every new carryover test | Alliance Final Test station | adhoc | 93 Box | Cap | PC | 7 | N/A | 50 / month | FLX-IX-2036-AR corporation. | Change new vial\cap every new carryover test | N/A | N/A | N/A | N/A
Waters PN | Description | PM:Routine/frequency | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ??? | ???
GUIDELINE :
- To trigger 2 months ahead
- During weekly Eng'g meeting and through email
Note on K11: 17 Oct 2018 trigger 3 months ahead 04 May 2019 - to trigger WPPL for replacement bottle due to expiration date

## OESR_v00 (hidden)
OESR ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | OESR # 193460 Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4: 74.1% Prime / 89.3 with Prime and Spare Q1-Q2 2020:Prime 79.3% / Prime + Spare 91.56% (HZ to update) 04Sep19 - HZL send reminder to MJ Milford 16Sep19 - MJ update the Milford is working concept to modernize this tester as duplicating it is impractical due to parts obsolecence 10Oct19: HZL to check for update from MJ for the concept of modernization of this tester 13Nov19:HZL to checked for update from Milford 27Nov19: HZL updated flowcell PN to MJ and prepare CEA for review. 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
2 | OESR # 193461 Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4 66.1% Prime / 74.7 with Prime and Spare Q1-Q2 2020: Prime 81.1% / Prime + Spare 85.5% (HZ to update) 04Sep19 : TE82520-3 CEA Approval loop 09Sep19: TE82520-3 CEA Approved/ HZL sent the CEA to Milford 10Oct19: HZL to check and advise the progress from Milford and respective completion date 16Oct19: ETD on Q2'20 advice from Chau Nguyen 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 31-Jan-20 31-Jul-20 | Open | Medium
3 | OESR # 193124 Additional Capacity for Alliance Pump Tray Tester (TE82526) 14Aug19: to put the TE number (JMD) (TE82526) 28Aug19: Milford in progress in building and updating the CPU to Windows 10 (HZ to update) 03Sep19: HZL check with Matt Brule for the tester BOM and pricing / CEA 11Sep19: Finance AT mentioned that able to use existing CEA 10Oct19: HZL to check and advise the progress from Milford 13Nov19: HZL to check for update 04Dec19: Pending for HZL comfirmation 11Dec19: On-track as per HZL feedback from Milford | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
4 | OESR # 190059 Alliance Final Test 3 Frequently Freeze During Final Test OESR # 193244 Milford OESR - 03 Apr 2019 04 Apr 2019 - Matt Brule request Flex to provide detail on Hang issue 05 Apr 2019 - Joel provide detail to Matt Brule for detail of Hang issue 04 Sep 2019 - Joel to monitor and update by 06 Sep 2019 06 Sep 2019 - JMD sent monitoring data to HZL (refer rto attach Alliance FT3 System Freeze Tracking Record) 09 Sep 2019 : HZL updated Matt Brule with attached Alliance FT3 System Freeze Tracking Record 18 Sep 2019: HZL last occurrence is 3 Sep 19, slow build for Alliance follows thus need to extend monitoring / with respective volume (min 100 units), HZL to inform Matt that Flex to further monitor this issue. 25 Sep 2019: JMD issue the tracker for 100 units to Prod'n and monitoring in progress | JMD HZL | 2019-09-04 | 13 Sep 19 27 Sep 19 30 Oct 19 | Monitoring | High
5 | OESR # 190079 TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. 15May2019 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04Sep2019 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11Sep2019 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18Sep2019 : JMD will inform YHC once EQID 0497 is available 19Sep2019: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25Sep2019 VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. | JMD HZL | 2019-09-04 | 2019-09-27 | Close | High
6 | OESR # 190102 - TE82988-2 Power Supply 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
7 | OESR 190104 - TE82988-2 Spare Parts holding at Senai - to request return back to WPPL 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10 Sep 2019 - Erzad update ETA back to WPPL by 23 Sep 2019 | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
8 | OESR # 190108 - TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep 2019 FLX Kallang JMD to follow-up to Edwin 05 Sep 2019 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
9 | OESR # S190109 - TSP having 735002531 with 735001613 does not have consistent content 04 Sep19: reviewed with HZL/EK/DA and JMD and agreed to update the OESR content and its Model PN. 05Sep19: JMD updated the OESR 190109 and send to HZL 18Sep19: HZL to checked for update 24Sep19: HZL to follow-up with E Turner 27Sep19: HZL to follow-up with E Turner reminder due as Waters Annual Audit for Flex 08Oct19 : HZL send reminder to E.T 08Oct19 - HZL ask ET if need to revise procedure/monitoring or uprev the next revision. 13Nov19: HZL to check for update | JMD HZL | 2019-09-04 | 06-Sep-19 10-Oct-19 30-Dec-19 | Monitoring | Medium
10 | OESR # 190117 - TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05 Sep 2019 Start 20 Aug 2019 used in 210000411 SN 010919110M 11 Sep 2019 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
11 | OESR #190118/ #190119 Reticule Plate Broken Target Mirror 19Aug19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30Aug19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04Sep19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23Sep19: SFC relelase PO to vendor 10Oct19 PM documentation approval in-progress 16Oct19: Part lead time is 4week from now (16Nov19) 20Oct19: GLY assembled and release to Production. | JMD SFC | 2019-09-04 | 2019-12-30 | Close | Medium
12 | OESR #190124 - 193544 Request for TE# for ICT Fixture 210001344 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05 Sep 2019 Erzad to send the documentation to JMD 06 Sep 2019 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11 Sep 2019 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
13 | OESR S190130 -193550 TE82882 PDA Final Test System Missing Spare OB PN 700011652 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18Sep2019 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25Sep2019 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26Sep2019 - TE82882-10SEP19-VAF_Approved and released tester for usage. | JMD HZL | 2019-09-04 | 06 Sep 19 30 Sep 19 | Close | High
14 | OESR # S190132 - TE82943-13 USB Device Error 03Sep19 JMD raised OESR for TE82943-13 USB Device Error 04Sep19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06Sep19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10Sep19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11Sep19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18Sep19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25Sep19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09Oct19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 10Oct19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10Oct19: Informed purchaser to release pay using Paypal. 16Oct19: JMD to follow-up E.T.A from buyer 23Oct19: Bliss target ETA to Flex is on 29Oct19 30Oct19: To install today (Shamsul) and update status 06Nov19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | JMD HZL | 2019-09-04 | 06 Sep 19 18 Sep 19 25 Sep 19 30 Oct 19 | Close | High
15 | OESR # 190133 Label Printing 04 Sep 2019 GLY to follow-up with DA | GLY / DA | 2019-09-04 | 2019-09-06 | Close | High
16 | OESR # S190156 TE82503-5 Injector Seal Pack Ramp and Decay Test Fail 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | JMD HZL | 2019-10-03 | 2019-10-21 | Close | High
17 | OESR # S190194 TE82943 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium
18 | OESR # S190195 TE82882 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium

## OESR_v01
OESR
Date Requested | OESR No # | Description | Model | Initiated By | Assigned To | Action/Remark | Status
2019-08-14 | 193460 | Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) | WAT057200, WAT057463, WAT080733, WAT080734, WAT080736, WAT081140, WAT081157, WAT081158, WAT081159, WAT081321, 289007906, 289010294, 289007898, 289007899, 289010306, 289007905, 289010304, 289003590, 289010277, 289004509, 289010022, 289000512, 289012264 | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | Open
2019-08-14 | 193461 | Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) | WAT271290, WAT270928 | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 22-Jan-2020 Milford update to receive by Q2 | Open
2019-08-14 | 193124 | Additional Capacity for Alliance Pump Tray Tester (TE82526) | 279000419, 279000524, 289000526, 700000713, 700001030, WAT270463, WAT271078 | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08-Jan-2020 Awaiting TRF from Mark Jussaume 15-Jan-2020 Shipment already out and awaiting for arrival for tester upgrade.(TRFMJUEBJYL4US) 22-Jan-2020 Arrived to Flex and update WPPL. Must open package with presence of WPPL OPS Team.HZL to check with M Jussaume for the approve VAF and protocol for perfroming validation. | Open
2019-09-04 | S190059 193244 | Alliance Final Test 3 Frequently Freeze During Final Test | 186269505, 186269506 | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | Monitoring (OFI)
2019-05-15 | S190079 | TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. | 186015001, 186015018,186015080 | JMD | HZL | 15-May-19 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04-Sep-19 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11-Sep-19 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18-Sep-19 : JMD will inform YHC once EQID 0497 is available 19-Sep-19: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25-Sep-19: VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. 13-Nov-19: Unit has been scrap and emailed Certificate of Destruction and sign-off scrap form to Flex and WPP OPS Team | Close 13-Nov-19
2019-09-04 | S190102 | TE82988-2 Power Supply | 210000415 | Erzad | HZL | 04-Sep-2019 FLX Kallang JMD to follow-up in Senai 05-Sep-2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06-Sep-2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | Close 06-Sep-19
2019-09-04 | S190104 | TE82988-2 Spare Parts holding at Senai - to request return back to WPPL | 210000415 | Erzad | HZL | 04-Sep-19 FLX Kallang JMD to follow-up in Senai 05-Sep-19 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06-Sep-19 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10-Sep-19 - Erzad update ETA back to WPPL by 23 Sep 2019 | Close 23-Sep-19
2019-09-04 | S190108 | TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts | 210000415 | Erzad | HZL | Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep-19 FLX Kallang JMD to follow-up to Edwin 05-Sep-19 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | Close 05-Sep19
2019-09-04 | S190109 | TSP having 735002531 with 735001613 does not have consistent content | 186015034, 186015053, 700009014, 700009532, 700011868, 700011869, | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | Monitoring (OFI)
2019-09-04 | S190117 | TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS | 21000411 | Erzad | HZL | Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04-Sep-19 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05-Sep-19 Start 20 Aug 2019 used in 210000411 SN 010919110M 11-Sep-19 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | Close 11-Sep-19
2019-09-04 | S190118 S190119 | Reticule Plate Broken Target Mirror | 279003224, 2791922 | JMD | HZL | 19-Aug-19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30-Aug-19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04-Sep-19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23-Sep-19: SFC relelase PO to vendor 10-Oct-19 PM documentation approval in-progress 16-Oct-19: Part lead time is 4week from now (16Nov19) 20-Oct-19: GLY assembled and release to Production | Close 20-Oct-19
2019-09-04 | S90124 193544 | Request for TE# for ICT Fixture 210001344 | 210000562, 210001164, 210001344 | Erzad | HZL | 04-Sep-19 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05-Sep-19 Erzad to send the documentation to JMD 06-Sep-19 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11-Sep-19 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | Close 11-Sep-19
2019-09-04 | S190130 | TE82882 PDA Final Test System Missing Spare OB PN 700011652 | 700011652 | JMD | HZL | 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18-Sep-19 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25-Sep-19 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26-Sep-19 - TE82882-10SEP19-VAF_Approved and released tester for usage. | Close 26-Sep-19
2019-09-04 | S190132 | TE82943-13 USB Device Error | 186015001, 186015018, 186015080 | JMD | HZL | 03-Sep-19 JMD raised OESR for TE82943-13 USB Device Error 04-Sep-19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06-Sep-19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10-Sep-19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11-Sep-19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18-Sep-19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25-Sep-19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09-Oct-19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 1-0Oct-19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10-Oct-19: Informed purchaser to release pay using Paypal. 16-Oct-19: JMD to follow-up E.T.A from buyer 23-Oct-19: Bliss target ETA to Flex is on 29Oct19 30-Oct-19: To install today (Shamsul) and update status 06Nov-19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | Close 06-Nov-19
2019-09-04 | S190133 | Label Printing | Generic | GLY | DA | 04-Sep-19 GLY to follow-up with DA | Close 04-Sep-19
2019-10-03 | S190156 | TE82503-5 Injector Seal Pack Ramp and Decay Test Fail | 279000925, 289001994, WAT270576, 279003915, 289012412 | JMD | HZL | 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | Close 21-Oct-19
2019-12-17 | S190194 | TE82943 Truncated (test time) on the Test Report | 186015001, 186015018, 18605080 | JMD | HZL | 17-Dec-19: Concall between JMD and HZL to inform Ed Jackson if OK 20-Dec-19: Waters Milford Matt Brule and Eric Turner aware and will update one time is available | Close 20-Dec-19
2019-12-17 | S190195 193814 | TE82882 Truncated (test time) on the Test Report | 186299800, 186002998, 186015032, 186015033, 186015034, 186015053 | JMD | HZL | 17-Dec-19: Concall between JMD and HZL to inform E Turner if OK 20-Dec-19: Waters Milford Matt Brule and Eric Turner aware and will update one time is available | Close 20-Dec-19
2019-12-18 | S190198 | TE82526-4 Pump Tray Not Able to Flow Solvent | 279000419, 279000524, 289000526, 700000713, 700001030, WAT270463, WAT271078 | JMD | HZL | 18-Dec-19: Raise TE82526-4 Pump Tray Not Able to Flow Solvent 20Dec19: Hua Zhen send the flow valve to Flex, replaced the vave and verified station OK | Close 20-Dec-19
2019-12-19 | S190199 | Instrument Serial Number print out with barcode with broken , alignment on logo and missing Waters logo) Affected label :- | 186269505, 289000560 | GLY | GYC | 19-Dec-2019 - Reprint the label 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08-Jan-2020 JMD received cuvettes from HZL | Close 19-Dec-19
2019-12-24 | S190201 | Replacement for Consumable (Cuvettes) | PDA and TUV Family Detector | JMD | HZL | 02-Jan-2020 REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | Close 08-Jan-20
2020-01-10 | S20003 | TE82526-1 Pump Tray Not Able to Flow Solvent | 279000419, 279000524, 289000526, 700000713, 700001030, WAT270463, WAT271078 | JMD | HZL | 12-Jan-2020: Raise TE82526-4 Pump Tray Not Able to Flow Solvent to Seal Wash Input 15-Jan-2020: Received the BSBA Shut-Off Valve | Close 15-Jan-2020
2020-01-15 | S20005 | Alliance Instrument Serial Number A20SM4286A print out with barcode with broken line) | 186269506 ASSY, e2695 WITH HEATER/COOLER. | GLY | HZL | 15-Jan-2020: Reprint Serial number A20SM4286A | Close 15-Jan-2020
2020-01-15 | TBD | QSP Instrument Serial Number G19QSP507A A print out with barcode with broken line) | 186015080 ACQUITY UPLC H-Class PLUS QSM S/N: G19QSP507A | GLY | HZL | 15-Jan-2020: Reprint Serial number S/N: G19QSP507A | Close 15-Jan-2020

## TMD
Last updated date: 22 Jan 2020
Last updated by: Joel MD
S/N | Deviation# | Version | Flex DE | Description | Affected P/N | Product Type | Originator | Effective Date | Expiration Date/ Criteria | Status | Has this been extended? | Column2 | Remark
1 | TMD010626 | v00 | DE-SKW-00000101 | ...allow the Alliance Instrument Build to use untested PCBA 210000411 from Flex Senai until the spare test fixture arrived and released to production line. | 210000411, 186269505, 186269506 | LC | Chan Yee Hwa | 2019-01-21 | 2019-02-28 | closed | N/A |  | na
2 | na (internal DE, with MPPL signed MEMO | v00 | DE-SKW-00000108 | ferrule 1125 | [5 empty] | closed | N/A |  | na
3 | TMD010644 | v00 | DE-SKE-00000109 | .. ..allow for P/N 410001125 to be dispositioned as "Use As Is" and consumed in the production of instruments, kits, and spares.... | 410001125 | LC | Ethan Peloquin | 2019-02-27 |  | closed
4 | TMD010640 | v00 | DE-SKW-00000110 | .. ..allow for P/N WAT022330 to be dispositioned as "Use As Is" and consumed in the production of instruments, kits, and spares... | WAT022330 | LC | Ethan Peloquin | 2019-02-22 | 2019-05-31 | closed | N/A |  | na
5 | TMD010639 | v00 | DE-SKW-00000114 | ..seal pack, to allow production to : i. Apply methanol on the washer. ii. Use of new holding jig iii. Use of alignment pin | 289012412 | LC | Chan Yee Hwa | 2019-02-25 | 2019-05-10 | closed | N/A |  | na
6 | NA | NA | DE-SKW-00000116 | To allow Flex use two PNs supplied by Waters Milford at older revision, old year code, and made by Waters US supplier, Atronix | 441000619: 42pcs WAT081179 : 45pcs | NA | NA | NA |  | closed |  |  | Hua Jie email dated 30 Mar: WATH-441000619-C & WATH-WAT081179-J (CMTi) - condition of Milford shipped inventory - DE # 116
7 | TMD010635 | v00 | DE-SKW-00000117 | ...to allows the use of flow cell covers with the newly designed boss feature | 415002122 (MRPC), WAT081201 (MRPC), 289012264 (Milford) | LC | Matt Franklin | 2019-02-13 | 2019-09-30 | closed | N/A |  | na
8 | TMD010527 | v00 | DE-SKW-00000120 | ...to evaluate on a new jig to aid the process of snapping the CPU PCB (PN:210000520) onto the Personality PCB... | various Detectors.... See TMD for the full list | LC | Chan Yee Hwa | 2019-03-28 | 10pcs each combination | closed | N/A
9 | TMD010661 | v00 | N/A as need to uprev | ...to allow for production testing using released test software versions and qualified (but unreleased) test software versions, until all stations have been upgraded. | 186015001 186015018 186015080 | LC | Emma Pierce | 2019-05-28 | 2019-08-30 | closed | Yes
10 |  | v01 | DE-SKW-00000140 | ...to allow for production testing using released test software versions and qualified (but unreleased) test software versions, until all stations have been upgraded. | 186015001 186015018 186015080 | LC | Emma Pierce | 2019-05-31 | 2019-08-30 | closed | Yes
11 |  | v02 | DE-SKW-00000157 | ...to allow for production testing using released test software versions and qualified (but unreleased) test software versions, until all stations have been upgraded. | 186015001 186015018 186015080 | LC | Emma Pierce | 2019-08-26 | 2019-09-27 | closed | N/A |  | 02 Sep 2019 - TE82943 Acquity SDS Final Tester has been updated
12 | TMD010654 | v00 | DE-SKW-00000184 | …PN WAT079185 upon reciept to Keypoint is to be routed to Richport for passivation per WA022306 | WAT079185 | LC | Monica Collins | 2019-10-25 | 2020-03-30 | Implementation | N/A |  | 15-Jan-2020 In-prgoress 22-Jan-2020 Engineering activity completed, Implementation Stage
13 | TMD010667 | v00 | DE-SKW-00000137 | .. to allow Flex convert 71 units of 186015018 ACQUITY UPLC QSM into 186015080 ACQUITY UPLC H-Class PLUS QSM. | 186015018 | LC | Yin Yin Soo | 2019-05-16 | 71 uints ony | closed | N/A
14 | TMD010367 | v06 | DE-SKW-00000145 | ...to the opening in the cover is always the feature to crack (see attachment B for location). The 3 other features seem to accept the screw without cracking. - Both the hole feature and the screw measure within specification. This TMD is needed to use a smaller screw in this one feature while root cause is determined. | 415002619; 415002125 415002122 | LC | Matt Franklin | 2018-06-19 | 2019-10-31 | closed | N/A
15 | TMD010578 | v02 | DE-SKW-00000141 | ...will allow the receipt of these parts until such time as the labeling can be added to the packaging. This TMD is being extended in order to give the supply chain adequate time to implement the labeling changes. | WAT212043 RETAINER, FRIT is used in: 289001994 ASSY, SEAL PACK, ALLIANCE 289012412 ASSY, SEAL PACK, ALLIANCE WAT270942 2690 SEAL PACK REPLACEMENT KIT | LC | Matt Franklin | 2019-05-31 | 2020-03-15 | Implementation | N/A |  | 15-Jan-2020 In-prgoress 22-Jan-2020 Engineering activity completed, Implementation Stage
16 | TMD010508 | v02 | DE-SKW-00000147 | ...will allow the receipt of these parts until such time as the labeling can be added to the packaging. This TMD is being extended in order to give the supply chain adequate time to implement the labeling changes. | More parts added | LC | Joe Gately | 2019-06-26 | 2020-06-30 | Implementation | N/A |  | 15-Jan-20 20 In-prgoress 22-Jan-2020 Engineering activity completed, Implementation Stage
17 | TMD010709 | v00 | DE-SKW-00000165 | ...will update the print and ASP and the two drip trays will get a new number assigned. | 289003609 , 289005242 | LC | Roland Poisson | 2019-09-24 | 2019-12-30 | closed | N/A |  | Extended - refer to DE-SKW-00000195
18 | TMD010709 | v01 | DE-SKW-00000195 | The assemblies print and assembly procedures (ASP) below need to be modified. An ECR will update the print and ASP and the two drip trays will get a new number assigned. | 289003609 , 289005242 | LC | Roland Poisson | 2019-12-11 | 2020-03-30 | Implementation | N/A |  | 15-Jan-2020 In-prgoress
19 | TMD10734 | v00 | DE-SKW-00000187 | ..Due to components obsolescence the power supply had to be revised to accommodate the new components. The new design has been tested by the manufacturer and verified by Waters. As a last step before releasing the new change to Manufacturing, a pilot run is being requested to build 20 instruments using the new revised lamp power supply (REV G) in some TUV and PDA detectors. | 279001105 | LC | Ziad El-Massih | 2019-11-14 | 20 pices | Closed | N/A |  | 15-Jan2020 -JMD to verify closuer date 1) Serial number for the instruments installed with 5 of the new lamp power supplies (LPS) P/N 279001105 in TUV detector model P/N 186248900 are : (a) L1987E928A; (b) L1987E929A; (c) L1987E930A; (d) L1987E931A & (e) L1987E932A. 2) Serial number for the instruments installed with 5 of the new lamp power supplies (LPS) P/N 279001105 in TUV detector model P/N 186015031 are : (a) L19TUV330A; (b) L19TUV331A; (c) L19TUV332A; (d) L19TUV333A; & (e) L19TUV333A. 3) Serial number for the instruments installed with 5 of the new lamp power supplies (LPS) P/N 279001105 in PDA detector model P/N 186015033 are : (a) L19UPL331A; (b) L19UPL332A; (c) L19UPL333A; (d) L19UPL334A; (e) L19UPL335A. 4) Serial number for the instruments installed with 5 of the new lamp power supplies (LPS) P/N 279001105 in PDA detector model P/N 186299800 are : (a) L19998288A; (b) L19998289A: (c) L19998290A: (d) L19998291A & (e) L19998292A. All the above instruments installed with the new lamp power supplies (LPS) P/N 279001105 passed all Manufacturing testing.
20 | TMD010753 | v00 | DE-SKW-00000192 | ...to convert the Non-Plus to the Plus models. | 186015018 186015080 | LC | Chan Guan Yee | 2019-11-28 | 15 Units only | Closed | N/A |  | 15-Jan-2020 In-prgoress 22-Jan-2020 Eng'g activity close
21 | TMD010761 | v00 | DE-SKW-00000197 DE-SKW-00000204 | ...to allow the use of the new pressure transducer (279010172 TRANSDUCER, PRESS, 5.5KSI, Z, ALLIANCE) in186269506 e2695 Separations Module w/SHC (Blue) prior to implementation of ECO 9232680 | 186269506 | LC | Gary Bertone | 2019-12-11 | 50 units only | Closed | N/A |  | 15-Jan-2020 in progress 21 Jan 2020 - Complete by-off, no reject observed related to transduceer
22 | TMD010739 | v00 | DE-SKW-00000203 | ...to replace TUV Personality PCB P/N 210000355. On that circuit board assembly a programmable chip (FPGA) from ALTERA (INTEL) had gone obsolete. The new PCB utilizes a newer technology FPGA by XILINX. To accommodate the change many circuits on the board had to be redesigned. | 210000355, 210001344 | LC | Ziad El-Massih | 2019-12-16 | 210001344(30), 210001232(30) | Implementation | N/A |  | 15-Jan-2020 in progress. WAT-186248900 - built WAT-186002487 – going to build after 15-Jan shipment WAT-186017002 – going to build after 15-Jan shipment
23 | N/A (WPPL sign FlexDE) | NA | DE-SKW-00000177 | Lean Project | a) 186015031 b) 186017002 c) 186017012 d) 186248900 | LC | Amy Xu Xia Jie |  | 4 units | Closed | N/A |  | 15-Jan-2020 Drop Test Concluded last Dec 2019 Affected product serial number for the above DE are : a) 186015031 AQUITY UPLC TUV Detector TC : K19TUV195A. b) 186017002 ACQUITY Arc 2489 UV_Vis Detector : K19VTU072A. c) 186017012 ACQUITY Arc Bio 2489 UV_Vis Detector : J19BTU090A. d) 186248900 2489 UV / Visible Detector ( BLUE ) : K1987E431A.
24 | TMD010756 | v00 | DE-SKW-00000202 | ...will pilot a new supplier's grating drive motor in the following instruments and quantities: | 2475 and 2489 Grating Drive Motors 2489 Detectors: 279003346 - Qty. 15 2475 Detectors: 279003347 - Qty. 15 | LC | Nicholas Szwaja | 2019-12-23 | 15 2475 / 15 2489 Detectors | Closed | N/A |  | 15-Jan-2020 in progress 22-Jan-2020 Buy-off completed by Eugene WPPL SQE on 20 Jan 2020- No issue found
25 | TMD010766 | v00 | DE-SKW-00000208 | ...will allow for a blank label (pn 615005253) to be placed over the existing silk screened EC Rep on non IVD instrument Chassis until ECO 9229664 is effective. | 186001515, 186001525, 186002487, 186002998, 186015001, 186015018, 186015031, 186015033, 186015034, 186015053, 186017002, 186017003, 186017013, 186248900, 186299800, 186269505, 186269506, 187269507, 186269508 | LC | Joe Gately | 2019-12-20 | 2020-07-15 | Active | N/A |  | 15 Jan 2020 Gary to drafted DE-SKW-00000208 workflow 22 Jan 2020 Pend drawing for new blank label and to include WAT20700
26 | TMD010770 | v00 | DE-SKW-00000207 | This TMD allows Flex to build 10 Optics Benches 279003224 without the foam insulation. Ship the non-installed foam insulation with optics bench. | 279003224 | LC | Matt Franklin | 2020-01-09 | 10 units only | Implementation | N/A |  | 15-Jan-2020 QDMS DE in aprroval loop 22-Jan-2020 : Approved and move to Implementation stage

## Pump Head (HotPart) (hidden)
Update as if 04 Dec 2019
Below data per latest sound measurement conducted today.
Present : Aldwin. Ivy, Desmond, Guan Yee, Eugene
As per discussion today the data gathered today shall be used as reference final reference for the Sound comparison of Hot Part vs Known Good Part..
Observations:
The sound level of the Hot Part Motor (68.2 dB) is still high when installed to production unit (Hot Part Motor + New#6)
There is an average of 5.0 dB difference between the Hot Part compared to 6 samples of production units.
After re-sitting ang tightening the pump coupling shaft (WAT271467) of the Hot Part Motor, the sound level is still higher.
SN | Date | Comparison | General Sound Level ( A ) dB (min) | General Sound Level ( A ) dB (max)
1 | 2019-12-04 | Hot Part (1st try) | 66.4 | 68.4
2 | 2019-12-04 | New#1-1 | 62 | 64.6
3 | 2019-12-04 | New#2-1 | 62.6 | 65.2
4 | 2019-12-04 | New#3-1 | 61.4 | 63.1
5 | 2019-12-04 | Hot Part (2nd try) | 68.4 | 70
6 | 2019-12-04 | Hot Part (3rd try) | 68.2 | 69.5
7 | 2019-12-04 | New#4-1 | 65.1 | 67.6
8 | 2019-12-04 | New#5-1 | 61.8 | 62.5
9 | 2019-12-04 | New#5-2 | 61.7 | 62.3
10 | 2019-12-04 | New#6-1 | 64 | 64.7
11 | 2019-12-04 | Hot Part (4th try) | 67.4 | 70.8
12 | 2019-12-04 | New 6 Motor + Hot Part Body (1st try) | 62.8 | 64.5
13 | 2019-12-04 | New#6-2 | 66.8 | 67.7
14 | 2019-12-04 | Hot Part Motor + Hot Part Body (after Tightened) | 69.3 | 70
15 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 1) | 68.1 | 69.4
16 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 2) | 67.9 | 69.5
17 | 2019-12-04 | Hot Part Motor + New#6 | 66.8 | 68.2
18 | 2019-12-04 | New#6 Motor + Hot Part Body | 63.7 | 64.2
2019-12-09 | Final report submitted to WPPL QA Team. Pending final dispotion. Eng'g Team activity Close

## Alliance FT3 Tracker (hidden)
Alliance Final Tester 3 Freeze OESR S190059/OESR 193244 Alliance Final Test 3 Frequently Freeze During Final Test
Item | Serial Number | Freeze (Y/N) | Time | Prepared By | Remarks
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50

## IQR TE82943 UPDATE (hidden)
Used Known Good Unit QSM+ (F19QSP275A) and BSM (F19UPB765A)
Week 0f (estimate) | System Number | 2019-06-03 | 2019-06-03 | 2019-06-10 | 2019-06-17 | 2019-06-24 | 2019-07-01 | 2019-07-08 | 2019-07-15 | 2019-07-22 | 2019-07-29 | 2019-08-05 | 2019-08-12 | 2019-08-19 | 2019-08-26 | 2019-09-02 | 2019-09-09 | 2019-09-16 | 2019-09-23
WPPL Flextronics | TE82943-11 | 2019-06-03: TMD 010161 v01; 2019-06-03: IQ
| ACTUAL | 2019-06-03: TEST SOFTWARE UPDATE 07 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-10: 10 Jun 2019 186018080 F19QSP275A 14 Jun 2019 186015001 F19UPB765A; 2019-06-17: PREPARED BY: CYH 19 Jun 2019; 2019-06-24: REVIEWED BY: EM 19 Jun 2019 EJ 20 Jun 2019 APROVED BY: TKC 28 Jun 2019
| TE82943-12 | 2019-06-03: PLAN; 2019-06-10: IQ
| ACTUAL | 2019-06-17: TEST SOFTWARE UPDATE 25 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-24: 26 Jun 2019 186018080 F19QSP275A 26 Jun 2019 186015001 F19UPB765A; 2019-07-08: PREPARED BY: CYH 09 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-13 | 2019-06-03: PLAN; 2019-06-17: IQ
| ACTUAL | 2019-07-01: TEST SOFTWARE UPDATE 09 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-07-08: 10 Jul 2019 186018080 F19QSP275A 11 Jul 2019 186015001 F19UPB765A PREPARED BY: CYH 12 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-14 | 2019-06-03: PLAN; 2019-06-24: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 18 Jul 2019 186018080 F19QSP275A 19 Jul 2019 186015001 F19UPB765A; 2019-07-22: PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-15 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software 18 Jul 2019 186015001 F19UPB765A; 2019-07-22: 22 Jul 2019 186018080 F19QSP275A PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-16 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 07 Aug 2019 186015080 F19QSP275A 08 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-17 | 2019-06-03: PLAN; 2019-07-08: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 08 Aug 2019 186015080 F19QSP275A 07 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-18 | 2019-06-03: PLAN; 2019-07-15: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-08-12: 14 Aug 2019 - testing completed; 2019-08-19: 21 Aug 2019 Resubmit : EQID 0497 Test report with updated calibration report and Evauation Result - Not allowed to used due to somer error on 90g reading during calibration; 2019-08-26: TMD010661 v 01 To be extended until end Sept 19 29 Aug 2019 Submitted TE89243-18 test report to YHC using AND HR 200 SN 12337042 (EQID 2092) 30 Aug 2019 WPPL ISOISO13485 Audit; 2019-09-02: WPPL ISO13485 Audit 30 Aug 2019-02 Sep 2019 TMD010661 v 02 Start at Flex 02 Sep 2019 extended until end Sept 19 AND HR 200 ITS verification completed - within specs. 02 Sep 2019 AND HR 200 Miing Deng re-calibration within specs 05 Sep 2019; 2019-09-09: KC Tan on business trip until 18 Sep 19; 2019-09-16: KC Tan sign-off 18 Sep 19

## Consined Alliance Training Plan (hidden)
Known Good Alliance Transfer Plan To Flex Senai
Mounting and dismantling of PCBA to UUT - 2695, Alarm Message (Related to Allianc 2695 only) - Instrument Specification (Basic related to PCBA) | WW 40 | WW 41 | WW 42 | WW 43 | WW 44 | WW 45 | REMARKS
Plan
Actual
Generic Preventive Maintenance Procedure For Test Equipment (735004259)
Plan
Actual
Transfer of Alliance Known Good Unit to Senai
Plan
Actual
Flex Senai to use Known Good Unit for Troubleshooting
Plan
Actual
Flex Senai to source Spare Part such as cables to enhance toubleshooting
Plan
Actual

## TUV Shunt Energy Level (hidden)
TUV Flow Cell | 55% to 150 %
Energy Setting as per ASP
Shunt ID

## Pump Spares Part (hidden)
Waters ID | Description | Part Number | Spare Part Description | OESR | SOURCE BY: |  | Remarks
[5 empty] | FLEX (Yes/No) | WPPL (Yes/No)
TE82705-2 | 1500 Series Final Test Station | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
|  |  | Waters Pump Control Module |  |  | Yes
|  | SC2072 ASSY181310-02 | National Instrument Controller Card |  |  | Yes
TE82534 | 515 Final Tester |  |  |  |  | Yes
|  | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes | No
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes | No
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
TE82613-2 | Pump Mechanic Run In Module |  | TE82613 PUMP MECHANIC RUN IN MODULE |  |  | Yes
TE82851 | 515 Bezel Assemby Tester |  | TE82851 , 515 BEZEL ASSEMBLY TESTER Both Waters customized PCBA PN is not available |  |  | Yes

## Alliance Pump HDD Imaging (hidden)
Alliance / pump |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.50 | TE 82534 - 2 | 515 Pump Final Test | 2019-02-12 | Completed
10.249.26.51 | TE 82705 - 2 | 1500 pump Final Test | 2019-03-12 | Completed
10.249.26.20 | Final Test - 1 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.21 | Final Test - 2 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.22 | Final Test - 3 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.23 | Final Test - 4 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.24 | Final Test - 5 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.25 | Final Test - 6 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.26 | TE 82526 - 1 | Waters Alliance Test Station ( Pumptray Test - 1 ) | 2019-03-13 | Completed
10.249.26.27 | TE 82526 - 2 | Waters Alliance Test Station ( Pumptray Test - 2 ) | 2019-03-13 | Completed
10.249.26.28 | TE 82526 - 3 | Waters Alliance Test Station ( Pumptray Test - 3 ) | 2019-03-13 | Completed
10.249.26.29 | TE 82519 | Waters Alliance Test Station ( Pump Head Mechanics ) | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.32 | TE 82573 - 1 | E-Alliance Rear Panel Test | 2019-03-12 | Completed
10.249.26.34 | AutoLoader - 2 | Alliance AutoLoader Station 2 | 2019-03-12 | Completed
10.249.26.33 | AutoLoader - 3 | Alliance AutoLoader Station 3 | 2019-03-12 | Completed
10.249.26.35 | TE 82520 - 1 | Alliance Carousel/Carrier Alignment Station -1 | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.36 | TE 82520 - 2 | Alliance Carousel/Carrier Alignment Station -2 | 2019-03-13 | Backup thru Windows OS, not able to use Acronis Imaging Software
10.249.26.37 | TE 82546 | Front Panel Station | 2019-03-13 | Completed
10.249.26.38 | TE 0084018-1 | Heater Cooler Tester | 2019-03-28
10.249.26.90 | QA Workstation | QA Workstation/SQT.ORT | 2019-02-11 | Completed
NA | PMS | Particle Measurement |  | Plan : 30 Apr 19
PLAN : Completion Date

## BSM.QSM HDD Imaging (hidden)
BSM /QSM |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.2 | Waters SISNL1 | Solectron Instrument Serial Number 1 | 2019-03-12 | Completed
10.249.26.3 | Waters SPNL1 | Solectron Part Number Labeling 1 | 2019-03-12 | Completed
10.249.26.4 | WATERSCOFP1 | WATERS Certificate Of Funtional Performance 1 | 2019-03-12 | Completed
10.249.26.9 | TE 82800 - 2 | Acquity Actuator Tester #2 | 2019-02-13 | Completed
10.249.26.10 | TE 82800 - 3 | Acquity Actuator Tester #3 | 2019-03-12 | Completed
10.249.26.11 | TE 82943 - 11 | Acquity SDS Tester #11 | 2019-02-13 | Completed
10.249.26.12 | TE 82943 - 12 | Acquity SDS Tester #12 | 2019-03-12 | Completed
10.249.26.13 | TE 82943 - 13 | Acquity SDS Tester #13 | 2019-02-13 | Completed
10.249.26.14 | TE 82943 - 14 | Acquity SDS Tester #14 | 2019-02-13 | Completed
10.249.26.15 | TE 82943 - 15 | Acquity SDS Tester #15 | 2019-02-13 | Completed
10.249.26.16 | TE 82943 - 16 | Acquity SDS Tester #16 | 2019-02-13 | Completed
10.249.26.17 | TE 82943 - 17 | Acquity SDS Tester #17 | 2019-03-12 | Completed
10.249.26.18 | TE 82943 - 18 | Acquity SDS Tester #18 | 2019-03-13 | Completed
10.249.26.55 | Genealogy SDS | Acquity SDS Genealogy Stn (BSM/QSM) | 2019-03-12 | Completed
PLAN : Completed

## Detector HDD Imaging (hidden)
Detectors / Cleanroom |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.40 | TE 82903 - 6 | TUV Final Test Station | N/A | White Lace \ CD ROM Not able to activate Acronis Imaging Software
10.249.26.42 | TE 82903 - 8 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.43 | TE 82903 - 9 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.44 | TE 82903 - 10 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.45 | TE 82964 - 1 | TUV Pretest System | 2019-03-13 | Completed
10.249.26.75 | TE 82969 - 1 | PDA Linearity Tester | 2019-03-13 | Completed
10.249.26.76 | TE 83057-4 | PDA Download Station | 2019-03-13 | Completed
10.249.26.78 | TE 82882 - 2 | 2998/PDA Final Test System | 2019-03-13 | Completed
10.249.26.79 | TE 82882 - 8 | 2998/PDA Final Test 8 | 2019-03-13 | Completed
10.249.26.80 | TE 82882 - 9 | 2998/PDA Final Test 9 | 2019-03-13 | Completed
10.249.26.81 | TE 82882 - 10 | 2998/PDA Final Test 10 | 2019-03-13 | Completed
10.249.26.82 | TE 82882 - 11 | 2998/PDA Final Test 11 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.84 | TE 82882 -13 | PDA Final Test System 13 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.85 | TE 82882 -14 | PDA Final Test System 14 | 2019-03-14 | Completed
10.249.26.63 | TE 82918 - 2 | 2998 Fine Alignment Sys | 2019-03-14 | Completed
10.249.26.64 | TE 82868 - 2 | 2998 Aperture Alignment System | 2019-03-14 | Completed
10.249.26.65 | TE 82951 - 2 | 2998 Mask Alignment Sys | 2019-03-14 | Completed
10.249.26.66 | TE 82916 - 2 | Flowcell Alignment Fixture | 2019-03-14 | Completed
10.249.26.68 | TE 82973 - 1 | 2998 Flowcell Noise & Drift System | NA | CD-ROM unable to read
10.249.26.72 | TE 82917 - 1 | 2998 Rough Alignment | 2019-03-14 | Completed
10.249.26.73 | TE 82880 - 2 | PDA Rough Alignment Sys | 2019-03-14 | Completed
10.249.26.74 | TE 82881 - 2 | PDA Fine Alignment Stn | 2019-03-14 | Completed
10.249.26.56 | Genealogy SDS | Acquity SDS Genealogy Stn (Detector) | 2019-03-15 | Completed: Backup thru Windows OS
10.249.26.59 | TE83096-2 | Universal PDA Alignment | 2019-03-16 | Completed: Backup thru Windows OS
PLAN : Completion Date