## Meeting Minutes
Waters Engineering Weekly Meeting | [7 empty] | Work Week 34
Item | Priority | Action and Updated Status | Owner | Status | Required Completion Date | Actual Completion Date | Priority | Reference document
1 |  | Engineering Evaluation
2.1 | Med | QDA site preparation - Provide weekly states update. [5 Oct] PC server ship date to revise to follow capability kits, need to check with Sidd justification from supplier on Nitrogen reservoir requirement. All items received except divert valve will be received on week 44. WPPL to send out the test equipment tools list. [16 Oct] Add standard tool list statues. [19 Apr] To add/update the required jig in the QDA list. | Alan/Sufern | Open | Ongoing
2.6 | Low | iBSM 1st Pass Yield. [08 Mar] Mimi assign to Detector troublshooting.[22 Mar] Micheal to - Learning and understanding the Assembly and Test process for BSM - Collecting, modifying(standardize) and analyzing data for the first pass yield [19 Apr] Applying 5M (man, machine, method, material & measurement) to study the effectiveness of 3 cycles of priming test and break down the study of seal wash prime failure .Benchmark the priming error and study to propose improvement on testing.From the presentation slide, there are "9x re-zero of weighting scale". Find out the contribution of the testers. | Chan | Open KIV change focus to item 9. | Ongoing
9 | Med | Flow Rate [17May] To understand the variation between the testers at BSM/QSM. [19Jul] To get the off the shelf tool for eval on the tightening of tubing to UUT. To update the fall out rate and data range from Sept to Jul on slide. | Chan | Open | Ongoing
2.7 | Med | BSM seal wear rate at FTN test stations. [23 Feb.] To come out a justification and target on this project improvement. [28 Feb] to come out a draft plan update by next week. 2 weeks to finilize the justification. [19 Apr] To comleted informationof PM for all tester in the able and to add require information in the presentation. To provide pareto chart for FTN tester BSM seal change. [03 May] To prepare data of seal change based on usage of individual tester vs seal change rate. To prepare proposal of swapping the seal between pump A and pump B of selected testers. | Eilwin | Open | Ongoing
2.9 | Low | 1st pass yield monthly compilation to monitor the tester variation. [28 Feb] To start TE82943 for Jan/Feb compilation. [17May] To complie for FTN for May data. [16Aug] To add BSM / QSM data. (6 testers). Detector (2 Tester) [1st week of Sep] CH/CHC (2 tester) CM/CMA (2 Tester). | Kok Leong / Eilwin/Chan/Mimi | Open | Monthly
6 | Med | Asset list for all [03 May] To provide asset list for infrastructure for Qda and other models. [16Aug] To add P.M. number in Assets Manual. | Chong Keong | Open | Ongoing
7 | Med | List of consumable for assembly and testing of all models. [03 May] To provide annually consumable items list for assembly and testing of all models. And P.M. calendar. To check the PM procedure on TUV Detector.[20 Jul] To add chemical list | Hong Cheng | Open | Ongoing
8 | Med | Hipot Safety Enclosure for each Hipot tester [17May] To get a quotation and plan for implementation. [31May] Fixture supplier to provide quote. [18Jun] Quote received, pending issue P.O. to supplier. [16Aug] Implement on production line. Need some modification on Enclosure. | Kok Leong | Open | 15/8/2018 16Sep18
9 | Med | Implementation of Electronic MI | Hong Cheng | Open | Q3, 2018
10 | Med | Migration of Tracebility system to DGS [05Jul] Production line schedule output. | Chong Keong | Open | Q3, 2018
11 | Med | Improvement on BSM packaging. [18Jun] Proposal to send by this week. [28Jun] To add cost benefit to the proposal slide. | Kok Leong | Open | 2018-07-11
12 | Med | Pre Heater Interface issue on Column Manager | Chan / Chong Keong | Open | 2018-07-31
13 | Med | Tubing Evaluation for BSM & QSM test stations | Chan / Chong Keong | Open | 2018-07-31
14 | Med | Needle Seal Force Study for FTNs | Eilwin/Ruby/Hong Cheng | Open | Ongoing
2 |  | ECR/ECO
2.1 | Med | ASP/TSP activities - Focus on column heater/cooler line - Maintain Matrix list - (13 July 2017) - to update ASP/TSP matrix - Done (21 Sep) - Pending WPPL review. - To add Dust Cover into affected TSP. | Chong Keong | Open | Ongoing

## Meeting Minutes (Closed)
Waters Engineering Weekly Meeting |  |  |  | Updated : 25/Jun/2018
Item | Action and Updated Status | Owner | Status | Required Completion Date | Actual Completion Date | Priority
1 | Engineering Evaluation
1.1 | Temperature Recording for FTN testing area - Need discussion with production to build additional ahead in two weeks and start the evaluation after quantity build up. - To compile the temperature reading into slides and submit to WPPL . - Send the data logger temperature reading into the slide as well. | Ei Lwin | Closed | 2017-07-05 | 2017-07-06
1.2 | Fixture Fabrication for CH-A door pin - Lead time required for fixture fabrication is 4 weeks upon PO received.(required by 01 Sep 2017) - Estimated ECO9213028 cut-in Oct 2017 for Fuyu part. [20 Jul 2017] - ECO9215089, GES to send requirement for fixture fabrication. [10 Aug 2017] - drafted a template for standard requirement for tooling fixture fabrication. -HCtan[14 Aug] - PO to be given by WPPL - Desmond[18 Aug] - Drawing to be submitted - HCTan [TDB] - Drawing approve - Desmond [1week] - Fabrication -[4 weeks] - drafted ASP to be submitted with drawing submission - [17-Aug] Desmond to provide supplier contact used by Flex to fabricate the fixture. [24-Aug] - Quote from JEG received and forwarded to WPPL - Done (24 Aug) - Check with Keypoint on their support. Get 2nd quote. - Requirement: 1) draft drawing base on Jig/fixture in existing production line. 2) Jig/fixture cannot be dismantled when taking measurement. 3) Use Solidworks 2017, provide 2D & 3D drawings editable, 2D PDF. 4) Drawings include explored views and individual designed part drawings. [30 Aug] - Compile two quotes and finalise who to fabricate the fixture. [14 SEP] Pending quote from Keypoint. [20 SEP] Received revise quote from Key Point (S) Pte Ltd. [21 SEP] Feedback to Keypoint on the scratch drawing. Pending update from Keypoint. [26 SEP] Received revise scratch drawing and forward to WPPL. [28 SEP] Revisit the design and confirm by next week. [5 Oct] Finalized Quote from keypoint. WPPL to send PO to Keypoint. [11 Oct] WPPL send PO to Keypoint. [27 Oct] Keypoint reply 6-Nov to delivery. [10 Nov] keypoint feedback shipment change to 17-Nov waiting machine shop to deliver the jig.[15 Nov] Delivery improve to 16 Nov.[21 Nov] Mark up ASP send to WPPL. Pending Keypoint to send in Solidworks drawing. [29 Nov] Keypoint feed back need more time to completed the drawing. [07 Dec] Keypoint send in the drawing pending review by GES.[20 Dec] Completed review. No issue found. | Chong Keong Hong Cheng | Open Closed | 2017-09-21 | 2017-12-20
1.3 | QSM Prime D failures study - To prepare problem statement and breakdown the possible root cause by 4M - Draft out the evaluation plan - (13 July 2017) - Mi Mi will share updated slide for prime D failure. - Monitor weekly for prime D failure. | Mi Mi | Closed | 2017-07-13 | 2017-07-24
1.3a | QSM Prime D failures study [20 Jul 2017] - Desmond to share Parker FA report on GPV. - Study and present the solvent line flow. | Mi Mi | Closed | 2017-07-13 | 2017-07-24
1.4a | - Draft out request for evaluation for FTN model and BSM model, Need TMD for this evaluation from Waters. - Red line the ASP to inculde this fixture and send to WPPL. - (13 July 2017) - Provide slide with photos for PCBA fixture. Put in the first page of PCBA matrix - Show how fixture help on improvement on preventing boards from warp [20 Jul 2017] - Tabulate the failure of 210000520 from 2016 till date including failure modes. [10 Aug 2017] - drafted a template for standard requirement for tooling fixture fabrication. -HCtan[14 Aug] - Fabricate a removing tool for 128 pins connector from 210000367 & 210000368 for internal troubleshooting use. | Desmond Hong Cheng | Closed | 7/17/2017 24 Aug 2017 | 2017-07-20
1.4b | - (13 July 2017) - perform one piece first and share with WPPL. - Take 30 units of PCBA assembly photos and video of current process of FTN & come out with the slide to present - Need to note down all the serial numbers for the PCBAs and instruments. - log in the results pass or fail. | Desmond Hong Cheng | Closed | 2017-07-20 | 2017-07-21
1.5 | Door Pin design information across Waters Models - Respective MEs to provide a Matrix(by models) and Photos with categories of Long pins with knurling, Short pins with Knurling, short pins without knurling, etc. | Hong Cheng | Closed | 2017-07-04 | 2017-07-06
1.3a | QSM Prime D failures study [10 Aug 2017] - Summarize the Prime D slides and continue monitoring results to Sep - Mi Mi [17 Aug] - Track the numbers of flushing/purging on check valves. - Mi Mi [17 Aug] - Standardize the purge requirement. - Mi Mi [17 Aug] - Continue to monitor and put into Aug and Sep data. - Provide the fallout rate for the tubing 4300004510 - Send a copy of check valve flushing procedure to Desmond for him to follow up with Milford on official instruction for flushing. - Mark up instruction and send to Desmond. Sent on 28 Aug. [14 SEP] Mi Mi to add detail and send out the copy to WPPL. [04 Oct] Mi Mi send out the final copy with three months data to WPPL. Report sent. | Mi Mi | Closed | 7/13/2017 Sep 2017 25 Aug 2017 18 Sep 2017 | 2017-07-24
1.6 | New Tubing Evaluation for FTN H & I-Class - Submit Request of evaluation to WPPL for review | Ei Lwin | Closed | 2017-06-29 | 2017-06-28
1.7 | Fixture drawing drafting and Fabrication plan for 2018 - Check with R & D if they can support the design and drafting of drawing. - WPPL to provide requirement - drawing version need to be Solidwork 2017 format - Define requirement on drawing requirement - Venture to check with supplier(JEG engineering) with existing capability. JEG reply their Solidworks is 2003 can meet all other requirements. - Solidworks 2017, 2D & 3D Solidworks drawing editable, 2D PDF. Example format to follow Tool: 999027246. Brad Fitting to Tube Assy. (It will be one full overview drawing with BOM included in the drawing, each individual part will have their drawing ). | KLLOO Desmond /Hong Cheng Hong Cheng | Closed | 8/17/2017 24 Aug 2017 21 Sep 2017 26 Sep 2017
1.8 | Column Manager area Shifting. - Option 2. - Send affected TE number to WPPL-Done (21 Sep) - Shifting of TCM/PCM, 5045, W600 completed. Test reports submitted to Siew Yee. - Shifting of Column Manager(5043) on 11 Oct and 5007 on 13 Oct. - Pending IQR approval - IQR Approval received. | Alan/Siew Yee | Closed |  | 2017-10-10
1.9 | Optics Bench Shifting. -Send affected TE number to WPPL-Done (21 Sep) - Shifting of Optic bench alignment stations completed. Test reports submitted to Siew Yee. - Siew Yee to compile IQR for above models (06 Oct 2017) - IQR approval on 10-Oct. | Chong Keong/Siew Yee | Closed |  | 2017-10-10
2.5 | Provide 210000852 Ethernet board failure trend in FTN & Classic Column Manager Jan 2017 till date. - Base on PCBA records, there is no RMA return for 210000852 from Jan’17 till date. | Hong Cheng | Closed |  | 2017-12-13
1.7a | PCBA fixture fabrication [24 Aug] - Quote from JEG received and forwarded to WPPL - Done (24 Aug) - Check with Keypoint on their support. Get 2nd quote. - Requirement: 1) draft drawing base on Jig/fixture in existing production line. 2) Jig/fixture cannot be dismantled when taking measurement. 3) Use Solidworks 2017, provide 2D & 3D drawings editable, 2D PDF. 4) Drawings include explored views and individual designed part drawings. Reply from Venture R&D not able to support to redraw jig/fixture in existing production line. [30 Aug] - Compile two quotes and finalise who to fabricate the fixture. [14 SEP] Pending quote from Keypoint. [20 SEP] Received quote from Keypoint but sketch drawing is incorrect pending update from Key Point. Check affected ASP if there is any tooling mention and update in the matrix.- Done (22 Sep). [26 SEP] Received revise quote (US$4800/-) pending revise sketch drawing. [05 OCT] Keypoint response no bidding for PCBA Press Jig. On hold. [02 Nov] -Invited 2 new supplier to provide quote . 1st supplier: Rumenn Pte Ltd reply not able to support solidework. 2nd supplier: NKS Technologies Pte Ltd pending update. [09 Nov]Supplier: NKS Technologies Pte Ltd submit quote. [16 Nov]Need to re-visit the design. [20 Nov]NKS asking for job order confirmation before arrange design study and look into the requirement.[29 Nov] Pending FFM Services Pte Ltd to submit quote by Work Week 50. [14 Dec] Keypoint send in quote and draft drawing. [18 Dec] FFM Services send in quote and draft drawing.[04 Jan]To finalize the supplier by next week.[08 Feb] Pending WPPL advise. | Hong Cheng | Closed |  | 2018-02-08
2.4 | SIT summary - To provide weekly update. | Hong Cheng | Closed |  | 2018-06-26
4 | Assets manual matrix by family target completion date. [28 Feb] To send Tooling master list and Chong Keong to send time line plan. [07 Mar] To add QDA and Beacon tooling [WK 11] Completed review on BSM.[WK 12] Completed review on QSM. [WK 13] Completed review on FTN.[17May] To update the tracker. | Chong Keong / Su Fern | Closed |  | 2018-07-19
14 | Provide summary of action items required for Detector test issue. | Mi mi | Closed |  | 2018-07-06
2 | ECR/ECO
2.2 | New Firmware release for ECR9215282, ECR9215298 and ECR9215322 - When the new firmware will be ready in the Testers. - When is the TEQ ready for Execution. - Any TMD for the run? No - One unit for each model will be used to verify on one of the tester from each affected model. | WPPL | Open | Closed
2.3 | New Firmware release for ECR9215282, ECR9215298 and ECR9215322 186017000 - Approved and released 186015043 - Approved and released 186015029 - waiting for training record to resubmit. | WPPL | Open | Closed
2 | FTN Test Software Up Version IQ -Immediately retest with new test program after UUT pass on existing test program without disconnect. -Plan to start on 21-Sep-17. -Venture to update qty remaining for Q3 commitment. Information send to WPPL-Done (22 Sep). -IQ completed on 26-Sep-17. | Ei Lwin | Open Closed |  | 2017-09-26
2.7 | Fabricate Pulley Set Block, Dwg No: 999032831 - To get 2 quote. - GES to issue purchase order to fixture supplier: JL Components Trading. [20 Nov] JL Compnents Trading delivery the Pulley Set Block. | Hong Cheng | Open Closed |  | 2017-11-20
1.6a | FTN Tubing evaluation '- To prepare workmanship evaluation(30 units of 5017 & 10 units of 5046 for each operators) prior to New tubing evaluation - New tubing for this workmanship evaluation delivery date target date on 7th July - Venture to share the schedule and plan of the evaluation - (13 July 2017) - First operator's evaluation started on 10-July-2017. Estimated to complete on 17-July-2017 - Complete data of first operator's will be presented on 20 July 2017. [10 Aug 2017] - To start third operator evaluation. Lwin [25 Aug] - Third operator evaluation completed. 1 out of 10 5046 and 1 out of 30 5017. - Compile and provide conclusion of this evaluation. [18 SEP] Eilwin send the revise copy to WPPL. [20 SEP] Reassessment on 2 operators for 1 month. Tubing change will be set as 1 month frequency for monitoring. -Evaluation will be KIV for 1 month. [06 Oct] -Start of assessment on workmanship in 02 October 2017. [12 oct] -Weekly review. | Ei Lwin | Closed |  | 2018-01-04
| TE Technology fault Heater Cooler use in Column Manager | Ei Lwin | Closed |  | 2018-01-04
3 | Engineering Role Matrix. [23-Feb] To add Master trainer in the matrix. [28 Feb] To update by Friday. | Kok Leong | Closed |  | 2018-03-08
2.3 | 186015017 Flex Cable Assembly Evaluation. - Study on the feasibility on using a Press Jig for the Flex Cable. - Prepare a power point slide. [27 Oct] pending new Fixture supplier to submit the quota and scratch drawing. [09 Nov] Supplier: NKS Technologies Pte Ltd submit quote [16 Nov]Need to re-visit the design. [20 Nov]NKS asking for job order confirmation before arrange design study and look into the requirement. [29 Nov] pending FFM Services Pte Ltd to submit quote on work week 50.[18 Dec] FFM service send in quote and draft drawing.[04 Jan] FFM service to revise the draft quote .[23 Feb] Milford going to send in a new jig. [27 Feb] Received the Press Jig. [08 Mar] Try out on 3 instrument no issue. Pending ECO release into production. [13 Mar] ECO 9221253 received pending implemention.Assets Manual and Tooling master list to be updated.[23 Mar] Implement. | Hong Cheng / Desmond | Closed |  | 2018-03-23
2.2 | 186015042 E-cord test failure - Place E-cord before start of testing. Start on work week 38. Failure still occur with the new test method, failure rate is 7/21. - In WK41, number of Ecord failure was 8 which is similar to previous weeks, WK 38 & WK39 [27 Oct] Prepare summary of E-cords failures by months Take the video of two test methods to be sent to Waters. Check connection from UUT to tester for any abnormality. [04 Jan] Received repaired board from CEI [11 Jan] Install board back to tester. [23 Feb] compile a report and raise an OESR for a old board. [08 Mar] Monitor until end of April. OESR will be send after monitoring is effective. [19 Apr] To update failure trending graph for Mar-2018. To close the activity and Desmond (WPPL) will send the file to Waters Test Engineer. Desmond (WPPL) will raise request for old PCBA 210000367 Rev G as a spare of TE83013-1. | Ei Lwin / Desmond | Open Closed |  | 2018-05-03
5 | SIT study on the Sample Manager improvement program. | Hong Cheng / Eilwin | Open Close |  | 2018-05-17
2.2 | Bifrost activities -FTN Plus 186015085 & 186015088 to redline the ASP. | Hong Cheng | Open Closed |  | 2018-05-17
2.3 | Bifrost activities -BSM plus 186015082 to redline the ASP. -Pending official release | Chong Keong | Open Closed |  | 2018-06-28
2.4 | Red lined and Submitted ASPs of Qda - To provide ASP submitted date to WPPL & to add part number/document number in the table | Alan | Open Closed

## OESR
Work Week | New this week | Closed | Outstanding | Detail
21 | 1 | 1 | 0
23 | 1 | 1 | 0
24 | 1 | 1 | 0
25 | 1 | 0 | 1
25 | 1 | 0 | 1
25 | 1 | 0 | 1

## Attendence
| Venture | [6 empty] | Waters
Date | Kok Leong | Hong Cheng | Alan | Chong Keong | Lwin | Mi Mi | Chan | Edwin | Desmond | Huazhen | Remarks
2018-05-17 | Yes | Yes | Yes | Due to urgent requirement to release ECO on BSM pump | Yes | Due to Qda Testing | Yes | No | Yes | No | N.A.
2018-05-24 | Yes | On leave | Yes | Yes | Yes | Yes | Yes | Yes | Yes | No | No engineering meeting, discussion with Edwin.
2018-05-31 | Yes | Yes | Yes | Yes | Yes | Yes | Yes | No | Yes | No | N.A.
2018-06-07 | No engineering meeting.
2018-06-14 | No engineering meeting.
2018-06-21 | No engineering meeting. (MTR meeting)
2018-06-26 | Yes | Yes | Due to Beacon build. | Yes | Yes | Due to Qda Testing | Yes | No | Yes | No | N.A.
2018-07-05 | Yes | Yes | Due to QDA MR audit | Medical Leave. | Yes | Due to Qda Testing | On leave | No | Yes | Yes | N.A.
2018-07-12 | No engineering meeting.
2018-07-19 | On leave | Yes | Due to QDA MR audit | Yes | Yes | Due to Qda Testing | Yes | No | Yes | Yes | N.A.
2018-07-26 | Yes | YES | QDA meeting | Review tooling with Keypoint | Yes | On leave | Yes | Yes | Yes | Yes | N.A.
2018-08-03 | No engineering meeting.
2018-08-09 |  |  |  |  | Public Holiday
2018-08-16 | No | Yes | Annual Leave | Medical Leave. | Yes | Due to Qda Testing | Yes | No | Yes | Yes | N.A.
2018-08-23 | No engineering meeting.