## Flex
[5 empty] | Shipped |  | OK |  | Potential delay, need to monitor |  |  |  | Delay, need immediate action
Updated On: | 2019-01-04 17:27 |  |  | 2019-01-01 |  |  |  | 2019-02-01 |  |  |  | 01.13.19 |  |  |  |  | DO NOT CHANGE
Instrument |  | Q4 Forecast |  | WW01 | WW02 | WW03 | WW04 | WW05 | WW06 | WW07 | WW08 | WW09 | WW10 | WW11 | WW12 | WW13
PN | Short Desc | Prod Plan | QTD Del'd | 2018-12-31 | 2019-01-07 | 2019-01-14 | 2019-01-21 | 2019-01-28 | 2019-02-04 | 2019-02-11 | 2019-02-18 | 2019-02-25 | 2019-03-04 | 2019-03-11 | 2019-03-18 | 2019-03-25 | Type of Issue | Status
186269505 | Alliance X | 350 |  | 20 | 0 | 40 | 0 | 20 | 0 | 30 | 0 | 40 | 40 | 50 | 50 | 60 | Material | Open
186269506 | Alliance X | 1150 |  | 100 | 120 | 80 | 120 | 100 | 100 | 90 | 120 | 80 | 80 | 70 | 70 | 20 | Quality | Close
186001515 | Pump | 30 | 8 | 8 | 0 | 0 | 8 | 0 | 4 | 0 | 6 | 4 | 0 | 0 | 0 | 0 | Tester
186001525 | Pump | 50 | 0 | 0 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 0 | 8 | 0 | 0 | 0 | Manufacturing
186001825 | Pump | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Tooling
WAT207000 | 515 Pump | 70 | 6 | 6 | 12 | 12 | 0 | 12 | 0 | 12 | 0 | 16 | 0 | 0 | 0 | 0
186015001 | BSM | 65 | 15 | 15 | 21 | 12 | 0 | 6 | 0 | 0 | 11 | 0 | 0 | 0 | 0 | 0
186015018 | QSM | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186015080 | ACQUITY UPLC H-Class QSM PLUS | 410 | 32 | 32 | 32 | 42 | 56 | 56 | 40 | 40 | 28 | 44 | 40 | 0 | 0 | 0
186002487 | 2489 Detector | 211 | 22 | 22 | 0 | 0 | 0 | 12 | 6 | 26 | 23 | 24 | 44 | 20 | 20 | 14
186248900 | Detector X | 950 |  | 64 | 84 | 96 | 80 | 80 | 72 | 84 | 50 | 56 | 64 | 70 | 80 | 70
186015031 | 5031 Detector | 283 | 0 | 0 | 48 | 0 | 48 | 48 | 0 | 0 | 32 | 32 | 0 | 32 | 28 | 15
186015032 | 5032 Detector | 250 |  | 24 | 44 | 28 | 0 | 64 | 24 | 0 | 16 | 10 | 0 | 8 | 8 | 24
186015033 | 5033 Detector | 220 | 12 | 12 | 0 | 36 | 0 | 24 | 12 | 36 | 32 | 0 | 40 | 16 | 0 | 12
186015034 | 5034 Detector | 20 | 0 | 0 | 8 | 8 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0
186015053 | 5053 Detector | 8 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0
186002998 | 2998 Detector | 70 |  | 12 | 0 | 0 | 12 | 0 | 0 | 0 | 24 | 0 | 0 | 0 | 22 | 0
186299800 | Detector X | 580 |  | 40 | 64 | 16 | 72 | 8 | 32 | 44 | 28 | 92 | 20 | 68 | 48 | 48
186017002 | Arc 2489 Detector | 176 | 29 | 29 | 0 | 20 | 0 | 0 | 0 | 24 | 24 | 24 | 24 | 16 | 8 | 7
186017003 | Arc 2998 Detector | 170 | 0 | 0 | 0 | 24 | 24 | 16 | 0 | 16 | 10 | 0 | 48 | 16 | 16 | 0
186017012 | Arc Bio 2489 UV/Vis Detector | 20 | 6 | 6 | 0 | 3 | 0 | 0 | 0 | 0 | 6 | 0 | 5 | 0 | 0 | 0
186017013 | Arc Bio 2998 PDA Detector | 15 | 5 | 5 | 0 | 4 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0
Part No | Part Description |  |  | Supplier | Instrument Affected | WPPL Team (Lead) | Milford Contact | Type of Issue | Brief Description of Issue | [5 empty] | Target Recovery Date | Status
| Linearity test software |  |  | 2018-12-31: Flex; 2019-01-07: 186015033 186015034 186017003; 2019-01-14: Desmond; 2019-01-28: Tester; 2019-02-04: 19 Jan 18 update: ECO released to Flex on 17-Jan-2018 05 Jan 18 update: Impacted by linearity test software. Desmond is working with Milford to see if there is possibility of TMD to resume production.; 2019-03-18: Monitoring; 2019-03-25: Close
405002568 | OPTICS BENCH, MACHINING, 2998 |  |  | 2018-12-31: JEP; 2019-01-07: 186299800 186002998 186015034 186017003 186017013; 2019-01-14: Bethany Yin Ying; 2019-01-28: Material; 2019-02-04: 13 Jul 2018 update : Based on current supply from ETLA, buffer stock is being estblished and shortages closed. Upon depletion of JEP OPO, Flex shall cut off the supply from JEP. 06 Jul 18 update : 22pcs (JEP) ETA 5-Jul, 75pcs ETA 5-Jul. Shortages closed. Flex to continue monitor delivery/support from ETLA. 29 Jun 18 update : ETLA will increase the shipment Qty to 75pcs/wk starting wk 2-Jul. Flex to closely monitor. 22 Jun 18 update : No impact for Q2 shipment. ETLA will confirm on possible to increase to 75pcs/wk. 14 Jun 18 update : working on 80pcs/wk starting 5-Jul. 6/14: 12pcs (JEP) & 60pcs (ETLA) received 08 Jun 18 update: 108pcs received from 1-Jun to 7-Jun, sufficient to cover until 20-Jun shipment. Confirming with ETLA for increased output to 60pcs/wk starting week 11-Jun 01 Jun 18 update : 146pcs received between 23-May to 31-May. Shipment impact on 6-Jun (48units). ETLA: 40pcs/wk to be delivered starting week 11-Jun. 25 May 18 update : Daily shipment from JEP & 40pcs/week from ETLA. Next delivery from ETLA on 31/May. Shipment impact on 30-May (64units). 18 May 18 update : Daily shipment from JEP & 30pcs/ week from ETLA. Next shipment from ETLA will be on 05/21. Shipment impact on 14u week 20(299800). 11 May 18 update : JEP provided WIP status, there is no firmed qty as in daily breakdown. Immediate impact on 9-May (49units). Due to cosmetic issue, 24pcs will only be delivered on 12-May (Sat) due to Malaysia’s holiday. ELTA will be delivery weekly basis starting week 7-May, 30pcs/wk. 04 May 18 update : Total 43pcs received on week 18. 30pcs from ETLA ETA on 10-May-2018. DC allocation :- 186299800 x 80 by week 19 186002998 x 10 by week 19 186017003 x 6 by week 19 27 Apr 18 update : 32pcs ETD 27-Apr. Flex will immediately release to production once all pass inspection check. DC allocation :- 186299800 x 80 by week 19 186002998 x 10 by week 19 186017003 x 6 by week 19 20 Apr 18 update : Based on current Qty received 127pcs at Flex, can full fill production requirement till week 17. DC allocation :- 186299800 - 50 units on week 17 (Top priority customer order + forecast) 186015034 - 2 units by week 17 186002998 - 6 units by week 17 186017003 - 16 units by week 18 13 Apr 18 update : Dimension corelation issue btw Flex and JEP after alignment Flex will use JEP data for 2 dimension.Total 39pcs can be released, 18pcs at Flex pending data, 44pcs to be delivered by JEP on Monday(16-Apr). Balance 24pcs on Tuesday(17-Apr). DC alloation :- 186299800 - 50 units on week 16 (Top priority customer order + forecast) 186015034 - 3 units by week 17 186002998 - 6 units by week 17 186017003 - 16 units by week 17 06 Apr 18 update : Total 55pcs rejected due to dimension out (dowel pin hole). JEP to deliver 20pcs on 06-Apr. ship impact: Immediate (33units) and 11-Apr, 66units DC alloation :- 186299800 - 40 units by next week (Top priority customer order + forecast) 186015034 - 3 units by week 17 186002998 - 6 units by week 17 186017003 - 16 units by week 17 29 Mar 18 update : Quality issue (Porosity) needs to be sorted out with Waters and Flex. 56 received, next shipment impact on 11Apr (55units). Alternately, ETLA has been approved and green light was given to Flex to issue PO to ETLA to drive material. 23 Mar 18 update : Continue to monitor. Next shipment impact on 4-Apr (47units) 16 Mar 18 update : Continue to monitor. Next shipment impact on 28-Mar: 1unit. 09 Mar 18 update: Continue to monitor. Next shipment impact on 28-Mar: 19units. 02 Mar 18 update : Continue to monitor. Next shipment impact on 28-Mar: 49units 25 Feb 18 update: Continue to monitor. Next shipment impact on 21-Mar. 14 Feb 18 update: 30pcs deliverd on 08 Feb. Continue to monitor. Next shipment impact on 28 Feb 14 units. 09 Feb 18 update : Weekly shipment from JEP. Continue monitoring. Next shipment impact on 28-Feb: 14units 02 Feb 18 update : Weekly shipment from JEP. Continue monitoring. Next shipment impact on 28-Feb (43units) 26-Jan-18 update : Next shipment impact 28-Feb-2018 19 Jan 18 update : 22units delivered on 18-Jan. Next impact will be on 14-Feb (56units). Flex is continue following up with supplier 12 Jan 18 update: 97pcs received on 04-Jan. 07-Feb: 11units. Flex is continue following up with supplier 05 Jan 18 update: Flex to continue to follow up with JEP. 29 Dec 17 update: 53pcs delivered on 22 Dec. Next potential impact 48units 17 Jan. 22 Dec 17 update: 24pcs delivered on 22 Dec. Next potential impact 10 Jan - 22units. With above Q4 shipment is covered. 15 Dec 17 update: 20pcs delivered on 14 Dec. 19pcs received 15 Dec. 20pcs to be delivered on 18 Dec. With above Q4 shipment is covered. 08 Dec 17 update: 25pcs delivered on 07 Dec. Potential shipment impact on 27-Dec: 39 units. Flex to continue to follow up with JEP. 01 Dec 17 update: 49 pcs delivered on 30/11. Revised shipment impact will be 19 units on 27/12. With consideration of the upside, potential shipment impact 64pcs for this quarter. Continue to follow up with JEP. 24 Nov 17 update: Rejection at JEP due to porosity issue. Potential shipment impact on 13 Dec, 36pcs. For the other 2 parts from JEP (405003295 & 405006788), sufficient qty to meet this quarter shipment and next shipment impact is on 10-Jan-18. Flex to continue follow up. 03 Nov 17 update: Flex to continue to follow up weekly delivery with JEP. 20 Oct 17 update: JEP committed to deliver 65 units weekly. Flex to continue to follow up weekly delivery with JEP. 13 Oct 17 update: 10/12: 57pcs delivered on 12-Oct 10/10: 60pcs delivered on 5-Oct 06 Oct 17 update: Potential shipment impact 71units, 01 Nov. 29 Sep 17 update: Shipment impact 25-Oct: 42units. Flex to continue to monitor weekly delivery. 22 Sep 17 update: Delivery is scheduled for 22/09. Moving forward JEP will deliver 65pcs every week on Thursday and gradually increase the qty if required. 15 Sep 17 update: 24pcs delivered on 14-Sep-17 late afternoon 48pcs scheduled to be delivered today late afternoon With this, new shipment impact is 18-Oct-17 for 46units Flex working with supplier for further schedule. 08 Sep 17 update : - 6pcs reject out of 76pcs from 5-Sep Delivery - the revised impact is 4-Oct-45 units. - JEP has not provided any further ship plan yet. Flex will keep WPPL updated about the status for this part. 31 Aug 17 update : Based on accepted parts, impacting WWK38, 17 units. According to Flex, a new guideline been set for acceptance of the parts with porosity. All the rejects can be used: 36pcs returned from JEP have been received 36pcs at JEP FG - delivered 31 Aug 2pcs delivery TBA STX has shipped raw casting (306pcs) as well to JEP with ETA as 31-Aug-17. Current Lead Time confirmed by JEP is 2Wks turn-around from casting received date. Flex will continue to work with JEP for arranging delivery from new castings as early as possible. In addition to the above qty, ACP has another 48pcs as on yesterday. 25 Aug17 update: High reject at incoming due to cosmetics and poor finishing. Impacting: 1. 120pcs RTV back to JEP and sorted. 2. 42pcs with good condition received. 3. 34pcs ETA on 25 Aug. As per Flex, JEP will continue do sorting and update result daily. Next shipment impact 13 Sep, 34 units. 18 Aug 17 update: Reject of defect Optic bench. Flex return all for sorting. Delivery commitment cover until end August. To monitor closely. 76pcs: 22 Aug 78pcs: 29 Aug; 2019-03-18: Monitoring; 2019-03-25: Close
WAT045248 | HPMV Sub Assy Testing |  |  | Flex | 186269505 186269506 WAT045438 | Desmond | Mike Higgins Fred Sarrifi | Manufacturing | 19 Jan 18 update: TMD 010368 has been closed with the effect of ECO9217360 released to Flex on 12.01.2018. 17 Nov 17 update: TMD expire on 01 Feb 2018. Desmond to follow up on ECO for closure prior to TMD expired targeted Jan. 18 Aug 17 update: To follow up on ECO closure, Jan 2018. 04 Aug 17 update: TMD010368 has been issued to allow Flex only perform 100% testing on WAT045248 sub assembly built into spare WAT045438. All sub assembly built into instrument will be tested 100% at final test, instead of sub-assembly level. This would free up the capacity constraint to support instrument build. To follow up with Flex on recovery plan for Spares. 28 July 2017 update : Due to obsolete of SOP WAT077886O Sampling Inspection / Testing, Flex is required to perform 100% testing per TSP735000287, instead of sampling check at 1 out of 8 units for sub assemblies that are built into instruments. This has increased the cycle time and impact the line capacity to fulfill both instruments and spares demand. ECO need to be raised to update ASP for WAT045248 to call for 100% testing per TSP735000287 for all spares only. | [5 empty] | Monitoring | Close
NA | Firmware issue |  |  | 2018-12-31: NA; 2019-01-07: 186015032 186015033; 2019-01-14: Edwin Koh; 2019-01-28: Tester; 2019-02-04: 14 Feb 18 update: 68pcs received on 12-Feb. Another 108pcs to be shipped on 16-Feb. Next shipment impact 28 Feb 72 units. Recovery in progress. 02 Feb 18 update : Recovery in progress. Please refer to recovery tab 26 Jan 18 update : TMD010453 being issued to rework all 186015032 and 186015033 built with On-Semi PDA chips. 5 units affected at ADC has returned to Flex on 26-Jan. Flex has implemented ECO9218579, resume production line with the configuration of Hamamatsu PDA chips and v1.69 firmware 19 Jan 18 update : Flex to put both 186015032 & 186015033 build with old PDA chips on hold, while waiting for Waters final decision if to reload the v1.65 old formware. Waters will provide TMD to Flex should the decision is for Flex to reload with v1.65 old frmware.; 2019-03-18: Monitoring; 2019-03-25: Close
425000540 | MASK, PDA |  |  | Tech-Etch | 186299800 186002998 186015033 186015034 186017003 | Bethany | John Gannon | Material | 02 Feb 18 update : Potential impact on 21-Feb(250pcs): 33units. Supplier confirmed to ship on 2-Feb with ETA 7-Feb to Flex. Salt spray test report will be available by end of this week. 26 Jan 18 update : 150pcs received at Flex and pending for the Salt spray test report. 19 Jan 18 update : Previous lot was rejected due to plating issue and parts were all scrapped and produce the new lot.. Supplier will be shipping the material on 19-Jan-2018 and follow by Salt spray test report. | [5 empty] | ASAP | Close
405003295 405006788 | OPTICS BENCH, MACHINING,2998LG OPTICS BENCH,MACHINING |  |  | 2018-12-31: JEP; 2019-01-07: 186015032 186015033; 2019-01-14: Eugene; 2019-01-28: Quality; 2019-02-04: 02 Feb 18 update : Material has been UAI but no acceptance criteria defined. Continue monitoring. No impact to production. 26-Jan-18 update : Part is being on-hold at Flex QA due to having minor porosity. They are neither having light passing through nor near any screw holes where there is no acceptance criteria established. Pending WPPL QA to review and advise; 2019-03-18: Monitoring; 2019-03-25: Close
999042549 | Upper Ram, Check valve Housing |  |  | 2018-12-31: Flex; 2019-01-07: 186269505 186269506; 2019-01-14: Yee Hwa Desmond; 2019-01-28: Tester; 2019-02-04: 10 Aug update: New tool received, under evaluation. 03 Aug update: Check valve tooling broken. New tool received on 03 Aug, under evaluation before commencing production.; 2019-03-18: Monitoring; 2019-03-25: Close
WAT270524 279001810 | GPV ASSEMBLY ASSY, GPV, DUAL CHANNEL, QSM |  |  | PARKER HANNIFIN | 186269505 186269506 186015018 | Bethany | John Gannon | Material | 20 Jul update: Flex to continue to monitor. 17 Jul update : There is a drastic demand increased for Alliance, Flex to review the supply against the requirement. Will follow up with Flex on the supply constraint. 14 Jun update : Q2 shipment no impact. Flex to continue monitoring Parker shipment to ensure ship per schedule. 08 Jun 18 update: no more allocation, next delivery per PO ETA. Confirming next delivery. 11 May 18 update : Based on the latest allocation plan from Parker, no impact to Q2 production plan. Flex will closely monitor shipment with parker. 04 May 18 update : Parker is having capacity constraint after moving to new plant. Material is being put under allocation. John from Milford is working with Parker on the allocation plan to ensure 3 site(Milford, Flex & GES) will be getting material to support production. Parker will be caught up by 3rd week in May, but they will still need to ship FedEx to support the one week in-transit time and expect that Parker will resume normal in-transit time of 2 weeks in early June. 27 Apr 18 update : Parker is having capacity constraint after moving to new plant. Material is being put under allocation. John from Milford is working with Parker on the allocation plan to ensure 3 site(Milford, Flex & GES) will be getting material to support production. Parker will be caught up by 3rd week in May, but they will still need to ship FedEx to support the one week in-transit time and expect that Parker will resume normal in-transit time of 2 weeks in early June. Allocation plan: 1st Priority is PN 279001810 @ 100 units - Flextronics 2nd Priority is PN 279003338 @ 50 units - GES 3rd Priority is PN WAT270524 @ 140 units - Flextronics 20 Apr 18 update : Part# 279001810 - 100 ETA 23/Apr. 150 pc ETD on 20/Apr and 125 pc ETD on 21/Apr(Qty to be confirmed). Part# WAT270524 - 97pcs received. Parker is working this weekend and they will have a ship plan update for next week on 23/Apr 06 Apr 18 update : Parker production is in recovery mode and expected recover to be the second half of April. Parker is arranging weekly shipment. Under monitoring. Next shipment impact on 25-Apr 29 Mar 18 update : Shipment impact on 11-Apr (7units). 100pcs shipping on 30-Mar-2018. 23 mar 18 update : No impact to Q1'2018 shipment. Continue to monitor.Next Shipment impact on 11-Apr (101units) 16 Mar 18 update : Close to for Q1'2018 shipment. Continue to monitor. Next shipment impact on 4-Apr: 53units 09 Mar 18 update: Continue to monitor. Next shipment impact on 28-Mar: 78units. 02 Mar 18 update : Parker allocation plan : 80pcs ETD 3/02, 108pcs ETD 3/06, 106pcs ETD 3/09, 56pcs ETD 3/12. 25 Feb 18 update: Continue to monitor. Next shipment impact on 7-Mar, but Parker is making weekly shipment on Friday to catch up. 09 Feb 18 update : Potential impact on week 09 shipment for Alliance. Parker is having capacity constraint, they are working to catch up the backlog and arrange shipment according to Flex minimum require Qty. 02 Feb 18 update : Potential impact on week 08 shipment for QSM & Alliance. Parker is having component supply difficulties and hence running into backlogs. Supplier is due to ship next lot on 2/2 for both Part#. Flex is closely monitoring. 92pcs shipped for Part# 279001810 on 31-Jan | [5 empty] | Monitoring | Close
NA | Wavelength verification failure |  |  | NA | 186002998 186299800 186017003 186015033 186015034 | Desmond | Eric Turner | Tester | 29 Mar 18 update : Backlog clear 16 Mar 18 update : To continue monitor the shipment. 09 Mar 18 update: Received agreement from Milford team to step down the testing - revert to normal 1x final test with additional power recycling step. To closely monitor production recovery. To complete shipment on last day of Q1 - 29-Mar-18. 02 Mar 18 update : Flex is still struggling to clear backlog. Current direction to Flex is to test final units 2 times to ensure there are no intermittent failure issue. refer to recovery tab. 25 Feb 18 update: Flex is still struggling to clear backlog. Current direction to Flex is to test final units 2 times to ensure there are no intermittent failure issue. To date, the backlog quantity affected are 186299800 102 units, 186017003 43 units, 186015033 29 units, 186002998 12 units and 186015034 3 units. 18 Feb 18 update: Suspected the wavelength verification failure is due to grating PN 402000198 supplied by Horiba. Flex is currently liaising with Horiba to get more testing data details on the recent lots from Horiba. Will need to verify if this is the root cause or due to any other issue. 09 Feb 18 update : Eric got a unit to fail while on hyperterm and got the details to send to the firmware engineer for analysis. Continue run and ship PDA/ELAMBDA units as we are not seeing the wavelength verification failure for new build instrument of PDA and eLAMBDA instruments. Flex will still perform additional test (recycle power and ensure initialization ok) for PDA/ ELAMBDA. Continue evaluation for 2998X, control run 2998x model with limited qty(4x) using 477 charge amp board from different lot code (200118) to see results. | [5 empty] | ASAP | Close
210000444 | ASSY, PCB ALLIANCE H/C |  |  | 2018-12-31: Flex Senai; 2019-01-07: 186269506; 2019-01-14: Edwin Shaynee; 2019-01-28: Quality; 2019-02-04: 13 Apr 18 update : No Impact to production. Shaynee to follow up with Failure analysis report. 06 Apr 18 update : Production continue and so far no impact to production. Pending for Failure analysis from Flex senai. report target to be receivde on end of 06-Apr. 29 Mar 18 update : Flex has started loading 444 PCBA by using IC date code 1736 on 24-Mar. Everything was OK until 26 Mar evening, whereby production reported 3u fail and on 27 Mar production reported another 1u fail. Flex is doing additional H/C testing to continue production. Flex to provide the failure analysis report. 23 Mar 18 update : 1723 date code additional testing no re-curring of the error code observed previously. Flex to continue the rest of the 1723 date code. When date code 1726 arrive Friday 23 Mar, Flex to switch to use date code 1726. Flex KP to work with Flex Senai on the root cause analysis. Flex KP production to continue monitoring of date code 1726. 16 Mar 18 update : Eugene & Edwin agreed using old DC 1723 to continue production, those with pass result is shippable and rejected need to be on hold. Result will be available on Monday. Flex is continue pulling in new DC for Avago MPN from manufacturer and open market. 09 Mar 18: During the weekly FPY review, it is observed that there is a surge in failure of 3 units of 210000444 in week of 02-Mar-18, and the failure analysis for previous months has not been forthcoming. Flex to purge current PCBA lot back to Senai and pull in fresh lot (new datecode) to resume production and access the performance of the new datecode PCBA.; 2019-03-18: Monitoring; 2019-03-25: Close
279003850 | ASSY, SOLENOID VALVE, 3-WAY |  |  | 2018-12-31: Bio-Chem; 2019-01-07: 186269505 186269506; 2019-01-14: Bethany Yin Ying; 2019-01-28: Material; 2019-02-04: 01 Jun 18 update : Shortage closed. Monitoring delivery. 25 May 18 update : Bio-chem shipped 600pcs on 25/May, pending tracking#. Supplier able to increase the output to 400pcs. Tentative potential shipment impact on 13-Jun (42units). 18 May 18 update : 100pcs ETA 05/21, weekly shipment of 200pcs. Shipment impact on on 13-Jun(41u). Next delivery 600pcs ETD : 05/25 (final shipping Qty to be confirmed) 11 May 18 update : Additional 300pcs added to existing demand. All additional PO loaded. 200pcs shipped with ETA 11-May, weekly shipment of 200pcs. 04 May 18 update : 100pcs shipped on 2-May, another 100pcs to ship 4-May. weekly shipment of 200pcs. Shipment impact on 16-May (54units) 27 Apr 18 update : 200pcs received on 21-Apr, 100pcs ship on 25-Apr (confirming tracking, and possible for Sat delivery), follow by balance on ETD 27-Apr. Shipment impact on 2-May (50units) 20 Apr 18 update : 200pcs intransit (ETA 23-Apr), Next ship date on 27-Apr; 2019-03-18: ASAP; 2019-03-25: Close
WAT026574 | SCR 8-32 SHOULDER |  |  | 2018-12-31: M&J; 2019-01-07: 186001515 186001525 WAT207000 WAS207033; 2019-01-14: Bethany; 2019-01-28: Material; 2019-02-04: 11 May 18 update : 500pcs (M&J) shipped on 11-May for ETA 14-May, FA to complete at this period. 04 May 18 update : May shipment is covered. Supplier is pending for Marcorubber part for risk build & FA. Marcorubber committed to ship 5/4, ETA 5/7 at M&J, and ETD M&J 5/9. In separate email from Soo Chin, Marcorubber has identified 500pcs that needs rework, suggesting to rework at M&J once received and see the outcome. Shipment impact on 30-May (4units for SPARE) 27 Apr 18 update : 100pcs support from Milford ETA 30-Apr. Flex to expedite with M&J for better delivery date to ensure production is able to catch up. 20 Apr 18 update : Localized from Ronco to asia supplier M&J. Due to short LT, M&J is pending for Marcorubber part for risk build & FA. Marcorubber committed to ship 5/4 with expediting fees paid, ETA 5/7 at M&J, and ETD M&J 5/9. Pre-alert, supplier’s expediting for improvement. Checking with Milford if able to support.; 2019-03-18: Monitoring; 2019-03-25: Close
279002660 | ASSY, PRE-AMP WITH CABLE 2489 |  |  | 2018-12-31: Flex Senai; 2019-01-07: 186015031 186248900 186017002; 2019-01-14: Hua Zhen; 2019-01-28: Tester; 2019-02-04: 27 Apr 18 update : Loan tester was received on 25-Apr. 10pcs received on 26-Apr, 30pcs in transit, and expecting approx.. 60pcs delivery daily starting 27-Apr. Immediate shipment impact on 25-Apr .50units is in-transit 20 Apr 18 update : Flex Senai will rent a unit to use for production, the rental unit will ETA Senai on 24-Apr (for noise and reference test).; 2019-03-18: ASAP; 2019-03-25: Open
210000444 | ASSY, PCB ALLIANCE H/C |  |  | 2018-12-31: Flex Senai; 2019-01-07: 186269506 WAT270804 WAT270802; 2019-01-14: Yee Hwa; 2019-02-04: 04 May 18 update : Flex to provide PCBA444 CRF Gantt chart (CFT) Flex to identify the IC date code and instrument SN to Waters to justify the remove the additional HC test today (CFT); 2019-03-18: ASAP; 2019-03-25: Close
210000411 | ASSY, PCB, EALLIANCE CPU ROHS |  |  | 2018-12-31: Flex Senai; 2019-01-07: 186269505 186269506; 2019-01-14: Bethany; 2019-01-28: Material; 2019-02-04: 22 Jun 18 update : 25pcs ETD 21-Jun, 40pcs ETD 22-Jun. Subsequent commit provided to close short. Potential Impact : 27-Jun (27units) 14 Jun 18 update : Q2 shipment no impact. 80pcs shipped on 13-Jun. Balance to ship on 19-Jun. Flex to continue follow up for the next delivery date 08 Jun 18 update: 11pcs received on 6-Jun. Component was shipped on 6-Jun, Senai to turn around with ETD 11-Jun (Mon), To monitor closely. 01 Jun 18 update : 200pcs received between 23-May to 29-May. Managed to secure at least 190pcs. Next ETA 8-Jun (40pcs), 9-Jun (40pcs), 12-Jun (80cs) and 13-Jun (30pcs). Following ETA TBA Shipment impact: 13-Jun (77units). 25 May 18 update : Flex kallang will receive total 200pcs boards on 05/28, next shipment impact will be on week 24 (06/13). Senai is able to pull-in and support next gap. Flex Kallang to continue monitor shipment from Senai.; 2019-03-18: Monitoring; 2019-03-25: Close
NA | NA |  |  | NA | 186269505 186269506 | Yee Hwa Desmond | Mike Cotter | Tester | 25 May 18 update : OESR-182578 raised due to Alliance Tester(1,3,5) at Flex Sg having intermittent freeze/hang. Milford request detail of failure. Pending Flex to provide the report. | [5 empty] | ASAP | Close
NA | NA |  |  | NA | 186269505 186269506 | Hua Zhen | Mike Cotter | Tester | 14 Jun 18 update : system has passed functionality verification and Yee Hwa is preparing IQ report. Target to have IQ approval by next week 25 May 18 update : Tester 6 is waiting for part from Milford to finish set up, pending Flex to build proper tester frame to put the Tester 6. Pending Milford to advise a Validation plan from Milford for Tester 6. | [5 empty] | ASAP | Close
279001055 | ASSY, VACUUM PUMP |  |  | 2018-12-31: IDEX; 2019-01-07: 186269505 186269506; 2019-01-14: Eugene; 2019-01-28: Quality; 2019-02-04: 25 May 18 update : Defect : PCBA failure issue. Idex was able to ship 70 units of 0001-6168 to Flex on 05/24, via IDEX FedEx account, priority mail. We expect to also be able to ship approximately 50 units to Flex tomorrow, via IDEX Fed-Ex account, priority. Flex to follow up with tracking#; 2019-03-18: ASAP; 2019-03-25: Close
WAT271804 | BALL SCREW ASSY, PUMP, 2690 |  |  | NSK | 186269505(2x) 186269506(2x) 700001031(1x) 700001030(2x) | Yin Ying Jie Hui | John Gannon | Material | 26 Oct update: 150 pcs shipped from NSK this week. Flex to continue to monitor shipment. 19 Oct update: 300 pcs shipped from NSK this week. Flex to monitor shipment. 12 Oct update: Next shipment impact WK45 (7-Nov) 118 units for instruments, 40u for spare 1031 and 5pcs allocated to Spares 1030. 05 Oct update: Next shipment Impact WK44, 95 untis, spare 50 units. 28 Sep update: 100pcs shipped from NSK this week. Next shipment impact WK42. 21 Sep update : No Impact to Q3 Instrument shipment. New allocation plan shared with Flex, Flex to continue working with NSK for Q4 delivery. Next impact will be on week 41, 17u. 14 Sep update : No Impact to Q3 Instrument shipment. Continue working with NSK for Q4 delivery. 07 Sep update : Q3 shipment closed for instrument. Spare Qty needed is 100sets, Flex will allocate 100pcs to spare and on week 37 shipment from NSK will allocate to Instrument to support for 1st 2 week Oct shipment. 31 Aug update: 271pcs shipped this week. Next shipment impact wk40. 24 Aug update: 100 pcs shipped this week. Next shipment impact WK37. 17 Aug update : Next potential impact on WK36, 34u. 10 Aug update : NSK shipment plan 350pcs. 100pcs received on 10-Aug, balance 250pcs to follow up. 03 Aug update: NSK committed to ship 300 pcs on 03 Aug. To closely monitor. 27 Jul update: 0 qty shipped from NSK this week. Next shipment impact, WK32 - 14 units of 186269505. 20 Jul 18 update: 200pcs shipped ETA 19 Jul. 100pcs shipped ETA 23 Jul. Next shipment impact will be week 32. 13 Jul 18 update: 50pcs shipped ETA 13-Jul. 100pcs shipped, pending for tracking. There is a drastic demand increased for Alliance, John Gannon is working with NSK for capacity and allocation plan to support for 3 site requirement. 06 Jul 18 update: 145pcs shipped, ETA 6-Jul. 50pcs ETA 09-Jul. 22 Jun 18 update : No impact to Q2 shipment. Milford is working with NSK for allocation for Q3 shipment. 08 Jun 18 update: 150 pieces in transit - to close Q2 shipment. Working with NSK now on capacity allocation. 01 Jun 18 update : Capacity issue at supplier end, working for partial, Next ETA TBA. Shipment impact on 27-Jun (56units). Seeing supply gap from week 2-Jul to week 20-Aug. Flex is working on pull-in and setup con-call with supplier. | [5 empty] | Monitoring | Close
289001324 | ASSY, BALLSCREW, 10mmX2.5mm LEAD |  |  | NSK | 186015001(4x) 186015018(2x) 186015080(2x) | Yin Ying Jie Hui | John Gannon | Material | 26 Oct update: 25 pcs shipped from NSK this week. Flex to continue to monitor shipment. 19 Oct update: 150 pcs shipped from NSK this week. Flex to monitor shipment. 05 Oct update: Flex to closely monitor. Next shipment impact WK46. Allocation plan will be share with Flex upon new forecast received. 28 Sep update: 150pcs shipped from NSK this week. Next shipment impact WK43, 22 units. 21 Sep update : No Impact to Q3 Instrument shipment. NSK put this part on kanban. Should be no issue for Q4, Flex to continue to monitor. Next impact on week 43, 46u. 14 Sep update : No Impact to Q3 Instrument shipment. NSK put this part on kanban. Should be no issue for Q4, Flex to continue to monitor. 07 Sep update : Q3 shipment closed for instrument. 31 Aug update: 150 pcs shipped this week. Next shipment impact wk43. 24 Aug update: No quantity shipped this week. Next shipment impact WK38. 17 Aug update : Next potential impact on WK38, 15u. 10 Aug update : NSK shipment plan 100pcs. 50pcs shipped on 08-Aug, balance 50pcs to follow up. 03 Aug update: 200 pcs shipped from NSk this week. 27 Jul update: 250 pcs shipped from NSK this week. Next shipment impact, WK35. 20 Jul 18 update: There is no impact to the production at the moment. Need to closely monitor the shipment plan from NSK. 13 Jul 18 update: John Gannon is working with NSK for capacity and allocation plan to support for 3 site requirement. | [5 empty] | Monitoring | Close
430000101 430000405 430001194 430001910 430001968 WAT270409 | ASSY,TUBING, VALVE TO FILTER ASSY, TUBE, MIXER TO MIXER TUBE ASSY, SAMP LOOP, 100ul, 2 TUBING, .062OD X .009ID X 20.0 TUBING, SS, CUT, .062 X .009 X TUBE ASSY, V2 TO XDUCER |  |  | Keypoint | 186269505 186269506 700001030 186001525 186002998 186299800 | Shaynee | Monica | Material | 07 Sep update : Part UAI and keypoint will resume delivery on 10 Sept (Monday). 31 Aug update: Part was being rejected by Keypoint due to supply of raw tube(Part# WAT024021) is found dirty and the supplier(Handytube) is not able to commit for immediate replacement. Waters SQE/Keypoint is working with Handytube for the replacement part. | [5 empty] | ASAP | Open
415002438 | INSULATION, BOTTOM, DETECTOR, 2489 |  |  | Unified Global Packaging (Rand Whitney) | 186248900 186017002 186017012 | Shaynee | Tim Norton | Quality | 07 Sep update : Supplier shipped wrong revision. Current rev F and supplier still shipping Rev D. Supplier is requesting Waters Milford QA - Tim Norton to issue TMD to use up the deviate inventory while working on a permanent solution. | [5 empty] | ASAP | Open
400007515 | BRACKET, POWER SWITCH, DETECTOR |  |  | 2018-12-31: Baiteng; 2019-01-07: 186248900 186299800 186017002 186017003 186017012 186017013; 2019-01-28: Quality; 2019-02-04: 07 Sep update : Reject defect : Color tone variance. Flex is perfoming sorting for 2 lots incoming (240pcs), will provide sorting result on Monday. Baiteng is also perform sorting for OH inventory (1363pcs) will update the result once available. Flex will arrange air shipment with Baiteng accordingly.; 2019-03-18: ASAP; 2019-03-25: Open
WAT270817 | FAN, 92MM X 25MM, 49 CFM |  |  | 2018-12-31: SAGER; 2019-01-07: 186015001; 2019-01-21: Bob Lincoln; 2019-01-28: Quality; 2019-02-04: 21 Sep update: Pending TMD. 14 Sep update : NC reject due to cable length issue. No OPO. Supplier is working with Bob Lincoln on TMD.; 2019-03-18: ASAP; 2019-03-25: Open
279000881 | X-DUCER ASSY, HEAD MOUNTED, 15 |  |  | 2018-12-31: DYNISCO - DJ INSTRUMENTS; 2019-01-07: 186015001 186015018 186015080; 2019-01-14: Bethany Jie Hui; 2019-01-28: Material; 2019-02-04: 19 Oct update: TMD010564 approved. 475 pcs ETA 19 Oct. Next shipment impact 28 Nov. Flex to continue to monitor shipment for DJ. 12 Oct update: 125 pcs received. 350 pcs from Milford ETA 12 Oct. Next shipment impact WK46. To continue monitor shipment for DJ. 05 Oct update: Milford is in the midst of processing the TMD. Shipment impact WK42 (17 Oct). 28 Sep update: With the below commited ship date and quantity from DJ, we should be all set to reconnect all production line. 75pcs: ETD 28 Sep 125pcs: ETD 08 Oct 125pcs: ETD 10 Oct 250pcs: ETD 22 Oct Flex to closely follow up with DJ to close this quarter.; 2019-03-18: ASAP; 2019-03-25: Open
210000414 210001164 | ASSY,PCB,ACQ,TUV,PERSONALITY ASSY, PCB HAMAMATSU 518 PIX PD |  |  | 2018-12-31: Flex Senai; 2019-01-07: 186015031 186015032 186015033; 2019-01-14: Hua Zhen; 2019-01-28: Tester; 2019-02-04: 04 Jan update: GTS failed self-test after installing the calibrated NI cards; Milford advised to run diagnostics to clean the relay resistance. WPPL tried to run cleaning by remote into GTS, but still not able to pass.
279003850 | ASSY, SOLENOID VALVE, 3-WAY |  |  | 2018-12-31: BIO-CHEM; 2019-01-07: 186269505 186269506; 2019-01-14: Eugene; 2019-01-28: Quality; 2019-02-04: 04 Jan update: Cosmetic reject not in accordance to drawing. Pending UAI.; 2019-03-18: ASAP; 2019-03-25: Open
210000520 | ASSY, PCB, 3VCCPU AHCT |  |  | 2018-12-31: CEI; 2019-01-07: Detectors and Pumps; 2019-01-14: Yin Ying Jie Hui; 2019-01-28: Material; 2019-02-04: 04 Jan update: Q4 upside missing compnent shortage. Under weekly allocation. 225pcs ETA 07 Jan 225pcs ETA 09 Jan Next shipment impact 23 Jan; 2019-03-18: ASAP; 2019-03-25: Open
279001810 | ASSY, GPV, DUAL CHANNEL, QSM |  |  | PARKER HANNIFIN CORPORATION | 186015018 186015080 | Yin Ying Jie Hui | John Gannon | Quality | 04 Jan update: 250pcs received, 75pcs shipped this week. 21 Dec update: 100 pieces ETA 24-Dec. Parker shipment has meet Flex MRP plan (materials arrival is 2 weeks ahead of shipment). Flex to continue monitor. 14 Dec update: 150pcs received by 13-Dec. 50pcs shipped ETA 14-Dec. Flex to continue to monitor subsequent shipment. 07 Dec update: Parker expects to ship 175 pieces by 7-Dec-18. Flex to continue to monitor subsequent shipment. 30 Nov update: 150 pcs received 26 Nov. 150 pcs ETA 03 Dec. Flex to continue to monitor subsequent shipment. 23 Nov update: 150pcs received 22 Nov. Flex to continue to monitor subsequent shipment. 16 Nov update: 100 pcs received. 75 pcs ETA 16 Nov. Flex to continue to monitor. 09 Nov upate: 100 pcs received. Flex to continue to monitor. 02 Nov update: 100 pcs ETD 02 Nov. Flex to continue to monitor. 26 Oct update: Parker is targeting to ship 100 pcs on 26-Oct. To monitor. 19 Oct update: Parker re-started production. John Gannon to advise Parker's shipment and production status. Tentative ETD 26 Oct. 12 Oct update: Shipment hold. Supplier is working with Milford to resolve issue. Recovery plan yet to provide. Next impact on WK44, 23u. 05 Oct update: Production line resume. Outstanding 16u on WK40, Niu Niu to provide | [5 empty] | Monitoring | Close
WAT270524 | GPV ASSEMBLY |  |  | PARKER HANNIFIN CORPORATION | 186269505 186269506 | Yin Ying Jie Hui | John Gannon | Quality | 04 Jan update: 112pcs received, 112pcs shipped this week. 21 Dec update: 140 pieces received. Short 33 pieces for 9-Jan-19 shipment. Flex to continue monitor. 14 Dec update: Parker to ship 28 pcs on 12-Dec; 56 pcs on 13-Dec; & 28 pcs on 14-Dec. Flex to continue to monitor subsequent shipment. 07 Dec update: Parker expects to ship 88 pieces by 7-Dec-18. Flex to continue to monitor subsequent shipment. 30 Nov update: 156 pcs received last week. Flex to continue to follow up. 23 Nov update: 320 pcs received. Flex to continue to monitor. 16 Nov update: 120 pcs from Milford received. 09 Nov update: Parker shipped 56 pieces, Milford to ship 120 pieces. 02 Nov update: Parker shipped 112 pcs this week, ETA 02 Nov. 56 pcs ETD 02 Nov. 26 Oct update: 56 pcs from Milford + 184 pcs shipped from Parker this week. Next shipment impact WW46. 19 Oct update: Parker re-started production. Milford to sell 56 units to Flext to avoid line done on Alliance. Flex would receive material on Monday, 22 Oct. John Gannon to advise Parker's shipment and production status. 12 Oct update: Shipment hold. Supplier is working with Milford to resolve issue. Recovery plan yet to provide. Next impact on WK44, 116u. | [5 empty] | Monitoring | Close
WAT270471 | POWER SUPPLY, 500 WATT |  |  | 2018-12-31: GE POWER; 2019-01-07: 186269505 186269506; 2019-01-14: Bethany Jie Hui; 2019-01-28: Material; 2019-02-04: 26 Oct update: 229 pieces received this week. 200 pieces ETA 2-Nov-18. 600 pieces ETA 12-Nov-18. Flex to continue monitor. 19 Oct update: Shipment delay. First 30 pcs ETA 20 Oct, another 159 pcs ETA 22 Oct. Flex to follow up with subsequent shipment.; 2019-03-18: ASAP; 2019-03-25: Close
352000115 | SENSATA THERMAL SWITCH |  |  | 2018-12-31: EMC RESOURCES PTE LTD; 2019-01-07: 186002998 186017002 186017003 186017012 186017013 186248900 186299800; 2019-01-14: Yin Ying Jie Hui; 2019-01-28: Material; 2019-02-04: 14 Dec update: 300pcs (441001398) received on 7-Dec. Total 573pcs will be received in wk 10 Dec. 350pcs (Sensata) will be shipping out by 14 Dec. Next shortages will in 21 Jan. 07 Dec update: 300 pieces of 441001398 cable from CMTi, received on 07 Dec. Flex to continue to follow up with subsequent delivery from Sensata. 30 Nov update: 100 pcs from new batch, ETA 11 Dec. To pull in another 300 pieces of 441001398 cable from CMTi, ETA 11 Dec. Flex to continue to follow up with subsequent delivery from Sensata. 23 Nov update: 100 pcs sample ETA 24 Nov. TMD010588 approved for 200 pcs from CMTI. Flex to expedite subsequent delivery from Sensata. 16 Nov update: 300 pcs received. 326 pcs ETA 19 Nov. Next shipment impact 05 Dec 146 units. To expedite subsequent delivery from Sensata. 09 Nov update: 300 pieces ETA 12-Nov, 500 pieces ETD 14-Nov, 200 pieces ETD 19-Nov. Flex to continue follow up shipment for another 500 pieces to close Q4 production requirement. 02 Nov update: Raw material received. 1000 pcs ETD 09 Nov, ETA 14 Nov. 26-Oct-18 update: EMC did not ship parts on 16-Oct-18 as committed. Current revised ETA date is 14-Nov. Will continue to expedite with Sensata.; 2019-03-18: ASAP; 2019-03-25: Close
289000237 | ASSY, COMPARTMENT, FOA |  |  | 2018-12-31: Foamtec; 2019-01-07: 186269505 186269506; 2019-01-14: Yin Ying Jie Hui; 2019-01-28: Material; 2019-02-04: 16 Nov update: 90 pcs ETA 22 Nov.Flex to continue expedite. 09 Nov update: Partial delivery ETA 22-Nov. Still have few days gap to connect production. Flex to continue expedite. 02 Nov update: Raw material shortage. Potential shipment impact 28 Nov. Flex to continue to expedite.; 2019-03-18: ASAP; 2019-03-25: Close
| High Pump Tray Test Failure |  |  | 2018-12-31: Flex; 2019-01-14: Yee Hwa; 2019-01-21: Alex Frazier; 2019-01-28: Quality; 2019-02-04: 16 Nov update: Flex encounter high reject on pump tray sub assembly test. Yield is good when Flex switch over to use new lot of transducer. Flex to continue to production and shipment of new transducer. Need further investigation for testing using of old lot of transducer.; 2019-03-18: ASAP; 2019-03-25: Close
410000760 | FITTING, ONE-PIECE, 10-32 |  |  | 2018-12-31: IDEX; 2019-01-07: 186002487 186269505 186269506 186248900; 2019-01-14: Yin Ying Jie Hui; 2019-01-28: Material; 2019-02-04: 21 Dec update: Parts received. 14 dec update: Missing cargo. Excluding 400pcs cargo, Flex will have following coverage for Q4’18, and shortage week will be 31 Dec. 800pcs – ETA 15-Dec 400pcs – ETA 17-Dec (FEDEX) 500pcs – ETD 17-Dec from IDEX 400pcs – ETA TBA; 2019-03-18: ASAP; 2019-03-25: Close
Note on B7: sgblim: New Seal pack implementation
Note on H7: Jie Hui Ng: BIO-CHEM
Note on B8: sgblim: New Seal pack implementation
Note on G8: Jie Hui Ng: BIO-CHEM
Note on H8: Jie Hui Ng: BIO-CHEM
Note on I9: Jie Hui Ng: CEI
Note on I10: Jie Hui Ng: CEI
Note on I15: Jie Hui Ng: CEI
Note on I17: Jie Hui Ng: CEI
Note on G18: Jie Hui Ng: GTS issue
Note on I18: Jie Hui Ng: CEI
Note on G19: Jie Hui Ng: GTS issue
Note on H19: Jie Hui Ng: GTS issue
Note on H20: Jie Hui Ng: GTS issue
Note on I22: Jie Hui Ng: CEI
Note on I23: Jie Hui Ng: CEI
Note on I24: Jie Hui Ng: CEI
Note on I26: Jie Hui Ng: CEI

## Venture
[5 empty] | Shipped |  | OK |  | Potential delay, need to monitor |  |  |  | Delay, need immediate action
Updated On: | 2019-01-04 17:27 |  |  | 2019-01-01 |  |  |  | 2019-02-01 |  |  |  | 2019-03-01 |  |  |  |  | DO NOT CHANGE
Instrument |  | Q4 Forecast |  | WW01 | WW02 | WW03 | WW04 | WW05 | WW06 | WW07 | WW08 | WW09 | WW10 | WW11 | WW12 | WW13
PN | Short Desc | Prod Plan | QTD Del'd | 2018-12-31 | 2019-01-07 | 2019-01-14 | 2019-01-21 | 2019-01-28 | 2019-02-04 | 2019-02-11 | 2019-02-18 | 2019-02-25 | 2019-03-04 | 2019-03-11 | 2019-03-18 | 2019-03-25 | Type of Issue | Status
186001500 | 1500 CH | 30 |  | 5 | 0 | 5 | 0 | 0 | 5 | 0 | 5 | 5 | 5 | 0 | 0 | 0 | Material | Open
186002441 | TCM | 60 |  | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 4 | 0 | 0 | 0 | 0 | 0 | Quality | Close
186002449 | PCM | 15 |  | 4 | 4 | 4 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Tester
186001791 | Alliance CH/C | 5 |  | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Manufacturing
186001863 | Alliance CH | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Tooling
186179100 | Alliance X CH/C | 830 |  | 64 | 74 | 72 | 76 | 68 | 24 | 76 | 70 | 66 | 70 | 56 | 60 | 54
186186301 | Alliance X CH | 650 |  | 40 | 52 | 60 | 60 | 68 | 60 | 44 | 44 | 48 | 42 | 48 | 48 | 36
186017008 | Arc 30cm CH | 115 |  | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 19
186017009 | Arc 30cm CH/C | 181 |  | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 12 | 24 | 12 | 25
186017014 | Arc Bio 30cm CH | 8 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0
186017015 | Arc Bio 30cm CH/C | 35 |  | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 0
186015007 | CM - Right | 23 | 0 | 0 | 6 | 6 | 6 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0
186015008 | CH/C | 10 |  | 4 | 4 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
186015009 | CM - Left | 5 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186015010 | HTCH | 27 | 0 | 0 | 6 | 0 | 6 | 0 | 0 | 0 | 9 | 0 | 6 | 0 | 0 | 0
186015011 | 30cm CH/C | 165 |  | 14 | 14 | 12 | 12 | 12 | 12 | 12 | 16 | 16 | 16 | 16 | 13 | 0
186015042 | CH-A | 550 |  | 42 | 42 | 42 | 42 | 42 | 42 | 63 | 21 | 63 | 49 | 60 | 42 | 0
186015043 | CM-A | 300 |  | 32 | 24 | 32 | 40 | 32 | 32 | 24 | 32 | 32 | 20 | 0 | 0 | 0
186015045 | HTCH-30A | 90 |  | 8 | 12 | 16 | 12 | 16 | 0 | 0 | 26 | 0 | 0 | 0 | 0 | 0
186015049 | CM-Aux | 75 | 8 | 8 | 8 | 6 | 8 | 24 | 5 | 0 | 8 | 8 | 0 | 0 | 0 | 0
186015038 | UPC2 CM-A | 10 |  | 4 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186015039 | UPC2 CM-Aux | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0
186015017 | FTN | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186015046 | I-Class FTN | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186015085 | ACQUITY UPLC H-Class FTN-H PLUS | 375 |  | 36 | 20 | 16 | 64 | 64 | 24 | 60 | 40 | 51 | 0 | 0 | 0 | 0
186015088 | ACQUITY UPLC I-Class FTN-I PLUS | 169 | 0 | 0 | 32 | 48 | 0 | 0 | 20 | 0 | 20 | 0 | 49 | 0 | 0 | 0
186015029 | FLR Detector | 132 | 0 | 0 | 12 | 12 | 12 | 24 | 6 | 0 | 24 | 8 | 24 | 0 | 0 | 10
186002475 | 2475 Detector | 20 |  | 6 | 6 | 0 | 0 | 0 | 6 | 0 | 2 | 0 | 0 | 0 | 0 | 0
186247500 | 2475 Detector | 150 |  | 18 | 6 | 18 | 18 | 0 | 12 | 24 | 0 | 18 | 0 | 18 | 18 | 0
186017006 | Arc 2475 Detector | 30 | 0 | 0 | 6 | 0 | 0 | 6 | 0 | 6 | 0 | 0 | 6 | 6 | 0 | 0
186015000 | I-BSM | 10 |  | 9+1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186017000 | ACQUITY Arc QSM-R | 320 |  | 30 | 26 | 24 | 26 | 28 | 20 | 36 | 24 | 24 | 24 | 24 | 18 | 16
186015082 | ACQUITY UPLC I-Class BSM PLUS | 275 |  | 9 | 22 | 24 | 22 | 20 | 18 | 12 | 24 | 24 | 24 | 24 | 28 | 24
289011842 | BEACON HTL MKII TOP ASSY | 5 |  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
289011912 | FRONT HOUSING FLANGE ASSY, BEACON | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
289011915 | OUTER SOURCE ASSY,BEACON | 15 |  | 8 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
289011937 | ASM, INNER SOURCE, BEACON | 7 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
186008511 | ACQUITY QDa Detector | 12 | 0 | 0 | 0 | 6 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0
Part No | Part Description |  |  | Supplier | Instrument Affected | WPPL Team (Lead) | Milford Contact | Type of Issue | Brief Description of Issue | [5 empty] | Target Recovery Date | Status
NA | NA |  |  | 2018-12-31: NA; 2019-01-07: 186001500; 2019-01-14: Edwin Desmond; 2019-01-28: Tester; 2019-02-04: 26-Jan-18 update : Tester issue resolved. 19 Jan 18 update : Failure symptom: UUT hang at Firmware download test step.; 2019-03-18: ASAP; 2019-03-25: Close
402000130 | Grating |  |  | 2018-12-31: OPTOMETRICS; 2019-01-07: 186002475 186247500 186017006 700002035 700003707 700009344; 2019-01-14: Edwin; 2019-01-28: Quality; 2019-02-04: 13 Apr 18 update : Based on the manufacturing data provided and the gratings samples shipped to US, it appears there is inconsistency in performance on the gratings manufactured by the new master grating that Waters approved in Sept 2017. The drift in performance likely happen on later lots from Dec 2017 onwards which leads Milford team to believe it a process related issue at this supplier Optometrics rather than a design issue. For Venture manufacturing - action is to continue alignment on 2475 models with sorted higher efficiency grating > 49.1% while Milford team and Optometrics works to address the problem. 06 Apr 18 update : Next incoming 177pcs ETA 9 Apr. Continue to use high efficiency grating for 247500. 29 Mar 18 update : Q1 customer orders is covered. However, Venture is failed to meet Q1 production plan. Venture is continuing build with new grating and provide daily update. 23 Mar 18 update : 10u+5spares away from customer orders. 53pcs grating received, 29pcs > 49.1% effeciency. Venture is continuing build with new grating and provide daily update. 16 Mar 18 update : As of 16-Mar-18 morning, 11u 186247500, 8u 186015029 and 1u 186017006 passed and shipped from Venture. Balance 43u 185247500, 8u 186015029, 1u 186017006 and 3u 186002475 to close the quarter. Milford is in progress to draft the TMD to allow Venture to use grating with energy level >900, with the condition it passed specification at final test level. 02 Mar 18 update : New batch is giving the same fallout results, Venture is continuing production by performing sorting. 30% of lot passing alignment to be used for 2475. Balance 70% will be used for 5029. Venture to pull in more material to support production. 25 Feb 18 update : New batch of parts ETA schedule (30 pieces ETA 26-Feb, and 70 pieces ETA 1-Mar). Upon receiving the new batch, Venture engineering will be assembling these into the Optics Bench, and from there, it will be determine whether does these still have the “Reference energy failure issue”. 15 Feb 18 update: FLR detector output are affected by output of optic bench - cannot achieve energy level after tuning. Suspected this is due to grating change. Still under investigation with supplier for more details.; 2019-03-18: Monitoring; 2019-03-25: Close
430002640 430003271 | NEEDLE, .039 OD x .005 ID W/C'SINK (405013771) |  |  | Machine shop | 186015017 186015046 | Jen | Lisa McNarama | Quality | 13 Ap 18 update : No impact to production at GES There is no stop ship at instrument level as determined by PIM (Lisa McNarama) and Point of Contact at Waters Milford (Scott Gibsob). TMD 010489 sent to Venture to allow part# 405013771 out of tolerence for the true position at the tip of the needle, and true position at counterbore and straightness. 06 Apr 18 update : it was discovered in the machine shop that the needle on assembly 430005254 may be potentially bent which may result in functional failure. Stop ship issued to purge spares and kits with 430005254 of date code 02.Mar.2018 to 03.Apr.2018. In progress to confirm actual material impact at GES. | [5 empty] | Monitoring | Close
279003338 | ASSY, GPV - GROUND BRACKET, VELOX |  |  | Parker | 186017000 | Bethany | John Gannon | Material | 18 May 18 update : Total received 55pcs. Q2 shipment is covered. Venture to continue monitoring Parker shipment. 11 May 18 update : 23pcs ETA 05/14. 32pcs shipped on 05/11 follow up with tracking# 04 May 18 update : Parker is having capacity constraint after moving to new plant. Material is being put under allocation. John from Milford is working with Parker on the allocation plan to ensure 3 site(Milford, Flex & GES) will be getting material to support production. Parker will be caught up by 3rd week in May, but they will still need to ship FedEx to support the one week in-transit time and expect that Parker will resume normal in-transit time of 2 weeks in early June. 27 Apr 18 update : Parker is having capacity constraint after moving to new plant. Material is being put under allocation. John from Milford is working with Parker on the allocation plan to ensure 3 site(Milford, Flex & GES) will be getting material to support production. Parker will be caught up by 3rd week in May, but they will still need to ship FedEx to support the one week in-transit time and expect that Parker will resume normal in-transit time of 2 weeks in early June. Allocation plan: 1st Priority is PN 279001810 @ 100 units - Flextronics 2nd Priority is PN 279003338 @ 50 units - GES 3rd Priority is PN WAT270524 @ 140 units - Flextronics | [5 empty] | Monitoring | Close
279001352 | POWER SUPPLY MERCURY XENON LAMP 150 W |  |  | 2018-12-31: WARNER; 2019-01-07: 186002475 186015029 186247500 186017006; 2019-01-14: Jie Hui Yin Ying; 2019-01-28: Material; 2019-02-04: 13 Jul 18 update: 100pcs ETD 20 Jul. Pending tracking number. 06 Jul 18 update: 100pcs ETA 20 Jul(ETD 14 Jul). Venture is working with supplier for partial delivery.; 2019-03-18: ASAP; 2019-03-25: Close
| Clean Room Aircon |  |  | 2019-01-07: 186015029 186017006 186247500; 2019-02-04: 03 Aug update: Clean room aircon breakdown 27 July and recovered on 02 Aug. Production in recovery mode.; 2019-03-18: Monitoring; 2019-03-25: Close
615005084 | ACQUITY UPLC I CLASS PLUS BSM |  |  | 2018-12-31: TECHPRINT INC; 2019-01-07: 186015082; 2019-01-14: Jie Hui Yin Ying; 2019-01-28: Material; 2019-02-04: 27 Jul update: 1K rejected due to colors discrepancy. ETA 27 Jul, 40 pcs from Milford.; 2019-03-18: Monitoring; 2019-03-25: Close
289001289 | ASSY, BALLSCREW, 4MM LEAD, PRELOADED |  |  | NSK | 186015017 186015046 186015085 186015088 700005201 | Jie Hui Yin Ying | John Gannon | Material | 26 Oct update: NSK is scheduled to ship 90 pieces on 26-Oct-18. To follow up tracking number. 19 Oct update: NSK is experiencing quality issues on a component supplied from Japan. 20U of spare (5201) is going to push out to 5 Nov, therefore, no issue with FTN shipments for Oct. If NSK to ship on 25 Oct (with enough qty), estimated ETA GES will be 29 Oct. If no any other issue arises, Venture can still turn around to ship on 7 Nov. 12 Oct update: 12 pcs shipped from NSK this week. Next shipment impact WK44. 05 Oct update: 120pcs shipped from NSK this week. Next shipment impact WK44. 28 Sep updatae: 60pcs shipped from NSK this week. Next shipment impact WK42. 21 Sep update: 100 pcs shipped from NSK this week. Next shipment impact WK41. 14 Sep update: 180 pcs shipped from NSK this week. Next shipment impact WK40. 07 Sep update: No shipment from NSK this week. Next shipment impact WK37. Continue to expedite. 31 Aug update: NSK had not ship as of this week. Continue to expedite. 24 Aug update: No quantity shipped from NSK this week. Next shipment impact WK37. 17 Aug update: 60 pcs shipped from NSK this week. Next shipment impact WK36. 10 Aug update: 60 pcs shipped from NSK this week. Next shipment impact WK35. 03 Aug update: 60 pcs shipped from NSK this week. Next shipment impact WK34. 27 Jul update: 120 pcs shipped from NSK this week. Next shipment impact WK32. 20 Jul update: received 30pcs on 20 Jul. 60pcs ETA 23 Jul. With the 90pcs, still short of 51pcs for 25 Jul shipment. 13 Jul update: received 60pcs on 10 Jul. 50 pcs ETA 16 Jul. John Gannon is working with NSK for capacity and allocation plan to support for 3 site requirement. 06 Jul update: 60pcss support from Milford will ETA : 09-Jul (Monday). Subsequent delivery to be updated on on 10-Jul once NSK resume work. | [5 empty] | ASAP | Close
289001324 | ASSY BALLSCREW 10MMX2.5MM LEAD |  |  | NSK | 186015000 186015082 186017000 | Jie Hui Yin Ying | John Gannon | Material | 26 Oct update: 75 pcs shipped from NSK this week. Venture to continue to monitor. 19 Oct update: 75 pcs shipped from NSK this week. Venture to continue to monitor. 12 Oct update: 75 pcs shipped from NSK this week. Venture to continue to monitor. 05 Oct update: 150pcs shipped from NSK this week. Venture to continue to monitor. 21 Sep update: Current inventory can cover till November. 14 Sep update: 175 pcs shipped from NSK this week. NSK put this part on kanban. Should be no issue for Q4, Venture to continue to monitor. 07 Sep update: 100 pcs shipped from NSK this week. Cleared for Q3. 31 Aug update: 200 pcs shipped from NSK this week. Next shipment impact WK43. 24 Aug update: 100 pcs shipped from NSK this week. Next shipment impact WK38. 17 Aug update: No quantity shipped from NSK this week. Next shipment impact WK38. 10 Aug update: 75 pcs shipped from NSK this week. Next shipment impact WK38. 03 Aug update: 125 pcs shipped from NSk this week. Next shipment impact WK37. 27 Jul update: 225 pcs shipped from NSK this week. Next shipment impact WK35. | [5 empty] | Monitoring | Close
415001314 | INSULATION FRONT SAMPLE COMPARTMENT |  |  | Unified2 Global Packaging Group, LL | 186015017 186015046 | Jie Hui Yin Ying | Tim Norton | Material | 03 Aug update: Pending Waters Milford to approve CPK report. ETD 03 Aug, 100 pcs from Milford to support production at Venture. 27 Jul 18 update: CPK report was already sent to Milford on Tuesday. Now pending for Waters to advise the next steps. Potential impact WK34 shipment. 13 Jul 18 update: Potential impact on first week of Aug shipment. | [5 empty] | ASAP | Close
289010629 | ASSY, CHASSIS, PUMP |  |  | 2018-12-31: Baiteng; 2019-01-07: 186017000; 2019-01-14: Wan Jing; 2019-01-28: Quality; 2019-02-04: 05 Oct update: 120pcs good part received from Baiteng. Next shipment impact WK47 (21 Nov). 28 Sep update: After sorting, 40pcs on hand. There will be 167pcs incoming, with this can cover till 24 Nov. Next shortage 28 Nov, 1pcs. 21 Sep update: DFM is in progress. Venture to continue to monitor. 14 Sep update: Baiteng will proceed production per current DWG until DFM has been reviewed and drawing is updated. Pending on Baiteng's recovery plan. 07 Sep update: Production detected quality issue loose QSM Bracket. Pending Venture recovery plan.; 2019-03-18: Monitoring; 2019-03-25: Close
615004795 | LABEL DOOR ACQUITY ARC |  |  | 2018-12-31: TECHPRINT INC; 2019-01-07: 186017000; 2019-01-14: Wan Jing Bethany; 2019-01-28: Quality; 2019-02-04: 05 Oct update: FA is in progress. 28 Sep updatae: Had trigger Venture to activate FA with Zephyr Silkscreen to localise this part. 21 Sep update: 100 pcs from Milford, verified ok. Continue to work with Venture. 14 Sep update: 'Discoloration' on label door. Had get inventory support of 100 pcs from Milford.; 2019-03-18: Monitoring; 2019-03-25: Close
410007145 | THIN FILM RTD SLEEVED |  |  | 2018-12-31: Omega; 2019-01-07: 289011842; 2019-01-28: Quality; 2019-02-04: 12 Oct update: New replacement part had been received (10 Oct) and released to production. 05 Oct update: New replacement parts had been built according to the corrected drawing ETD 05 Oct. SCAR had been issued to Omega. 28 Sep update: OMEGA will ship the replacement on WK40. 21-Sep: Venture highlighted that there is quality issue encountered during build - sleeve diameter unable to fit into the mating part.; 2019-03-18: ASAP; 2019-03-25: Close
405012136 | TROUGH RETAINER, BOTTOM, REWORK, CM-A |  |  | 2018-12-31: MRPC; 2019-01-07: 186015043 186015038 186015049; 2019-01-14: Jie Hui Wan Jing; 2019-01-28: Quality; 2019-02-04: 05 Oct update: 951pcs received. Venture to continue to follow up. 28 Sep update: Received 336pcs. Next shipment impact WK43. 21 Sep update: 336 shipped from MRPC ETA 24 Sep. MRPC will continue to produce 1092 pcs and ship via expedite mode. Venture to follow up on recovery plan.; 2019-03-18: Monitoring; 2019-03-25: Close
279003338 | ASSY GPV - GROUND BRACKET VELOX |  |  | PARKER HANNIFIN CORPORATION | 186017000 | Jie Hui Wan Jing | John Gannon | Quality | 30 Nov update: 60 pcs received last week. 25 pcs ETA 03 Dec. Venture to continue to follow up. 23 Nov update: Parker shipped 60 pcs ETA 26 Nov. Venture to continue to follow up. 16 Nov update: Parker shipped 40 pcs on 09 Nov. Venture to continue to follow up. 09 Nov update: Parker shipped 56 pcs om 05 Nov. Venture to continue to follow up. 02 Nov update: Partial delivery 27 pcs ETD 05 Nov. Parker to ship 56 pcs every week. 26 Oct update: Venture has received 57 pieces of 279003338. Production recovery in progress. Target to have first 46 units shipped on 31-Oct-18. 19 Oct update: Parker re-started production. John Gannon to advise Parker's shipment and production status. Target to ship 56 pcs to Venture on 20 Oct. 05 Oct update: Production line stop. High production reject, Venture do not have any good part to run production. Pending Parker to advise delivery date. | [5 empty] | Monitoring | Close
279001951 | XDUCER PRESSURE FLEX CABLE 20K PSI |  |  | DYNISCO - DJ INSTRUMENTS | 186015000 186015082 | Yin Ying Jie Hui | Dave Reed | Material | 19 Oct update: TMD010564 approved. 298 pcs received. Next shipment impact 28 Nov. 12 Oct update: ETD 15 Oct qty 199pcs. No line impact. | [5 empty] | Monitoring | Close
279003625 | ASSY PRESSURE TRANSDUCER TI 15KPSI |  |  | DYNISCO - DJ INSTRUMENTS | 186017000 | Yin Ying Jie Hui | Dave Reed | Material | 19 Oct update: TMD010564 approved. 240 pcs received. Next shipment impact 05 Dec. 12 Oct update: ETD 24 Oct qty TBA. 05 Oct update: Milford is in the midst of processing the TMD. Shipment impact WK43 (24 Oct). | [5 empty] | Monitoring | Close
441000392 | ASSY CABLE HV ANODE PANEL MOUNT |  |  | 2018-12-31: ATRONIX; 2019-01-07: 186002475 186015029 186017006 186247500; 2019-01-14: Jie Hui Yin Ying; 2019-01-28: Material; 2019-02-04: 09 Nov update: Atronix shipped 300 pcs on 09 Nov. ETA 12-Nov. 02 Nov update: Atronix improve shipment to ETD 02 Nov from Mexio. Venture to follow up. 26 Oct update:Atronix is facing shortage of one component, ETA 7-Nov-18 at Atronix. Atronix committed to ship on 10-Nov-18. To continue expedite.; 2019-03-18: ASAP; 2019-03-25: Close
415002153 415002154 | GUIDE COVER TOP NANO ASM-BSM GUIDE COVER BOTTOM NANO ASM-BSM |  |  | 2018-12-31: GREENE RUBBER; 2019-01-07: 186017000; 2019-01-14: Jie Hui Wan Jing; 2019-01-28: Quality; 2019-02-04: 23 Nov update: TMD010589 aprroved. 16 Nov update: TMD010589 pending approval. Venture raise internal DE to allow production build and test instrument with the PNs: 415002154 & 415002153. However shipment cannot be made without approve TMD. 09 Nov update: Wrong part received on 415002154 & 415002153 from Greene Rubber. Venture to get replacement, pending for supplier respond. 02 Nov update: 415002153 - 250 pcs received. 415002154 - ETA 02 Nov. Venture to continue to follow up. 26 Oct update: Greene Rubber is facing quality issue with the parts, and they are currently working on correcting the material after discuss with Waters Milford. Greene Rubber is recovering the production right now, and commit to ship on 31-Oct-18. Had requested to ship partial on 26-Oct-18, with allocation done for both Milford and Venture.; 2019-03-18: ASAP; 2019-03-25: Close
415000680 | ELBOW DRAIN 1/8 NPT TOBARBED |  |  | 2018-12-31: MRPC; 2019-01-07: 186015010 186015029 186015042 186017006; 2019-01-14: Wan Jing; 2019-01-28: Quality; 2019-02-04: 16 Nov update: TMD010584 approved, Venture to proceed to use the burr elbows in new production and to continue with shipment. 09 Nov update: Pending TMD approval to allow usage of on hand parts at Venture. 02 Nov update: Fitting issue (end of Elbow uneven with basin bottom). Replacement part ETA on 2 Nov. Pending inspection result.; 2019-03-18: ASAP; 2019-03-25: Close
425000888 | GASKET THERMAL GRAPHITE CM |  |  | 2018-12-31: LAIRD; 2019-01-07: 186015043 186015049 186015038; 2019-01-14: Wan Jing; 2019-01-28: Quality; 2019-02-04: 09 Nov update: Total 140 pieces and 700 pieces received and accepted. 02 Nov update: Wrong part received. 140 pcs from Milford to support inventory. 700 pcs shipped out 1 Nov from Laird Reynosa Mexio. To monitor shipment and inspection result.; 2019-03-18: ASAP; 2019-03-25: Close
402000122 | MIRROR 40 DEG PARABOLA MOLDED |  |  | GS Plastic Optics | 186002475 186015029 186017006 186247500 | Jie Hui Yin Ying | Matt Franklin Nicholas Szwaja | Quality | 23 Nov update: Received 35u, next supplier ETD 5'Dec , ETA 10'Dec 16 Nov update: Cosmetic issues on mirror. 35 pcs ETD 16 Nov. Venture to follow up with subsequent shipments. | [5 empty] | ASAP | Close
410004033 | COUPLING, SHAFT, .250 X 6MM |  |  | 2018-12-31: HELICAL; 2019-01-07: 186015082; 2019-01-14: Jie Hui Yin Ying; 2019-01-28: Material; 2019-02-04: 23 Nov update: Rcvd 684pcs 19/Nov. Shortage on Jan'19 shipment. Will build ahead in Dec. 16 Nov update: Suppler delinquent, ETA 19 Nov. Venture to follow up.; 2019-03-18: Monitoring; 2019-03-25: Close
289010629 | ASSY CHASSIS PUMP |  |  | 2018-12-31: Baiteng; 2019-01-07: 186017000; 2019-01-14: Jie Hui Yin Ying; 2019-01-28: Material; 2019-02-04: 22 Nov update: TMD010590 approved. Venture to work with Baiteng for rework. 16 Nov update: TMD010590 pending approval. Venture to work with Baiteng to rework chassis for production use.; 2019-03-18: Monitoring; 2019-03-25: Close
405004811 | Tray Pre-Heater Four Column H/C |  |  | 2018-12-31: Thermo Cool; 2019-01-07: 186015007 186015008 186015009 700003816 700005536 700003155; 2019-01-14: Jie Hui Wan Jing; 2019-01-28: Quality; 2019-02-04: 21 Dec update: Current delivery date for replacement lot stays on 22-Feb-19. These parts are manufacture in Thermocool China but the coating is done in the US at Light Metals in Connecticut (Waters authorized supplier for this coating). To continue expedite. Last date WPPL can accept is on 28-Jan-19 - 13u of 700003816 and 6u of 186015007. 14 Dec update: Pending replacement from Thermo cool, LT 9-10 weeks. Next shipment impact: spares: 17 Dec (short of 13u out of 30u) instrument: 02 Jan 07 Dec update: Different surface finishing issue. 25pcs good parts to use after sorting. Next shipment impact will be on 17-Dec.; 2019-03-18: ASAP; 2019-03-25: Open
410003141 | Magnetic Target |  |  | 2018-12-31: Timken; 2019-01-07: 186015017 186015046 186015085 186015088; 2019-01-14: Jie Hui Wan Jing; 2019-01-28: Quality; 2019-02-04: 21 Dec update: Set screw used found to be wrong material - steel alloy instead of stainless steel material. Instruments shipped with wrong set screw are under stop ship and will be recall back from DC for rework. Venture is in production recovery for total 105 units to close Q4 shipment (54x 186015088, 47x 186015085 and 4x 186015046).; 2019-03-18: ASAP; 2019-03-25: Open
425000430 | GASKET, DEGASSER CHAMBER |  |  | 2018-12-31: Foamtec; 2019-01-07: 186015000; 2019-01-14: Jie Hui Wan Jing; 2019-01-28: Quality; 2019-02-04: 21 Dec update: Venture is facing reject during production (easily torn). Need to explore if Milford Engineering can add alternate material in the drawing specification - KT412 AD (existing material before ECO change). Only affecting last 9 units of 186015000.; 2019-03-18: ASAP; 2019-03-25: Close
Note on C14: sgblim: CH = Column Heater
Note on C15: sgblim: CH/C = Column Heater/Cooler
Note on C16: sgblim: ACQUITY Arc Bio 30cm Column Heater
Note on C17: sgblim: ACQUITY Arc Bio 30cm Col. Heater / Cooler
Note on G18: Jie Hui Ng: Thermo cool
Note on F19: Jie Hui Ng: Thermo Cool
Note on G19: Jie Hui Ng: Thermo Cool
Note on G20: Jie Hui Ng: Thermo cool
Note on H31: Jie Hui Ng: 16U ETD 18 Jan
Note on H32: Jie Hui Ng: 32U OK 16U ETD 18 Jan
Note on F37: Jie Hui Ng: 9u ETD 4 Jan 1u ETD TBA, rejected due to Fail function
Note on I39: Jie Hui Ng: 6u ETD 23-Jan 16u ETD 27-Feb

## CEI
[5 empty] | Shipped |  | OK |  | Potential delay, need to monitor |  |  |  | Delay, need immediate action
Updated On: | 2019-01-04 17:27 |  |  | 2019-01-01 |  |  |  | 2019-02-01 |  |  |  | 2019-03-01 |  |  |  |  | DO NOT CHANGE
|  | Q4 Forecast |  | WW01 | WW02 | WW03 | WW04 | WW05 | WW06 | WW07 | WW08 | WW09 | WW10 | WW11 | WW12 | WW13 | Type of Issue | Status
PN | Short Desc | Prod Plan | QTD Del'd | 2018-12-31 | 2019-01-07 | 2019-01-14 | 2019-01-21 | 2019-01-28 | 2019-02-04 | 2019-02-11 | 2019-02-18 | 2019-02-25 | 2019-03-04 | 2019-03-11 | 2019-03-18 | 2019-03-25 | Material | Open
186015014 | CPSO | 80 | 6 | 6 | 7 | 10 | 10 | 10 | 6 | 6 | 8 | 10 | 7 | 0 | 0 | 0 | Quality | Close
186015020 | ASO | 9 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 4 | 2 | 0 | 0 | 0 | 0 | 0 | Tester
186015021 | ASO | 3 | 3 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Manufacturing
Tooling
Part No | Part Description |  |  | Supplier | Instrument Affected | WPPL Team (Lead) | Milford Contact | Type of Issue | Brief Description of Issue | [5 empty] | Target Recovery Date | Status
400006352 400001357 | SUPPORT TRAY, QSM, CPSO PANEL, SIDE, ELEVATOR |  |  | 2018-12-31: TREND; 2019-01-07: 186015014; 2019-01-14: Jen; 2019-01-28: Quality; 2019-02-04: 15-Jan-2016 update : CEI revised their incoming inspection checklist to remove criteria for gloss since it's not a requirement for them to check. QA has approved. Closed 08-Jan-16 update : CEI is working with Trend to optimize their process to meet gloss requirements for CEI parts and share the supporting data. Jen to follow up with Waters Milford on TMD to cover gloss levels of 3.8 to 4.7 for powder coat batch # 3R15I0033. 31-Dec-15 update : 405002979 - Supplier Accurounds shipment delinquent to ETD US 29-Dec, ETA 04-Jan-2016. AWB# 7193146295. (orignal committed ETA 28-Dec-2015). 400006352 rejected due to Delta E* Out Of Spec, physical has already returned to Trend. Status TBA until Monday due to Trend shut down.; 2019-03-18: ASAP; 2019-03-25: Close
290001063 | BOX, SHIP, SAMPLE ORG, DUAL H/C ENGINE |  |  | 2018-12-31: J&J; 2019-01-14: Yin Ying; 2019-01-28: Material; 2019-02-04: 23-Sep-2016 update : Milford team will review the report follow by ECR submission. Target to close by end of OCT 2016. 15-July-2016 update : Evaluation report already sent to reliability group for review. Status : TBA 13-May-2016 update : Status remain no progress. 15-Apr-2016 update : Report submitted to Mike Cook & Randy Chin. Will follow up with reply return date. 01-Apr-2016 update : Report received on 31-Mar-2016. Review in progress. 24-Mar-2016 update : Drop test performed on 22-Mar, pending report compilation from J&J and CEI. 11-Mar-2016 update: Waters dwg updated with correct dimension - Approved FA based on dwg prelim Rev B. Ethafoam 600 is used instead of Ethafoam 450. CEI to finalize and submit plan for WPPL review before proceed with physical drop test. 04-Mar-2016 update: Jen and WJ went to CEI for FA sample buyoff on 03-Mar. Unable to buyoff due to wrong dimension on Waters dwg – YinYing to send revised prelim dwg to CEI. FAI report also require amendment - CEI to resubmit. 26-Feb-2016 update : FA sample has been reschedule to buy off on 02-Mar. 19-Feb-2016 update : FA sample to be ready on 29-Feb for buy off. 15-Jan-2016 : Has activated CEI to proceed with FA samples using localized proposed material Ethafoam 600. FA samples and drop test is schedule to be completed on 15-Feb-16.; 2019-03-18: Oct; 2019-03-25: Open
405002950 | SHAFT, GUIDE, A/S TRANSFER |  |  | 2018-12-31: Rei Promax Accurounds; 2019-01-07: 186015020; 2019-01-14: Bethany Lim Wan Jing; 2019-01-28: Material Quality; 2019-02-04: 22-July-2016 update : Localized supplier Rei Promax is having problem to support due to "straightness issue after heat treatment process" (1st lot) & "surface roughness"(2nd lot) causing 100% part rejected.Next shipment from Accurounds ETD on 29-July-2016. Has extended LOI with Accurounds to continue support. 15-July-2016 update : Localized supplier Rei Promax is having problem to support the material due to straightness issue and causing 100% part rejected. Next shipment from Accurounds will ETD on 29-July-2016. Buyer is expediting.; 2019-03-18: ASAP; 2019-03-25: Close
279001917 | ASSY, PANEL, PNEUMATICS, CPSO |  |  | 2018-12-31: AIRLINE HYDRAULICS CORP; 2019-01-07: 186015014; 2019-01-14: Bethany; 2019-01-28: Material; 2019-02-04: 30-Sep-2016 update : CEI to continue follow up delivery date from Airline to ensure no impact to Q4 production. 15pcs - ETD 10/03. 23-Sep-2016 update : 4pcs(reworked Qty) of 279001917 received on 19-Sep. Q3 shipment no impact. Next incoming 30pcs on 21-Sep & 15pcs on 23-Sep. Pull-in 10units from ww40 to ww39, CEI is reviewing on the capacity and will advise. 16-Sep-2016 update : 186015014 upside. Supplier airline imposed expediting cost of USD 2717.50 to pull-in the material. - 1 unit shipment affected on ww 39 is due to fail decay test. - CEI is still working with supplier on the delivery schedule.; 2019-03-18: Monitoring; 2019-03-25: Close
NA | NA |  |  | 2018-12-31: NA; 2019-01-07: 186015014; 2019-01-14: Wesley Desmond; 2019-01-21: NA; 2019-01-28: Tester; 2019-02-04: 24-Feb-2017 update : Per CEI feeback, they can meet 120unit per quarter (simulation floor space, headcout and tooling, will activate when require) 6-Feb-17 update: Production Meeting discussion - attach rate probably at 430 units per year. Operations team to work with CEI to simulate at a realistic volume (150 or 200 per quarter). 20-Jan-2017 update : Wesley and Desmond to advise the Production and Tester capacity for Insturment 186015014 if were to support for 300units per quarter.; 2019-03-25: Close
289007707 (289004744) | ASSY, SHELF, SAMPLE ORGANIZER |  |  | 2018-12-31: Technimetal; 2019-01-07: 186015014 186015020; 2019-01-21: Dave A Mark Loney; 2019-01-28: Material; 2019-02-04: 10-Mar-2017 update : Bearing (child part#289004744) from NBM managed to improved to 1000pcs ETD Japan 20-Mar-2017. CEI is following up with Technimetal to ensure shipment is deliver on time to support for April shipment. Closed. 03-Mar-2017 update : Impacting 10-Apr-2017 shipment and onwards. Technimetal has shortage of bearing (child part#289004744) from NBM. Current committed date from NMB (bearing supplier) as below:- 3000pcs ETD Japan 07-Apr-2017 4000pcs ETD Japan 22-May-2017 Currently Dave A and Mark Loney are helping to expedite bearing from NBM, however there is no improvement at the moment. Will continue follow up and expedite. CEI is checking with Technimetal the turn around time to assemble into 289007707 upon bearing receive.; 2019-03-18: ASAP; 2019-03-25: Close
210000420 441000878 | ASSY, PCB, 2200 ELEVATOR PERSONALITY ASSY, CABLE, ELEVATOR D.C. POWER |  |  | CEI | 186015014 186015020 700002606 | Wan Jing | Bob Lincoln Gene Tancrell | Quality | 30 June 2017 update : Both ECO9214851 and ECO9214563 to CEI and implemented. 09 June 2017 update: Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) - expiration will be extended to 23 Jun 2017 (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Current status: TMD is active. Will be in progress for all fresh builds until TMD expires or upon ECO release. ECO9214851 – released new board. ECO9214563 – released new cable. - ECO 9213776 hves been released to have both part# 210001161 & 441001987 structured into BOM and update ASPs. - Status for spare new P/N - ECR9215383,to release spare part# 700011556 (Kit, 2200 Elevator Personality PCB/Cable) & 700011557 (ASSY, CABLE, ELEVATOR 24VDC POWER) ECO released. 21-Apr-2017 update: Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) - expiration will be extended to 23 Jun 2017 (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Current status: TMD is active. Will be in progress for all fresh builds until TMD expires or upon ECO release. ECO9214851 – released new board. ECO9214563 – released new cable. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM and update ASPs. Current ECR status : Assigned - Status for spare new P/N TMD010308 - TMD has been closed. 07-Apr-2017 update: Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Current status: TMD is active. Will be in progress for all fresh builds until TMD expires or upon ECO release. ECO9214851 – released new board. ECO9214563 – released new cable. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM and update ASPs. - Status for spare new P/N TMD010308 For rework of units sent back to CEI from DC - TMD010308 (expire on 20-Jun-2017) (1)Remove shield only -WJ had buyoff CEI's rework instruction and first unit reworked. -Current status: TMD is active. -Total RMA is 101x RMA from ADC - 12x 20-Mar, 30x 22-Mar, 25x 23-Mar (Total 67x) RMA from Japan – 2x – 29-Mar (Total 2x) RMA from EDC - TBC (Total 32x) – Shipment delayed, scheduled to return to CEI on 10-Apr (Mon). -Follow up: (1) Rework for units from GDC to be confirmed 30-Mar-2017 update : Customer reported burned/melted power connector in several sample organizer instruments. TMD010283 (expire on 05-May-2017) (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Current status: TMD has been implemented. Will be in progress for all fresh builds until TMD expires or upon ECO release. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM. - Current spare P/N will not be reworked until spare new P/N is released. Randy Ryder will follow up on this ECR. TMD010308 (expire on 20-Jun-2017) FOr rework of units sent back to CEI from DC (1)Remove shield only -Current status: TMD has been implemented. -Total RMA is 101x RMA from ADC - 12x 20-Mar, 30x 22-Mar, 25x 23-Mar (Total 67x) RMA from Japan - 2x 29-Mar (Total 2x) RMA from EDC - TBC (Total 32x) (arriving Singapore mid next week) -Follow up: (1) Rework for units from GDC to be confirmed - probably rework in Milford. Randy to confirm. 24-Mar-2017 update : Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Current status: TMD has been implemented. Will be in progress for all fresh builds until TMD expires or upon ECO release. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM. - Status for spare new P/N TMD010308 FOr rework of units sent back to CEI from DC - TMD010308 (expire on 20-Jun-2017) (1)Remove shield only -WJ had buyoff CEI's reowkr instruction and first unit reworked. -Current status: TMD has been implemented. -Total RMA is 101x RMA from ADC - 12x 20-Mar, 30x 22-Mar, 25x 23-Mar (Total 67x) RMA from Japan - TBC (Total 2x) RMA from EDC - TBC (Total 32x) -FOllow up: (1) Rework for units from GDC to be confirmed 17-Mar-2017 update : Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -New cable FA has been approved. -WJ and Edwin had buyoff CEI's rework instruction. -Current status: TMD has been implemented. 6 units of CPSO had completed buyoff under TMD010283. WJ to release list of S/N of first 6 units of CPSO. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM. 10-Mar-2017 update : Customer reported burned/melted power connector in several sample organizer instruments. - TMD010283 (expire on 05-May-2017) (1) Rework existing PCB 210000420 into 210001161. (2) New PCB 210001161 must use with new cable 441001987. -Desmond and WJ to review CEI's TCN and rework instruction. -Current status: TMD has not been implemented. Implementation after cable 441001987 approval granted from Milford. - New cable 441001987, 2pcs FA sample will ETA Milford 13-Mar-2017, result should be available on 14-Mar-2017. - Has instructed CEI to proceed to build new PCBA 210001161. - Follow up to release ECR 9213776 to have both part# 210001161 & 441001987 structured into BOM. | [5 empty] | ASAP | Close
405002031 | BLOCK, CARRIAGE, LEADSCREW |  |  | 2018-12-31: Multisource; 2019-01-07: 186015020; 2019-01-14: Bethany/ Jie Hui; 2019-01-28: Material; 2019-02-04: 01 Dec 17 update: FA approved. 24 Nov 17 update: FA sample received, CEI to perform measurement and will arrange with WPPL QA for buyoff. 17 Nov 17 update: FA sample + production lot estimate to dock at CEI 20 Nov. CEI will arrange with Jen for buy off. Potential impact 2 units ASO. 10 Nov 17 update: Reschedule to 10 Nov to resubmit FA sample to Milford , affecting 2 units ASO.; 2019-03-18: Monitoring; 2019-03-25: Close
400004470 | PLATE, INNER, DOOR, SAMPLE ORGANIZER |  |  | 2018-12-31: Mechanical Plastics; 2019-01-07: 186015014; 2019-01-14: Bethany/ Jie Hui; 2019-01-28: Material; 2019-02-04: 19Jan 18 update: Supplier has shipped the material on 20-Jan(Waybill: 9387413991) 29 Dec 17 update: Supplier short of raw material. Impacting shipment on 22-Jan-2018.; 2019-03-18: Monitoring; 2019-03-25: Close
405002031 | BLOCK, CARRIAGE, LEADSCREW |  |  | 2018-12-31: Multisource; 2019-01-07: 186015020; 2019-01-14: Bethany/ Jie Hui; 2019-01-28: Material; 2019-02-04: 29 Dec 17 update: Impacting shipment on 12-Feb-2018 due to pull-in of Spare 700004468, shortage : 1pc; 2019-03-18: Monitoring; 2019-03-25: Close

## Fu Yu
|  |  |  | Potential delay, need to monitor |  |  |  | Delay, need immediate action |  |  |  |  | Shipped |  | OK
Updated On: | 2019-01-04 17:27 |  |  | 2019-01-01 |  |  |  | 2019-02-01 |  |  |  | 2019-03-01 |  |  |  |  | DO NOT CHANGE
Instrument |  | Q4 Forecast |  | WW01 | WW02 | WW03 | WW04 | WW05 | WW06 | WW07 | WW08 | WW09 | WW10 | WW11 | WW12 | WW13
PN | Short Desc | Prod Plan | QTD Del'd | 2018-12-31 | 2019-01-07 | 2019-01-14 | 2019-01-21 | 2019-01-28 | 2019-02-04 | 2019-02-11 | 2019-02-18 | 2019-02-25 | 2019-03-04 | 2019-03-11 | 2019-03-18 | 2019-03-25 | Type of Issue | Status
205001247 | Bottle tray | 1364 |  | 0 | 77 | 44 | 198 | 154 | 77 | 99 | 143 | 132 | 132 | 132 | 99 | 77 | Material | Open
WAT271082 | Bottle tray | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Quality | Close
205001139 | Bottle tray | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Tester
N2971044 | PL - 1139 BT | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Manufacturing
N2971045 | PL - 1247 BT | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Tooling
Part No | Part Description |  |  | Supplier | Spare Affected | OPS Team (Lead) | Milford Contact | Type of Issue | Brief Description of Issue | [5 empty] | Target Recovery Date | Status
405019835 | ADAPTER, BARBED, ELBOW |  |  | 2018-12-31: FUYU; 2019-01-07: 205001247; 2019-01-14: Bethany; 2019-01-28: Material; 2019-02-04: 15-Jan-2016 update : 405019835 to be received by FUYU on 20-Jan-2015. WW03 remaining 33pcs shipment to be shipped on WW04. 08-Jan-2016 update : ECO 9200214, Supplier Sample only ETA FUYU on 13-Jan-2105. Fuyu is checking with supplier whether they still have inventory to support at old revision before they can implement with new Rev. TBA; 2019-03-18: ASAP; 2019-03-25: Close
205001139 | BOTTLE TRAY REWORK |  |  | 2018-12-31: FUYU; 2019-01-14: Zack/Kevin; 2019-01-28: Material; 2019-02-04: 17-Mar-2017 Update: Damaged inventory scrapped in system in plant SG20. Pending issuance of claim settlement from Expeditors, Claim Ref: "Claim20170222 – 205001139 damaged in ADC" 24-Feb-2017 update : Claim issued to Expeditors for the full cost of the 14 damaged units. Pending credit from Expeditors before scrapping the units 17-Feb-2017 update : Provision for scrap - TBA. - 14 units damaged at expeditor due to mishandling. - FUYU quoted USD 950 to cover the cost incurred for re-inspection / re-box as well as left over balance of MOQ of carton box (336pcs) 20-Jan-2017 update : Provision for scrap - TBA. - 14 units damaged at expeditor due to mishandling. - Getting FUYU to check on the condition of the physical - Pending FUYU to advise the quotation for the rework and re-box of 14pcs x 205001139 base on MOQ of 350pcs for the carton box. 05-Aug-2016 update : Rework activity need to be further push out until further notice. 29-Jul-2016 update : Final Qty to be salvaged is 182pcs(13 pallets). Working with FUYU on when to start to perform rework activity. Total cost : USD 7672.22 (USD 900 (labor cost) + USD 6772.22 (Rework cost)) 08-Apr-2016 update : Dismantle and rework activity estimated to be performed by early July 2016. Final Qty to be finalised. 01-Apr-2016 update : Revised quotation as attached to be reviewed on 05-Apr-2016 04-Mar-2016 update : Revised quotation as attached. 19-Feb-2016 update : Getting FUYU to re-quote for rework cost. (Dismantle the bottle tray). TBA 05-Feb-2016 update : Quotation obtain as attached 29-Jan-2016 update : Pending Fuyu to provide dismantle quotation and new quote for Part# 205001139. To explore feasibility to rework 205001139 to 205001247. Total 14 purchased parts involved. 4 parts - common to 205001247. 3 parts - unique, need to scrap. 7 parts - common to Milford or other CMs. Important note: The rework require major effort (strip / dismantle the bottle tray till individual piece part level) - may impose high rework cost and possess risk in part damage as well.; 2019-03-18: TBA; 2019-03-25: Close
400006520 400006447 | ASSY, BOTTLE TRAY SUPPORT/STIFFENER BRKT, TUBING GUIDE HOLDER |  |  | 2018-12-31: Paradigm; 2019-01-07: 205001247; 2019-01-14: Jen; 2019-01-28: Quality; 2019-02-04: 23 Mar 18 update : No Impact to shipment. Defect :40006520 & 400006447 found finishing paint surface rough/ contaminated/ pitted. Fuyu is working with Paradigm to sort part based on inspection at 30 inches for SPECK per spec WAT022557.; 2019-03-18: ASAP; 2019-03-25: Close
400006520 | Assy, Bottle Tray Support/Stiff |  |  | 2018-12-31: Paradigm; 2019-01-07: 205001247; 2019-01-14: Bethany; 2019-02-04: 13 Apr 18 update : Paradigm receive few batches(0n 02/14) of Sherwin William powder(PN: ULM2-90001) which is in harden condition. Estimated the new batch of powder will be reaching Paradigm on 17-Apr. Fuyu to continue expedite and follow up.; 2019-03-18: ASAP; 2019-03-25: Close
Note on D6: Jie Hui Ng: Quarterly amt

## Summary
Update as of WW01 2019:
Q1 2019 CTB Status
1 | Flextronics production CTB for Q1 2019 :
Material Shortage
210000520 ASSY, PCB, 3VCCPU AHCT
04 Jan update: Q4 upside missing compnent shortage. Under weekly allocation. 225pcs ETA 07 Jan 225pcs ETA 09 Jan Next shipment impact 23 Jan
Quality
279003850 ASSY, SOLENOID VALVE, 3-WAY
Cosmetic reject not in accordance to drawing. Pending UAI.
Tester
210000414 ASSY,PCB,ACQ,TUV,PERSONALITY & 210001164 ASSY, PCB HAMAMATSU 518 PIX PD
GTS failed self-test after installing the calibrated NI cards; Milford advised to run diagnostics to clean the relay resistance. WPPL tried to run cleaning by remote into GTS, but still not able to pass.
2 | Venture production CTB for Q1 2019 :
Material Shortage
NA
Quality
405004811 Tray Pre-Heater Four Column H/C
Pending replacement from Thermo cool, LT 9-10 weeks. Next shipment impact: Spares: 17 Dec (short of 13u out of 30u) Instrument: 02 Jan
3 | CEI production CTB for Q1 2019 :
Material Shortage
NA
Quality
NA
4 | Fu Yu production CTB for Q1 2019 :
Material Shortage
NA
Quality
NA

## QSM recovery
c
2998/PDA/eLambda Detectors Recovery Plan
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [12 empty] | Monthly - Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual => | [17 empty] | audit | audit | [10 empty] | 6-Nov: Deepavali
Models | Monthly | Oct | Nov | Dec | Weekly | Open + | [11 empty] | Monthly | Oct | Nov | Dec |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-10 | 2018-10-17 | 2018-10-24 | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | [75 empty] | 2018-12-17 | 2018-12-18 | 2018-12-19 | 2018-12-20 | 2018-12-21 | 2018-12-22 |  | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29 | time axis 2018-10-08 to 2018-12-15
[93 empty] | Dec | Dec | Dec | Dec | Dec | Dec |  | Dec | Dec | Dec | Dec | Dec | Dec | time axis 2018-10-08 to 2018-12-15
186015080 | Total Reqmt | 23 | 176 | 108 | Weekly Reqmt | 0 | 0 | 0 | 23 | 64 | 48 | 0 | 64 | 8 | 36 | 64 | 0 | Total Reqmt | 23 | 176 | 108 | Plan | 2018-10-13: 0; 2018-10-14: 0; 2018-10-31: 15; 2018-11-07: 40; 2018-11-14: 50; 2018-11-21: 26; 2018-11-28: 56; 2018-12-05: 16; 2018-12-12: 4 |  |  |  | 58 | [6 empty] | 42 |  |  |  | 307
| Total Plan | 0 | 187 | 120 | Total Plan | 0 | 0 | 0 | 15 | 40 | 50 | 26 | 56 | 16 | 4 | 58 | 42 | Total Act Till Date | 15 | 188 | 12 | Actual | 2018-10-31: 15; 2018-11-03: 18; 2018-11-05: 12; 2018-11-07: 10; 2018-11-09: 10; 2018-11-12: 18; 2018-11-14: 22; 2018-11-16: 9; 2018-11-19: 4; 2018-11-21: 13; 2018-11-23: 10; 2018-11-26: 15; 2018-11-28: 31; 2018-11-30: 16; 2018-12-07: 4; 2018-12-14: 8 |  | 10 |  | 40 |  | 16 | [8 empty] | 281
| Delta | -23 | -12 | 0 | Delta | 0 | 0 | 0 | -8 | -32 | -30 | -4 | -12 | -4 | -36 | -42 | 0 | Balance to go | -8 | 4 | -92 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 |  | 30 | 30 | 0 | 0 | 10 | 10 |  | 28 | 28 | 0 | 0 | 9 | 9 |  | 13 | 13 | 0 | 0 | 10 | 10 |  | 25 | 25 | 0 | 0 | 16 | 16 |  | 16 | 16 | 0 | 0 | 4 | 4 |  | 4 | 4 | 0 | 0 | 8 | 8 |  | 18 | 18 | 0 | 0 | 16 | 16 |  | 16 | 16 | -26 | -26 | -26 | -26
186015018 | Total Reqmt | 32 | 44 | 67 | Weekly Reqmt | 0 | 0 | 0 | 32 | 0 | 0 | 44 | 0 | 39 | 28 | 0 | 0 | Total Reqmt | 32 | 44 | 67 | Plan | 2018-10-13: 0; 2018-10-14: 0; 2018-10-21: 0; 2018-11-21: 24; 2018-12-05: 41; 2018-12-12: 66 |  |  |  | 12 | [6 empty] | 0 |  |  |  | 143
| Total Plan | 0 | 24 | 119 | Total Plan | 0 | 0 | 0 | 0 | 0 | 0 | 24 | 0 | 41 | 66 | 12 | 0 | Total Act Till Date | 0 | 24 | 119 | Actual | 2018-11-21: 14; 2018-11-23: 10; 2018-12-03: 4; 2018-12-05: 30; 2018-12-07: 16; 2018-12-10: 20; 2018-12-12: 37; 2018-12-14: 12 | [14 empty] | 143
| Delta | -32 | -52 | 0 | Delta | 0 | 0 | 0 | -32 | -32 | -32 | -52 | -52 | -50 | -12 | 0 | 0 | Balance to go | -32 | -52 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | -10 | -10 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 4 | 4 | -7 | -7 | 9 | 9 |  | 29 | 29 | 0 | 0 | 12 | 12 |  | 12 | 12 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
186015001 | Total Reqmt | 0 | 15 | 13 | Weekly Reqmt | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 0 | 13 | 0 | 0 | 0 |  |  |  |  | Plan | 2018-11-07: 4; 2018-11-21: 11; 2018-12-05: 13
| Total Plan | 0 | 15 | 13 | Total Plan | 0 | 0 | 0 | 0 | 4 | 0 | 11 | 0 | 13 | 0 | 0 | 0 |  |  |  |  | Actual | 2018-11-05: 4; 2018-11-16: 2; 2018-11-19: 4; 2018-11-21: 5; 2018-11-30: 5; 2018-12-03: 2; 2018-12-05: 6
| Delta | 0 | 0 | 0 | Delta | 0 | 0 | 0 | 0 | 4 | 4 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  |  | Delta | [21 empty] | 0 | 0 | 0 | 0 | 0 | 0 |  | 4 | 4 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 2 | 2 |  | 6 | 6 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 5 | 5 |  | 7 | 7 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [12 empty] | Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual ==>
Recovery Plan | Monthly | Oct | Nov | Dec | Weekly | [12 empty] | Monthly | Apr | May | Jun |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-10 | 2018-10-17 | 2018-10-24 | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | [5 empty] | 2018-10-08 | 2018-10-09 | 2018-10-10 | 2018-10-11 | 2018-10-12 | 2018-10-13 | 2017-04-02 | 2018-10-15 | 2018-10-16 | 2018-10-17 | 2018-10-18 | 2018-10-19 | 2018-10-20 | 2017-04-09 | 2018-10-22 | 2018-10-23 | 2018-10-24 | 2018-10-25 | 2018-10-26 | 2018-10-27 | 2017-04-16 | 2018-10-29 | 2018-10-30 | 2018-10-31 | 2018-11-01 | 2018-11-02 | 2018-11-03 | 2017-04-23 | 2018-11-05 | 2018-11-06 | 2018-11-07 | 2018-11-08 | 2018-11-09 | 2018-11-10 | 2017-04-30 | 2018-11-12 | 2018-11-13 | 2018-11-14 | 2018-11-15 | 2018-11-16 | 2018-11-17 | 2017-04-30 | 2018-11-19 | 2018-11-20 | 2018-11-21 | 2018-11-22 | 2018-11-23 | 2018-11-24 | 2017-04-30 | 2018-11-26 | 2018-11-27 | 2018-11-28 | 2018-11-29 | 2018-11-30 | 2018-12-01 | 2017-04-30 | 2018-12-03 | 2018-12-04 | 2018-12-05 | 2018-12-06 | 2018-12-07 | 2018-12-08 | 2017-04-30 | 2018-12-10 | 2018-12-11 | 2018-12-12 | 2018-12-13 | 2018-12-14 | 2018-12-15 |  | 2018-12-17 | 2018-12-18 | 2018-12-19 | 2018-12-20 | 2018-12-21 | 2018-12-22 |  | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29
All | Total Reqmt | 55 | 235 | 188 | Weekly Reqmt | 0 | 0 | 0 | 55 | 64 | 48 | 59 | 64 | 60 | 64 | 64 | 0 | Total Reqmt | 55 | 235 | 188 | Plan | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 15 | 0 | 0 | 0 |  | 0 | 0 | 44 | 0 | 0 | 0 |  | 0 | 0 | 50 | 0 | 0 | 0 |  | 0 | 0 | 61 | 0 | 0 | 0 |  | 0 | 0 | 56 | 0 | 0 | 0 |  | 0 | 0 | 70 | 0 | 0 | 0 |  | 0 | 0 | 70 | 0 | 0 | 0 |  | 0 | 0 | 70 | 0 | 0 | 0 |  | 0 | 0 | 42 | 0 | 0 | 0 | 478
| Total Plan | 0 | 226 | 252 | Total Plan | 0 | 0 | 0 | 15 | 44 | 50 | 61 | 56 | 70 | 70 | 70 | 42 | Total Act Till Date | 15 | 232 | 0 | Actual | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 15 | 0 | 0 | 18 |  | 16 | 0 | 10 | 0 | 10 | 0 |  | 18 | 0 | 22 | 0 | 11 | 0 |  | 8 | 0 | 32 | 0 | 20 | 0 |  | 15 | 0 | 31 | 0 | 21 | 0 |  | 6 | 0 | 36 | 0 | 20 | 0 |  | 20 | 0 | 37 | 0 | 20 | 0 |  | 10 | 0 | 40 | 0 | 16 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 452
| Delta | -55 | -64 | 0 | Delta | 0 | 0 | 0 | -40 | -60 | -58 | -56 | -64 | -54 | -48 | -42 | 0 | Balance to go | -40 | -43 | -231 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 |  | 34 | 34 | 0 | 0 | 10 | 10 |  | 28 | 28 | 0 | 0 | 11 | 11 |  | 19 | 19 | -10 | -10 | 10 | 10 |  | 25 | 25 | 0 | 0 | 21 | 21 |  | 27 | 27 | -7 | -7 | 13 | 13 |  | 33 | 33 | 0 | 0 | 20 | 20 |  | 30 | 30 | 0 | 0 | 16 | 16 |  | 16 | 16 | -26 | -26 | -26 | -26
WATH-279001810-L GPV receipt | as of 10/24, short of total 36sets for 31-Oct PO. Next 100pcs (50sets) ETA on 10/30. Each Tue expected 50sets |  |  |  | ETA date |  |  | 2018-10-30 | 2018-11-08 | 2018-11-15 | 2018-11-16 | 2018-11-22
[5 empty] | Qty |  |  | 100 | 100 | 100 | 75 | 150
[5 empty] | Cum shipped to Waters |  |  | 409
[5 empty] | Bal at Flex (sets) |  |  | -146.5
Formulas:
M = [S]+2
T = [M]+1
Note on BB8: 15u passed & packed. 1u failed, troubleshooting. The rest of 3ets are either reject or on-hold due by QA due to not match set during TMD implementation.
Note on BP8: 11/08: 10u pushed to wk 11/21 due to GPV received late.
Note on CK13: - some headcount on medical leave.
Note on BI16: 4sets have materials.
Note on N28: afternoon time.
Note on O28: after lunch time.
Note on P28: Amy Xu Xia Jie: Lunch time
Note on Q28: Amy Xu Xia 150pcs received on 22-Nov

## Detector recovery-Waters
c
2998/PDA/eLambda Detectors Recovery Plan
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [9 empty] | Monthly - Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual ==> | [8 empty] | Holiday | [14 empty] | 184 |  |  |  |  | 0 |  | 266 | [6 empty] | 231 | [6 empty] | 188 | [6 empty] | 227 | [5 empty] | Holiday | 211 |  |  |  | 681
Models | Monthly | Oct | Nov | Dec | Weekly | Open + | [8 empty] | Monthly | Oct | Nov | Dec |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | time axis 2018-10-29 to 2018-12-29
 | time axis 2018-10-29 to 2018-12-29
186002998 | Total Reqmt | 0 | 24 | 33 | Weekly Reqmt | 0 | 6 | 6 | 6 | 6 | 0 | 6 | 16 | 11 | Total Reqmt | 0 | 24 | 33 | Plan | 2018-11-28: 12; 2018-12-05: 6; 2018-12-12: 12; 2018-12-19: 16; 2018-12-21: 0; 2018-12-22: 0; 2018-12-23: 0; 2018-12-24: 0; 2018-12-25: 0; 2018-12-26: 11; 2018-12-27: 0; 2018-12-28: 0; 2018-12-29: 0 | 100% drift test
| Total Plan | 0 | 12 | 45 | Total Plan | 0 | 0 | 0 | 0 | 12 | 6 | 12 | 16 | 11 | Total Act Till Date | 0 | 12 | 28 | Actual | 2018-11-23: 6; 2018-11-26: 6; 2018-11-30: 6; 2018-12-12: 12; 2018-12-19: 16
| Delta | 0 | -12 | 0 | Delta | 0 | -6 | -12 | -18 | -12 | -6 | 0 | 0 | 0 | Balance to go | 0 | -12 | -17 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 6 | 6 |  | 12 | 12 | 0 | 0 | 6 | 6 |  | 6 | 6 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | -11 | -11 | -11 | -11
186299800 | Total Reqmt | 0 | 190 | 210 | Weekly Reqmt | 0 | 40 | 50 | 50 | 50 | 50 | 55 | 55 | 50 | Total Reqmt | 0 | 190 | 210 | Plan | 2018-11-07: 12; 2018-11-21: 67; 2018-11-28: 99; 2018-12-05: 62; 2018-12-12: 55; 2018-12-19: 43; 2018-12-26: 62; 2018-12-29: 0 | 100% drift test | testers are not stable (lost comm)
| Total Plan | 0 | 178 | 222 | Total Plan | 0 | 12 | 0 | 67 | 99 | 62 | 55 | 43 | 62 | Total Act Till Date | 0 | 178 | 152 | Actual | 2018-11-05: 12; 2018-11-19: 24; 2018-11-21: 43; 2018-11-23: 8; 2018-11-26: 40; 2018-11-28: 51; 2018-11-30: 24; 2018-12-03: 8; 2018-12-05: 30; 2018-12-07: 24; 2018-12-10: 16; 2018-12-12: 15; 2018-12-14: 8; 2018-12-17: 24; 2018-12-19: 11; 2018-12-21: 16
| Delta | 0 | -12 | 0 | Delta | 0 | -28 | -78 | -61 | -12 | 0 | 0 | -12 | 0 | Balance to go | 0 | -12 | -70 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 24 | 24 | 0 | 0 | 8 | 8 |  | 48 | 48 | 0 | 0 | 24 | 24 |  | 32 | 32 | 0 | 0 | 24 | 24 |  | 40 | 40 | 0 | 0 | 8 | 8 |  | 32 | 32 | 0 | 0 | 16 | 16 |  | 16 | 16 | -46 | -46 | -46 | -46
186017003 | Total Reqmt | 16 | 48 | 22 | Weekly Reqmt | 16 | 16 | 0 | 16 | 16 | 16 | 6 | 0 | 0 | Total Reqmt | 16 | 48 | 22 | Plan | 2018-11-12: 0; 2018-11-21: 0; 2018-11-28: 16; 2018-12-05: 16; 2018-12-12: 16; 2018-12-19: 14; 2018-12-26: 24 | 100% drift test | testers are not stable
| Total Plan | 0 | 16 | 70 | Total Plan | 0 | 0 | 0 | 0 | 16 | 16 | 16 | 14 | 24 | Total Act Till Date | 0 | 16 | 46 | Actual | 2018-11-23: 8; 2018-11-26: 8; 2018-12-03: 8; 2018-12-05: 8; 2018-12-12: 16; 2018-12-19: 14
| Delta | -16 | -48 | 0 | Delta | -16 | -32 | -32 | -48 | -48 | -48 | -38 | -24 | 0 | Balance to go | -16 | -48 | -24 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 8 | 8 |  | 16 | 16 | 0 | 0 | 0 | 0 |  | 8 | 8 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | -24 | -24 | -24 | -24
186017013 | Total Reqmt | 4 | 7 | 0 | Weekly Reqmt | 4 | 4 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | Total Reqmt | 4 | 7 | 0 | Plan | 2018-11-07: 4; 2018-11-12: 0; 2018-12-19: 7; 2018-12-29: 0 | 100% drift test | testers are not stable
| Total Plan | 0 | 4 | 7 | Total Plan | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | Total Act Till Date | 0 | 4 | 7 | Actual | 2018-11-09: 4; 2018-12-19: 7
| Delta | -4 | -7 | 0 | Delta | -4 | -4 | -4 | -7 | -7 | -7 | -7 | 0 | 0 | Balance to go | -4 | -7 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -4 | -4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
186248900 | Total Reqmt | 0 | 338 | 392 | Weekly Reqmt | 0 | 64 | 80 | 102 | 92 | 96 | 104 | 111 | 81 | Total Reqmt | 0 | 338 | 392 | Plan | 2018-11-07: 59; 2018-11-21: 117; 2018-11-28: 131; 2018-12-05: 131; 2018-12-12: 86; 2018-12-19: 123; 2018-12-21: 0; 2018-12-22: 0; 2018-12-23: 0; 2018-12-26: 83; 2018-12-29: 0 | sampling drift test
| Total Plan | 0 | 307 | 423 | Total Plan | 0 | 59 | 0 | 117 | 131 | 131 | 86 | 123 | 83 | Total Act Till Date | 0 | 307 | 324 | Actual | 2018-11-02: 24; 2018-11-05: 24; 2018-11-07: 11; 2018-11-19: 32; 2018-11-21: 85; 2018-11-23: 32; 2018-11-26: 32; 2018-11-28: 67; 2018-11-30: 32; 2018-12-03: 32; 2018-12-05: 67; 2018-12-07: 8; 2018-12-10: 24; 2018-12-12: 54; 2018-12-14: 24; 2018-12-17: 40; 2018-12-19: 59; 2018-12-21: 16
| Delta | 0 | -31 | 0 | Delta | 0 | -5 | -85 | -70 | -31 | 4 | -14 | -2 | 0 | Balance to go | 0 | -31 | -99 | Delta | 0 | 0 | 0 | 0 | 24 | 24 | 24 | 48 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 32 | 32 | 0 | 0 | 32 | 32 |  | 64 | 64 | 0 | 0 | 32 | 32 |  | 64 | 64 | 0 | 0 | 8 | 8 |  | 32 | 32 | 0 | 0 | 24 | 24 |  | 64 | 64 | 0 | 0 | 16 | 16 |  | 16 | 16 | -67 | -67 | -67 | -67
186017002 | Total Reqmt | 16 | 44 | 38 | Weekly Reqmt | 16 | 16 | 0 | 16 | 12 | 12 | 8 | 8 | 10 | Total Reqmt | 16 | 44 | 38 | Plan | 2018-11-21: 0; 2018-11-28: 8; 2018-12-05: 16; 2018-12-12: 19; 2018-12-19: 24; 2018-12-21: 0; 2018-12-22: 0; 2018-12-23: 0; 2018-12-26: 31; 2018-12-27: 0; 2018-12-28: 0; 2018-12-29: 0 | 100% drift test
| Total Plan | 0 | 8 | 90 | Total Plan | 0 | 0 | 0 | 0 | 8 | 16 | 19 | 24 | 31 | Total Act Till Date | 0 | 8 | 51 | Actual | 2018-11-23: 8; 2018-11-30: 8; 2018-12-03: 7; 2018-12-05: 1; 2018-12-12: 19; 2018-12-17: 16; 2018-12-19: 8
| Delta | -16 | -52 | 0 | Delta | -16 | -32 | -32 | -48 | -52 | -48 | -37 | -21 | 0 | Balance to go | -16 | -52 | -39 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 8 | 8 |  | 8 | 8 | 0 | 0 | 8 | 8 |  | 15 | 15 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 16 | 16 | 0 | 0 | 0 | 0 |  | 0 | 0 | -31 | -31 | -31 | -31
186002487 | Total Reqmt | 0 | 44 | 61 | Weekly Reqmt | 0 | 12 | 0 | 6 | 26 | 18 | 31 | 12 | 0 | Total Reqmt | 0 | 44 | 61 | Plan | 2018-10-31: 0; 2018-11-07: 9; 2018-11-14: 3; 2018-11-16: 6; 2018-11-21: 44; 2018-12-12: 29; 2018-12-19: 0; 2018-12-21: 0; 2018-12-22: 0; 2018-12-23: 0; 2018-12-24: 0; 2018-12-26: 26; 2018-12-28: 0; 2018-12-29: 0 | sampling drift test
| Total Plan | 0 | 62 | 55 | Total Plan | 0 | 9 | 3 | 50 | 0 | 0 | 29 | 0 | 26 | Total Act Till Date | 0 | 62 | 43 | Actual | 2018-11-07: 9; 2018-11-09: 3; 2018-11-16: 18; 2018-11-19: 24; 2018-11-21: 8; 2018-12-07: 12; 2018-12-10: 12; 2018-12-12: 5; 2018-12-21: 14
| Delta | 0 | 18 | 12 | Delta | 0 | -3 | 0 | 44 | 18 | 0 | -2 | -14 | 12 | Balance to go | 0 | 18 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 3 | 3 | 3 | 3 | 0 | 0 | 12 | 12 |  | 36 | 36 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 12 | 12 |  | 24 | 24 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 14 | 14 |  | 14 | 14 | -12 | -12 | -12 | -12
186015031 | Total Reqmt | 4 | 72 | 74 | Weekly Reqmt | 4 | 20 | 16 | 20 | 16 | 20 | 22 | 16 | 16 | Total Reqmt | 4 | 72 | 74 | Plan | 2018-11-07: 24; 2018-11-14: 46; 2018-11-19: 0; 2018-11-21: 26; 2018-12-12: 22; 2018-12-19: 16; 2018-12-22: 0; 2018-12-23: 0; 2018-12-24: 0; 2018-12-25: 0; 2018-12-26: 16; 2018-12-27: 0; 2018-12-28: 0; 2018-12-29: 0 | 100% drift test
| Total Plan | 0 | 96 | 54 | Total Plan | 0 | 24 | 46 | 26 | 0 | 0 | 22 | 16 | 16 | Total Act Till Date | 0 | 96 | 42 | Actual | 2018-11-09: 22; 2018-11-12: 18; 2018-11-14: 30; 2018-11-16: 16; 2018-11-19: 10; 2018-11-30: 4; 2018-12-07: 16; 2018-12-10: 2; 2018-12-14: 8; 2018-12-17: 8; 2018-12-21: 8
| Delta | -4 | 20 | 0 | Delta | -4 | 0 | 30 | 36 | 20 | 0 | 0 | 0 | 0 | Balance to go | -4 | 20 | -12 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -24 | -24 | -2 | -2 | -2 | 16 | 16 | 0 | 0 | 16 | 16 |  | 26 | 26 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 4 | 4 |  | 4 | 4 | 4 | 4 | 20 | 20 |  | 22 | 22 | 0 | 0 | 8 | 8 |  | 16 | 16 | 0 | 0 | 8 | 8 |  | 8 | 8 | -8 | -8 | -8 | -8
186015032 | Total Reqmt | 8 | 72 | 72 | Weekly Reqmt | 8 | 24 | 8 | 16 | 24 | 20 | 20 | 16 | 16 | Total Reqmt | 8 | 72 | 72 | Plan | 2018-11-07: 32; 2018-11-14: 16; 2018-11-21: 21; 2018-11-28: 3; 2018-12-05: 28; 2018-12-12: 20; 2018-12-19: 16; 2018-12-26: 16; 2018-12-29: 0 | 100% drift test | testers are not stable (lost comm)
| Total Plan | 0 | 72 | 80 | Total Plan | 0 | 32 | 16 | 21 | 3 | 28 | 20 | 16 | 16 | Total Act Till Date | 0 | 72 | 56 | Actual | 2018-11-09: 24; 2018-11-14: 24; 2018-11-16: 8; 2018-11-19: 8; 2018-11-21: 5; 2018-11-28: 3; 2018-11-30: 8; 2018-12-03: 8; 2018-12-05: 12; 2018-12-10: 8; 2018-12-12: 12; 2018-12-14: 8; 2018-12-19: 8
| Delta | -8 | -8 | 0 | Delta | -8 | 0 | 8 | 13 | -8 | 0 | 0 | 0 | 0 | Balance to go | -8 | -8 | -24 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -32 | -32 | -8 | -8 | -8 | -8 | -8 | 0 | 0 | 8 | 8 |  | 16 | 16 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 8 | 8 |  | 16 | 16 | 0 | 0 | 0 | 0 |  | 8 | 8 | 0 | 0 | 8 | 8 |  | 8 | 8 | 0 | 0 | 0 | 0 |  | 0 | 0 | -16 | -16 | -16 | -16
186015033 | Total Reqmt | 16 | 74 | 64 | Weekly Reqmt | 16 | 16 | 24 | 18 | 16 | 24 | 16 | 16 | 8 | Total Reqmt | 16 | 74 | 64 | Plan | 2018-11-07: 32; 2018-11-14: 40; 2018-11-21: 18; 2018-11-28: 8; 2018-12-05: 16; 2018-12-12: 16; 2018-12-19: 16; 2018-12-26: 8; 2018-12-28: 0; 2018-12-29: 0 | 100% drift test | testers are not stable (lost comm)
| Total Plan | 0 | 98 | 56 | Total Plan | 0 | 32 | 40 | 18 | 8 | 16 | 16 | 16 | 8 | Total Act Till Date | 0 | 90 | 48 | Actual | 2018-11-09: 40; 2018-11-12: 16; 2018-11-14: 16; 2018-11-19: 8; 2018-11-21: 10; 2018-11-30: 8; 2018-12-05: 16; 2018-12-10: 16; 2018-12-17: 8; 2018-12-19: 8
| Delta | -16 | 8 | 0 | Delta | -16 | 0 | 16 | 16 | 8 | 0 | 0 | 0 | 0 | Balance to go | -16 | 0 | -16 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -32 | -32 | 8 | 8 | 8 | 24 | 24 | 0 | 0 | 0 | 0 |  | 8 | 8 | 0 | 0 | 0 | 0 |  | 0 | 0 | -8 | -8 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 16 | 16 | 0 | 0 | 0 | 0 |  | 8 | 8 | 0 | 0 | 0 | 0 |  | 0 | 0 | -8 | -8 | -8 | -8
186015034 | Total Reqmt | 0 | 4 | 15 | Weekly Reqmt | 0 | 0 | 0 | 4 | 0 | 5 | 5 | 5 | 0 | Total Reqmt | 0 | 4 | 15 | Plan | 2018-11-21: 4; 2018-12-05: 5; 2018-12-12: 5; 2018-12-19: 5 | 100% drift test
| Total Plan | 0 | 4 | 15 | Total Plan | 0 | 0 | 0 | 4 | 0 | 5 | 5 | 5 | 0 | Total Act Till Date | 0 | 4 | 15 | Actual | 2018-11-21: 4; 2018-12-05: 5; 2018-12-10: 5; 2018-12-19: 4; 2018-12-21: 1
| Delta | 0 | 0 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Balance to go | 0 | 0 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 5 | 0 | 0 | 0 | 0 |  | 0 | 0 | -1 | -1 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
186015053 | Total Reqmt | 0 | 7 | 0 | Weekly Reqmt | 0 | 0 | 0 | 4 | 3 | 0 | 0 | 0 | 0 | Total Reqmt | 0 | 7 | 0 | Plan | 2018-11-09: 0; 2018-11-21: 4; 2018-11-28: 3 | 100% drift test
| Total Plan | 0 | 7 | 0 | Total Plan | 0 | 0 | 0 | 4 | 3 | 0 | 0 | 0 | 0 | Total Act Till Date | 0 | 7 | 0 | Actual | 2018-11-21: 4; 2018-11-28: 3
| Delta | 0 | 0 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Balance to go | 0 | 0 | 0 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [9 empty] | Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual ==>
Recovery Plan | Monthly | Apr | May | Jun | Weekly | [9 empty] | Monthly | Oct | Nov | Dec |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | [5 empty] | 2018-10-29 | 2018-10-30 | 2018-10-31 | 2018-11-01 | 2018-11-02 | 2018-11-03 | 2017-04-09 | 2018-11-05 | 2018-11-06 | 2018-11-07 | 2018-11-08 | 2018-11-09 | 2018-11-10 | 2017-04-16 | 2018-11-12 | 2018-11-13 | 2018-11-14 | 2018-11-15 | 2018-11-16 | 2018-11-17 | 2017-04-23 | 2018-11-19 | 2018-11-20 | 2018-11-21 | 2018-11-22 | 2018-11-23 | 2018-11-24 | 2017-04-30 | 2018-11-26 | 2018-11-27 | 2018-11-28 | 2018-11-29 | 2018-11-30 | 2018-12-01 | 2017-04-30 | 2018-12-03 | 2018-12-04 | 2018-12-05 | 2018-12-06 | 2018-12-07 | 2018-12-08 | 2017-04-30 | 2018-12-10 | 2018-12-11 | 2018-12-12 | 2018-12-13 | 2018-12-14 | 2018-12-15 | 2017-04-30 | 2018-12-17 | 2018-12-18 | 2018-12-19 | 2018-12-20 | 2018-12-21 | 2018-12-22 | 2017-04-30 | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29
All | Total Reqmt | 64 | 924 | 981 | Weekly Reqmt | 64 | 218 | 184 | 261 | 261 | 261 | 273 | 255 | 192 | Total Reqmt | 64 | 924 | 981 | Plan | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 172 | 0 | 0 | 0 | 0 | 0 | 0 | 105 | 0 | 6 | 0 | 0 | 0 | 0 | 301 | 0 | 0 | 0 | 0 | 0 | 0 | 280 | 0 | 0 | 0 | 0 | 0 | 0 | 280 | 0 | 0 | 0 | 0 | 0 | 0 | 280 | 0 | 0 | 0 | 0 | 0 | 0 | 280 | 0 | 0 | 0 | 0 | 0 | 0 | 277 | 0 | 0 | 0
| Total Plan | 0 | 864 | 1117 | Total Plan | 0 | 172 | 105 | 307 | 280 | 280 | 280 | 280 | 277 | Total Act Till Date | 0 | 856 | 812 | Actual | 0 | 0 | 0 | 0 | 24 | 0 | 0 | 36 | 0 | 20 | 0 | 93 | 0 | 0 | 34 | 0 | 70 | 0 | 42 | 0 | 0 | 106 | 0 | 159 | 0 | 62 | 0 | 0 | 86 | 0 | 124 | 0 | 90 | 0 | 0 | 63 | 0 | 139 | 0 | 60 | 0 | 0 | 83 | 0 | 133 | 0 | 48 | 0 | 0 | 96 | 0 | 135 | 0 | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Delta | -64 | -124 | 12 | Delta | -64 | -110 | -189 | -143 | -124 | -105 | -98 | -73 | 12 | Balance to go | -64 | -132 | -301 | Delta | 0 | 0 | 0 | 0 | 24 | 24 | 24 | 60 | 60 | -92 | -92 | 1 | 1 | 1 | 35 | 35 | 0 | 0 | 36 | 36 |  | 142 | 142 | 0 | 0 | 62 | 62 |  | 148 | 148 | -8 | -8 | 82 | 82 |  | 145 | 145 | 4 | 4 | 64 | 64 |  | 147 | 147 | 0 | 0 | 48 | 48 |  | 144 | 144 | -1 | -1 | 54 | 54 |  | 54 | 54 | -223 | -223 | -223 | -223
PDA Family | [5 empty] | 44 | 106 | 88 | 117 | 115 | 115 | 108 | 108 | 85 | [7 empty] | 0 | [6 empty] | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 114 | 0 | 0 | 0 | 0 | 0 | 0 | 141 | 0 | 0 | 0 | 0 | 0 | 0 | 133 | 0 | 0 | 0 | 0 | 0 | 0 | 124 | 0 | 0 | 0 | 0 | 0 | 0 | 117 | 0 | 0 | 0 | 0 | 0 | 0 | 121
TUV family | [5 empty] | 20 | 112 | 96 | 144 | 146 | 146 | 165 | 147 | 107 | [7 empty] | 0 | [6 empty] | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 49 | 0 | 6 | 0 | 0 | 0 | 0 | 187 | 0 | 0 | 0 | 0 | 0 | 0 | 139 | 0 | 0 | 0 | 0 | 0 | 0 | 147 | 0 | 0 | 0 | 0 | 0 | 0 | 156 | 0 | 0 | 0 | 0 | 0 | 0 | 163 | 0 | 0 | 0 | 0 | 0 | 0 | 156
Critical part: | Receive date | 2018-11-15 | 2018-11-19 | 2018-11-26
WAT-352000115: 300 + 326pcs received. | Qty | 300 | 326 | 200
OH stock covers until 11/29 build. Short of 55u for 5-Dec recovery. | Ship qty | 1128
200pcs of alternate part cover until 12/05 build. | Balance qty | -302
Next lot ETD Mexico on 12/05, ETA 12/10.
Note on BL4: short of 55pcs of Switch. To be convered by alternate part.
Note on AS6: built Qty based on 13-Nov CTB.
Note on AX6: including TUV, PDA, eLambda pulled in CTB
Note on BZ12: utilize capacity to build 5u for optic bench TMD - urgent evaluation.
Note on BS16: 16u wait for next lot of Sensata Switch or use next 300pcs of alternate part. ETA 7-Dec PM time.
Note on AL21: impacted by MRPC fitting quality issue.
Note on BS24: 8u wait for next lot of Sensata Switch or use next 300pcs of alternate PN, ETA 7-Dec PM time. 12u adjusted to 19-Dec as 2487 has OH materials to build
Note on BZ24: 12u from wk 12-Dec to 19-Dec after swap with 2487.
Note on BS28: original 24u, revised to 19u as 5u pushed out to give way to 2487. 19u wait for next lot of Sensata Switch or use next 300pcs of alternate PN, ETA 7-Dec PM time.
Note on N32: including 19u pulled in frm Jan.
Note on AQ32: FA control build that needed Waters SQE sign off first.
Note on BS32: 12u + 5u pulled in frm 12-Dec & 2-Jan, CTB on-hand.
Note on CG32: 14u pulled in frm Jan 2-Jan to 26-Dec. 12u pulled in frm 16-Jan to 2-Jan then to 31-Dec.
Note on AQ36: additional 30sets CTB.
Note on AL37: impact from MPRC fitting issue.
Note on BG37: from WIP
Note on AQ40: additional 8sets CTB.
Note on AL41: impact from MRPC fitting issue.
Note on AQ44: additional 16sets CTB.
Note on BE44: Amy Xu Xia Jie: 8u pull in from 5 Dec to 28 Nov for tester validation . Push out 8u 5032 from 28 Nov to 5 Dec
Note on AL45: impacted by MRPC fitting quality issue.
Note on BE45: 8u assy completed since 11/26. Waiting for Eric's return to office on 11/28 to put to testers.
Note on BZ50: 1u not yet pass drift test at final test.
Note on K59: Amy Xu Xia Jie: 340 -Internal
Note on L59: Amy Xu Xia Jie: Internal weekly 300
Note on C64: Amy Xu Xia Jie: Receive at lunch time
Note on D64: Amy Xu Xia Jie: Receive at lunch time
Note on E64: Amy Xu Xia Jie: from CMTi/Waters under TMD

## Alliance recovery
c
2998/PDA/eLambda Detectors Recovery Plan
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [12 empty] | Monthly - Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual => | [17 empty] | audit | audit | [10 empty] | 6-Nov: Deepavali
Models | Monthly | Oct | Nov | Dec | Weekly | Open + | [11 empty] | Monthly | Oct | Nov | Dec |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-10 | 2018-10-17 | 2018-10-24 | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | [75 empty] | 2018-12-17 | 2018-12-18 | 2018-12-19 | 2018-12-20 | 2018-12-21 | 2018-12-22 |  | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29 | time axis 2018-10-08 to 2018-12-15
[93 empty] | Dec | Dec | Dec | Dec | Dec | Dec |  | Dec | Dec | Dec | Dec | Dec | Dec | time axis 2018-10-08 to 2018-12-15
186269505 | Total Reqmt | 40 | 40 | 160 | Weekly Reqmt | 0 | 0 | 0 | 40 | 0 | 0 | 40 | 0 | 40 | 40 | 40 | 40 | Total Reqmt | 40 | 40 | 160 | Plan | 2018-10-13: 0; 2018-10-14: 0; 2018-10-31: 2; 2018-11-14: 38; 2018-11-21: 40; 2018-12-05: 40; 2018-12-12: 40 |  |  |  | 40 | [6 empty] | 90
| Total Plan | 0 | 80 | 210 | Total Plan | 0 | 0 | 0 | 2 | 0 | 38 | 40 | 0 | 40 | 40 | 40 | 90 | Total Act Till Date | 2 | 71 | 87 | Actual | 2018-10-31: 2; 2018-11-14: 30; 2018-11-16: 21; 2018-11-19: 15; 2018-11-21: 3; 2018-11-30: 2; 2018-12-03: 11; 2018-12-05: 36; 2018-12-12: 31; 2018-12-14: 9 |  |  |  | 35 |  | 35
| Delta | -40 | 0 | 50 | Delta | 0 | 0 | 0 | -38 | -38 | 0 | 0 | 0 | 0 | 0 | 0 | 50 | Balance to go | -38 | -7 | -80 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | -8 | -8 | 13 | 13 |  | 28 | 28 | -9 | -9 | -9 | -9 |  | -9 | -9 | -9 | -9 | -7 | -7 |  | 4 | 4 | 0 | 0 | 0 | 0 |  | 0 | 0 | -9 | -9 | 0 | 0 |  | 0 | 0 | -5 | -5 | 30 | 30 |  | 30 | 30 | -60 | -60 | -60 | -60
186269506 | Total Reqmt | 198 | 440 | 331 | Weekly Reqmt | 0 | 0 | 118 | 80 | 120 | 100 | 90 | 130 | 90 | 90 | 90 | 61 | Total Reqmt | 198 | 440 | 331 | Plan | 2018-10-13: 0; 2018-10-14: 0; 2018-10-21: 0; 2018-10-24: 20; 2018-10-31: 138; 2018-11-07: 120; 2018-11-14: 112; 2018-11-21: 118; 2018-11-28: 130; 2018-12-05: 90; 2018-12-12: 90 |  |  |  | 90 | [6 empty] | 11
| Total Plan | 20 | 618 | 281 | Total Plan | 0 | 0 | 20 | 138 | 120 | 112 | 118 | 130 | 90 | 90 | 90 | 11 | Total Act Till Date | 158 | 496 | 194 | Actual | 2018-10-24: 20; 2018-10-26: 43; 2018-10-29: 35; 2018-10-31: 60; 2018-11-03: 35; 2018-11-05: 35; 2018-11-07: 32; 2018-11-09: 45; 2018-11-12: 40; 2018-11-14: 45; 2018-11-19: 15; 2018-11-21: 66; 2018-11-23: 40; 2018-11-26: 40; 2018-11-28: 61; 2018-11-30: 42; 2018-12-03: 31; 2018-12-05: 33; 2018-12-07: 20; 2018-12-10: 20; 2018-12-12: 50; 2018-12-14: 40 |  | 34 |  | 25 |  | 2
| Delta | -178 | 0 | -50 | Delta | 0 | 0 | -98 | -40 | -40 | -28 | 0 | 0 | 0 | 0 | 0 | -50 | Balance to go | -40 | 16 | -121 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 43 | 43 | 78 | 78 | 0 | 0 | 0 | 35 |  | 70 | 70 | -18 | -18 | 27 | 27 |  | 67 | 67 | 0 | 0 | 0 | 0 |  | 15 | 15 | -37 | -37 | 3 | 3 |  | 43 | 43 | -26 | -26 | 16 | 16 |  | 47 | 47 | -10 | -10 | 10 | 10 |  | 30 | 30 | -10 | -10 | 30 | 30 |  | 64 | 64 | -1 | -1 | 1 | 1 |  | 1 | 1 | -10 | -10 | -10 | -10
| Monthly - Total Reqmt Vs Planned Recovery |  |  |  | Weekly - Total Reqmt Vs Planned Recovery | [12 empty] | Total Reqmt Vs Actual |  |  |  | Daily Recovery Plan Vs Actual ==>
Recovery Plan | Monthly | Oct | Nov | Dec | Weekly | [12 empty] | Monthly | Apr | May | Jun |  | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S | S | M | T | W | T | F | S
[6 empty] | 2018-10-10 | 2018-10-17 | 2018-10-24 | 2018-10-31 | 2018-11-07 | 2018-11-14 | 2018-11-21 | 2018-11-28 | 2018-12-05 | 2018-12-12 | 2018-12-19 | 2018-12-26 | [5 empty] | 2018-10-08 | 2018-10-09 | 2018-10-10 | 2018-10-11 | 2018-10-12 | 2018-10-13 | 2017-04-02 | 2018-10-15 | 2018-10-16 | 2018-10-17 | 2018-10-18 | 2018-10-19 | 2018-10-20 | 2017-04-09 | 2018-10-22 | 2018-10-23 | 2018-10-24 | 2018-10-25 | 2018-10-26 | 2018-10-27 | 2017-04-16 | 2018-10-29 | 2018-10-30 | 2018-10-31 | 2018-11-01 | 2018-11-02 | 2018-11-03 | 2017-04-23 | 2018-11-05 | 2018-11-06 | 2018-11-07 | 2018-11-08 | 2018-11-09 | 2018-11-10 | 2017-04-30 | 2018-11-12 | 2018-11-13 | 2018-11-14 | 2018-11-15 | 2018-11-16 | 2018-11-17 | 2017-04-30 | 2018-11-19 | 2018-11-20 | 2018-11-21 | 2018-11-22 | 2018-11-23 | 2018-11-24 | 2017-04-30 | 2018-11-26 | 2018-11-27 | 2018-11-28 | 2018-11-29 | 2018-11-30 | 2018-12-01 | 2017-04-30 | 2018-12-03 | 2018-12-04 | 2018-12-05 | 2018-12-06 | 2018-12-07 | 2018-12-08 | 2017-04-30 | 2018-12-10 | 2018-12-11 | 2018-12-12 | 2018-12-13 | 2018-12-14 | 2018-12-15 |  | 2018-12-17 | 2018-12-18 | 2018-12-19 | 2018-12-20 | 2018-12-21 | 2018-12-22 |  | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29
All | Total Reqmt | 238 | 480 | 491 | Weekly Reqmt | 0 | 0 | 118 | 120 | 120 | 100 | 130 | 130 | 130 | 130 | 130 | 101 | Total Reqmt | 238 | 480 | 491 | Plan | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 20 | 0 | 0 | 0 |  | 0 | 0 | 140 | 0 | 0 | 0 |  | 0 | 0 | 120 | 0 | 0 | 0 |  | 0 | 0 | 150 | 0 | 0 | 0 |  | 0 | 0 | 158 | 0 | 0 | 0 |  | 0 | 0 | 130 | 0 | 0 | 0 |  | 0 | 0 | 130 | 0 | 0 | 0 |  | 0 | 0 | 130 | 0 | 0 | 0 |  | 0 | 0 | 130 | 0 | 0 | 0 |  | 0 | 0 | 101 | 0 | 0 | 0
| Total Plan | 20 | 698 | 491 | Total Plan | 0 | 0 | 20 | 140 | 120 | 150 | 158 | 130 | 130 | 130 | 130 | 101 | Total Act Till Date | 160 | 567 | 0 | Actual | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 20 | 0 | 43 | 0 |  | 35 | 0 | 62 | 0 | 0 | 35 |  | 35 | 0 | 32 | 0 | 45 | 0 |  | 40 | 0 | 75 | 0 | 21 | 0 |  | 30 | 0 | 69 | 0 | 40 | 0 |  | 40 | 0 | 61 | 0 | 44 | 0 |  | 42 | 0 | 69 | 0 | 20 | 0 |  | 20 | 0 | 81 | 0 | 49 | 0 |  | 34 | 0 | 60 | 0 | 37 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
| Delta | -218 | 0 | 0 | Delta | 0 | 0 | -98 | -78 | -78 | -28 | 0 | 0 | 0 | 0 | 0 | 0 | Balance to go | -78 | 9 | -482 | Delta | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 43 | 43 | 78 | 78 | 0 | 0 | 0 | 35 |  | 70 | 70 | -18 | -18 | 27 | 27 |  | 67 | 67 | -8 | -8 | 13 | 13 |  | 43 | 43 | -46 | -46 | -6 | -6 |  | 34 | 34 | -35 | -35 | 9 | 9 |  | 51 | 51 | -10 | -10 | 10 | 10 |  | 30 | 30 | -19 | -19 | 30 | 30 |  | 64 | 64 | -6 | -6 | 31 | 31 |  | 31 | 31 | -70 | -70 | -70 | -70
| 23-Oct status | [21 empty] | W: will recover 23u; W: will recover 23u
short 38pcs for 31-Oct PO
short 47pcs for 31-Oct PO
Formulas:
M = [S]+2
T = [M]+1
Note on DF8: 14-Dec: Waters swaps 50u frm 506 to 505.
Note on BR9: Amy Xu Xia Jie: 9 on hold due to pump tray test failed. Pump tray test low yield start from 14 Nov
Note on BW9: Amy Xu Xia Jie: 9u rework planned to 28th Nov
Note on CB9: waiting for Waters to sign off rework instruction.
Note on CD9: pending Rework Instruction approval by WPPL Eng & QA.
Note on CF9: remaining 7u rework in progress after finalized rework instruction and training done.
Note on CR9: 9 units have not passed final test as of 10AM count. Impact from carry over failures.
Note on CY9: estimate abt 15u would need time until 21-Dec due to some HC on sick leave. Managed to reduce gap to 5u.
Note on BP10: 8u not yet passed final test. To be recovered on 16-Nov.
Note on DF12: swap with 505.
Note on BW13: Amy Xu Xia Jie: 37pcs miss due to pump tray test low yeild with DJ instrument part
Note on CD13: 3u cancelled due to Box supplier did not deliver shipping box and did not alert Flex.
Note on CK13: - some headcount on medical leave. - slow down of build than expected with handling of 9u of 186269505 in rework.
Note on CR13: impact from carry over failures.
Note on BI14: 18u to be recovered on 9-Nov shipment: still in testing due to no HC willing to OT during holiday period.
Note on CY14: 1u needs rework.
Note on CD26: including 9u of 505

## FTN Recovery Plan (2)
FTN Model | 2018-12-19 | 2018-12-26 | 2018-01-04 | 2018-01-09 | Total | ADC | Total |  |  | GES FG Qty to rework
5088 | 6 | 12 | 20 | 48 | 86 | 48 | 134 |  |  | Fresh Build
5085 | 15 | 32 | 36 | 20 | 103 | 46 | 149 |  |  | ADC Qty to rework/ship
5046 | 4 |  |  |  | 4 | 4 | 8 |  |  | Ship plan
5017 |  |  |  |  | 0 | 40 | 40 |  |  | Actual shipped
| Tester |  | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 7 | 7 | 7 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
| Day |  | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat
FTN Model | Date | Total QTY | 2018-12-20 | 2018-12-21 | 2018-12-22 | 2018-12-23 | 2018-12-24 | 2018-12-25 | 2018-12-26 | 2018-12-27 | 2018-12-28 | 2018-12-29 | 2018-12-30 | 2018-12-31 | 2019-01-01 | 2019-01-02 | 2019-01-03 | 2019-01-04 | 2019-01-05 | 2019-01-06 | 2019-01-07 | 2019-01-08 | 2019-01-09 | 2019-01-10 | 2019-01-11 | 2019-01-12
5088 | Fresh Build | 80 | 2018-12-20: 18; 2018-12-21: 18; 2018-12-22: 18; 2018-12-26: 4; 2018-12-27: 18; 2018-12-28: 4
| Rework | 54 | 2018-12-31: 6; 2019-01-07: 15; 2019-01-08: 15; 2019-01-09: 15; 2019-01-10: 3
| Ship Plan/Qty | 137 | 2018-12-21: 18; 2018-12-24: 39; 2019-01-04: 32; 2019-01-11: 48
| Actual Shipped | 15 | 2018-12-21: 15
| Balance to ship |  | 134 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119 | 119
5085 | Fresh Build | 56 | 2018-12-23: 14; 2018-12-24: 18; 2018-12-25: 14; 2018-12-26: 10
| Rework | 93 | 2018-12-28: 14; 2018-12-29: 18; 2018-12-30: 12; 2018-12-31: 3; 2019-01-03: 4; 2019-01-04: 15; 2019-01-05: 15; 2019-01-06: 12
| Ship Plan/Qty | 149 | 2018-12-26: 32; 2018-12-28: 15; 2019-01-04: 36; 2019-01-07: 46; 2019-01-09: 20
| Actual Shipped | 0
| Balance to ship |  | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149 | 149
5046 | Fresh Build | 0
| Rework | 8 | 2018-12-26: 4; 2019-01-03: 4
| Ship Plan/Qty | 8 | 2018-12-28: 4; 2019-01-04: 4
| Actual Shipped | 0
| Balance to ship |  | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8
5017 | Fresh Build | 0
| Rework | 40 | 2018-12-31: 6; 2019-01-01: 12; 2019-01-02: 15; 2019-01-03: 7
| Ship Plan/Qty | 40 | 2019-01-04: 40
| Actual Shipped | 0
| Balance to ship |  | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40 | 40
| Total Output/Day |  | 18 | 18 | 18 | 14 | 18 | 14 | 18 | 18 | 18 | 18 | 12 | 15 | 12 | 15 | 15 | 15 | 15 | 12 | 15 | 15 | 15 | 3 | 0 | 0
Note on E15: Celeste De Villa: 3U failed in FT. Will ship on 24/Dec

## Foamtec FA (hidden)
| 2017-10-01 | 2017-11-01 | 2017-12-01 | 2018-01-01 | 2018-02-01 | 2018-03-01
Venture Total: 43 parts | Buffer stock |  |  | Until March 2018
Site audit Done in Sept
FA submission Timeline |  | 22 | 21
Flex Total: 42 parts | Buffer stock |  |  | Until June 2018
Site audit 03.11.2017
FA submission Timeline |  |  | 21 | 19 | 2
CEI Total: 18 parts | Buffer stock |  |  | Until June 2018
Site audit Done on 10.10.2017
FA submission Timeline |  |  | 12 | 6
Foamtec shut down end Dec 2017

## FA Tracker (hidden)
No | CM | Supplier | Description | Status | CM FA submission date | WPPL FA buy off plan | To do/Owner | Target Phase-in date | Action item
1 | Flex/Venture/CEI | Foamtec | Foamtec will be relocating their office to Woodlands area and the manufacturing operation to Johor Bahru, Malaysia | Closed | Refer to Foamtec FA tab |  | Wan Jing - Venture Eugene - Flex Jen - CEI | 2018-03-01 | ECO 9222207 been released to 3 CMs.
2 | Venture | Baiteng | Sheetmetal - Total 29 parts Changed from supplier Metal component to Baiteng - CR | Closed | 2017-08-01 | 2017-09-01 | Wan Jing | Q1 2018 | 1. All 28 FA parts from BT at various dates/lots - Total 28 parts FA approved and implemented in Q1&Q2 2018. - 1 parts(400000938) - exploring new supplier JinPao
3 | Venture/Flex | Baiteng | Manufacturing Relocation Total 168 parts | Closed | TBA | TBA | Wan Jing - Venture Eugene - Flex | TBA | Baiteng has confirmed the Manufacturing Relocation plan will be pushing out to year 2020.
4 | Venture | B&E | 3 Optic benches Qualifying new raw cast supplier from TPI to China Synergy - CR, eliminate Quality and logistic issue | Closed | 2017-08-01 | 2017-08-10 | Wan Jing | 2017-11-01
5 | Flex | B&E | Launching new tool for child part# WAT207031(Child part for WAT207032) - Tool replacement | Closed | 2017-12-01 | 2017-12-01 | Flex SQE | Q2 2018 | 1. New tooling from B&E is approved for WAT207031 2. Flex is checking the ETA with B&E on the first lot using the new tool, TBA.
6 | Venture/Flex | PARKER HANNIFIN | Manufacturing Relocation | Closed |  |  | Milford QA - Rob Z | Q4 2017 | Parker has started to ship material from new plant.
7 | CEI | Multisource (Haus Specialty) | Manufacturing Relocation | Closed |  |  | Milford - Site buy off (John/Tom) WPPL FA - Jen | 2017-10-01 | 405002031 - FA approved, release to production.
8 | Flex | ETLA (Frencken Group Limited) | Change from JEP to ETLA (New supplier) - CR, eliminate Quality issue | Closed | 2017-09-01 | 2017-09-01 | Flex/Eugene 1st pass FA - Eugene site audit - Jen | Q3 18 | 1. Site audit - WPPL QA (Jen) --> Done 2. 405002568 - FA approved. Pending for final Quotation confirmation to ETLA. 3. 405006788 - FA approved 4. 405003295 - FA approved
9 | Flex | JEP | Optic bench - 405002568 - Faclity relocation | Closed | TBA | TBA | Flex/Eugene | TBA | ECO to match drawing therefore no need to raise TMD. (Void)
10 | Flex | Trend | Changed of child part WAT060671 (childpart for WAT057996) from A&one to Yangbum - CR | Closed | 2017-08-01 | 2017-08-01 | Eugene | 2017-11-01
11 | Flex | Flex Senai | PCBA conversion non-clean to Clean process | Closed | TBA | TBA | Shaynee | Q4 2017 | Kallang 1st shipment date Group 1 - FA reviewed and approved 210000411- 13 Dec, 210000416 - 15 Nov, 210000425 - 22 Nov, 210000384 - 15 Nov Group 2 - FA reviewed and approved 210000355 - 6 Dec, 210000444 - 13 Dec, 210000429 - 6 Dec, 210000428 - 6 Dec Group 3 - FA reviewed and approved 210000477 - 29 Nov, 210000404 - 29 Nov, 210000414 - 20 Dec, 210000415 - 6 Dec Group 4 - FA reviewed and approved 210000562 - 27 Dec, 210000345 - 27 Dec, 210000396 - 20 Dec, 210000417 - 20 Dec 210000421 - 7 Mar. Kallang is currently working with Senai on the spare demand from Milford, and will provide the revised implementation date with scrap quantity again