## New Pump Tray CT (hidden)
New Pump Tray Tester Cycle Time
Item | Description | Detail | CT (mins)
1 | Handling Time | Pick-up from trolley | 0.3
|  | Wiring\Tubing to tester | 2
|  | Serial Number - Manual Priming | 2
2 | Auto Time | Priming the Pump Mechanic | 3.58
Ramp and Decay
|  | Accumulator | 3.8
|  | Primary | 3.8
|  | Ripple \ Flow Test | 6.155
3 | Handling Time | Unmount the Pump Tray
|  | Unplug cable from tester | 2
|  | Print Report \ Signoff | 1.2
|  | Transfer to trolley | 0.3
| Total |  | 25.135

## Minutes Of Meeting-CLOSED (hidden)
ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | ASP/TSP Matrix Development - REFER HISTORY from 20th Jun 2018 email 30 May 2018 - To update and maintain the tracker to be on current state (6 Jun 2018 onwards) 11 Jul 2018 - target ASP BSM\QSM update by 31 July 2018. 15 Aug 2018: Pending ECR/ECO . 17 OCt 2018 : 1 ECO submitted and rest are pending/ to trigger Gary once approve 16 Jan 2019: Details are updated at "ASP TSP Summary" | Gary/Yee Hwa | 2017-07-06 | TBD | Closed | Low
7 | Implentation for Alliance Template 1.9 - Hua Zhen to verify the chemical part number as per WAT271129 Rev ? - after resolving the Carryover Test (High Blank) - 18 Jul 2018 24 Oct 2018 ONHOLD ACTIVITY 16 Jan 2019: No Template 1.9 will be impletemented. New testers are expected to be send to flex - TBA | Hua Zhen | 2018-06-20 | 6/27/2018 7/11/2018 12/30/2018 | Closed | Low
8 | Consigned Equipment Health Review Identify the list of consigned equipment.Total 56 equipment. Identify the downtime data on the consigned equipment Based on Jan-Jun data. Establish a checklist to review the health of the equipment.Avaialable Start the review of the health of the equipment. Review the status of the health with Waters team. Take actions and close on the problematic testers. Continue monitor for effectiveness - Dashbord update in progress for Alliance - 18 Jul 2018 - Alliance Dashboard : completed for submissiont 01 Aug 2018 (time frame CEHR) - Pump Dashboard completed - submitted on 31 Aug 2018 (resend the original file) - 17 Oct 2018 Identify and try to source locally available parts (PUMPS Testers) - 21 Nov 2018 Prioritize PUMP Tester (WPPL to source for PCBA and Pump Control Module). Target 07 Dec 2018 28 Nov 2018: Joel to follow up with a OESR (Pump) so that WPPL can work on the sourcing. 05 Dec 2018: Joel will send out the OESR by today 19 Dec 2019: To send Detector Health review 16 Jan 2019: Will be review through Project meeting | TAN CK/JOEL | 2018-07-07 | 2018-08-31 | Closed | Medium
9 | eDHR into Engg discussion Kick off meeting and align with CFT team on proposal Work with IT on configuration setup Matrix preparation Proposal and change request to customer Verify in test environment Release to production continue monitor and update 16 Jan 2019: Will be review through Project meeting | TAN CK | 2018-07-07 | 2018-11-30 | Closed | Low
10 | Monitor chemical for Carryover Test as per Rev 1.8 - 19 Sep 18 Hua Zhen to follow-up with Milford the implementation of Alliance Template 1.9 - Flex to continue the tracking and record potential issue with new Seal Pack, no details from Eng'g to share but higher suspect. Once design review finish with new seal pack and may have FW upgrade later. If FT performance did well in Flex, Engg will accelerate the upgrade process. - 17 Oct 2018 - To trigger 2 months ahead / During weekly Eng'g meeting and through email - 24 Oct 2018 refer to "CONSUMABLE TAB" - 21 Nov 2018 - To follow up with MCotter for chemical - Yee Hwa - 28 Nov 2018 - Yee Hwa will follow up to get additional 16 boxes that can last till early Jan19 - 05 Dec 2018: Joel to check and update on the qty. - 12 Dec 2018: return 1 box YH 16 Jan 2019: To review based on sheet "Consumables" | JOEL MD | 2018-07-25 | TBD | Closed | Medium
11 | Consigned Health review on Tooling Identify the list of tooling (approx. 275) from Su Fern Priority by family (alliance, pumps, BSM and Detectors) 2 weeks / family - 19 Sep 2018 Alliance looks OK, Test Equipment will be to prioritized - 26 Sep 2018 Onhold until tester - 09 Oct 2018 to discuss further the bandwidth? - 17 Oct 2018 - may need to push thrugh because of Oct'18 Audit - 31 Oct 2018 - resume for tooling with plan and GANTT Chart - 28 Nov 2018 - Discussed on the GANTT Chart - 12 Nov 2018 Item 1 completed (refer to CJFT Tab) - 19 Dec 2018 To provide details CK Tan as per CJFT Tab " Analyze downtime data " 16 Jan 2019: Will be review through Project meeting | Gary /CK | 8/15/2018 01 / 11/ 18 | 2018-12-31 | Closed | Medium
16 | Pump eEvent Log 05 Dec 2018: To generate a configuration file for Pump family 10 Jan 2019 : Awaiting review from YH / file submitted 27 Dec 2018 16 Jan 2019: Will be review through Project meeting | CK/ Joel | 2018-12-05 | 2018-12-31 | Closed | Medium
17 | CPU 2000 – 210000520 - REFER HISTORY from 20th Jun 2018 email - 210000520 to try on other model and pass is shippable. (YH) - to prepare a plan for 7 pcba return to Flex 05 Dec 2018: Yee Hwa to review with Melody to close this action item 19 Dec 2018 To prepare slide comparission for Year to Year CPU 2000 Performance 16 Jan 2019: Joel to consolidate the performance and close the monitoring YYYY - 2016 43/12604 (0.3411%) YYYY - 2017 32/12591 (0.2541%) YYYY - 2018 20/12315 (0.1624%) as of Nov 2018 (Dec 2018 Not Available) | Team | 2017-07-27 | 7/27/2018 8/31/2018 12/31/2018 02/28/2019 | Close | Medium

## Flex
ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY | WPPL Discussion
1 | Detector PCBA Fixture Qualification Plan:TMD execution, Fixture registration, ASP update 16 Jan 2019: Yee Hwa to trigger TMD for eval run by 25Jan19 27 Feb2019: Yee Hwa to trigger TMD for eval run by 01 Mar 19 05 Apr 2019: TMD 010527v00 start implentation date - Gary to send tracker (excel spread sheet) and schedule 24 Apr 2019: Gary to share the SN List of UUT 21 May 2019 : Serial Number to publish by Gary 12 Jun 019 : TMD 010527v00 evaluation Closed | GLY SFC | 2019-01-16 | 30 Jun 19 | Close | Medium
2 | BSM/QSM PCBA Fixture qualification - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 05-Feb-2020: PCBA Solidworks drawing to be forward to WPPL (GLY to check with Keypoint and advise date). 12-Feb-2020 GYC received the drawing from Keypoint 06-May-2020 GYC completed review for 22 drawings (BSM PCBA jig drawings), send all feed back to Keypoint | GLY GYC | 2019-06-12 | 2020-05-13 | Close | Low
3 | Weekly Downtime Update: Flex to the detail for down equipment every Thursday 5:30PM -Monthly - every end of the month: Dashboard / end of the month -to add summary page on the downtime report / weekly 24 Apr 2019: To include tools/equipment 11Dec2019: Target Closure change from Routine to 24-Apr-19 | JOEL MD | 2018-06-20 | 2019-04-24 | Routine | Low
4 | HDD Back Up/Imaging of all CPU as per PM Checklist: History: 2TB External HDD/Acronis imaging s/w provided by WPPL to backup all consigned testers. Objective :Flex to check with internal IT to HDD Back Up/Imaging of all CPU - In progress - 09 Oct 2018 - Raise Request for 3 Qoutation for 2 HDD - 17 Oct 2018 to discuss further as Flex not allowed to have external HDD inside facility - 24 Oct 2018 WPPL will consign HDD with memo to backup/imaging. Cloud service can also be considered. - 16 Jan 2019: Joel to follow up with Hua Zhen on this. - 14 Feb 2019: Refer to attached excel file) - 27 Feb 2019 To sched one unit a day and plan and create a separate tab for imaging - 12 Jun 2019 After ECO/Validation/Update - shall update the Back-up HDD accordingly - 19 Jun 2019 To backup TE82969-1 and all TE82882 Stations after validation To backup TE82943-11 to 18 after validation - 26 Jun 2019 Perform imaging periodically (6 months as a guideline and if there's ECO \ Firmware update. (JMD) 11Dec2019: Target Closure change from Routine to 26-June-19 and Status change from Onhold to Routine. To submit lastest backup/ imaging list and plan or schedule. David to discuss with Joel. - 17 Jul 2020 To share the schedule add here in from 9/19/2018 restart to 15-Jul-20 - 12-Aug-20 - to share schedule to complete - 19-Aug-20 - Refer to attched file 26-Aug-20 - JMD shared the schedule. - Feedback from HZL received last 2019. - To review together the guideline JMD and HZL. 09-Sep-2020 Refer extenal attachment for Backup schedule and completed the backup | Joel | 2020-07-15 | 6/26/2019 based on Schedule | Open | Low
5 | AIR COMPRESSOR requirement: History: Found filter inside Pump tray tester with oil/Water Objective : To provide Oil/Water free compressed air to Waters testers Refer to last Jan 2019 Updates. 12 Feb 2019 - Measurement of Outlet / Inlet of Compress Dry Air at Pump Tray Area (suspected to be the worst condition as last position of the air supply) 27 Feb 2019 To install air filter after recieved the report and monitor the performance, with date of installation and PM. 10 Apr 19 - to identify filter to suit the requirement 21 May 2019 - Equipment Calibration by vendor 29 May 2019 - Target to execute by 03 Jun 19 12 Jun 2019 - Measurement completed 03 Jun 2019 and report to share by 26 Jun 2019 (JMD) 26 Jun 2019 - Tan CK and Vendor in progress of verification of test result (Tan CK) 03 Jul 2019 - Oil and dew pass in requirement but fail in particulate count / get the raw data and recalculate (TCK) 10 Jul 2019 - Flex Facilities will install SMC filter which will control at less than 100 particles of 0.3um at the outlet and re-conduct a third party verification again. 17 Jul 2019 - Third party verfication is completed dated 11Jul19, pending for report. 24 Jul 2019 - As committed by vendor, report will send to flex by Thursday (25Jul19) - CK 31 Jul 2019 - Used filter SMC , Model : AMF350C-N04B-R (cost S$270 each) showing positive result and able to meet Waters specification. 14 Aug 2019 - To verify if filter in place (Joel) 21 Aug 2019 - varified that mainline is using SMC , Model : AMF350C-N04B-R and test result shows all within Waters specification. Refer to attached test result | Joel / Hua Zhen / CK Tan | 2018-09-19 | 31-Dec-2018 26-Jun-2019 03-Jul-2019 26-Jul-2019 31-Jul-2019 21-Aug-2019 | Close | Low
6 | Commissioning of New Seal Pack Jig Plan:TMD execution, Fixture registration, ASP update 16 Jan 2019: Yee Hwa to trigger TMD for eval run 27 Feb 2019 : TMD 010639 has been approved at Waters, release to Flex 28-Feb : Expedite return of fixture from supplier. 01-Mar : DE routed for approval in Flex System. 05-Mar : Received the fixture from supplier. 05-Mar : PM Documentation complete. Training completed. 06-Mar : Implemented TMD010639 v00 P/N: 186269506 : SN C19SM4920A P/N: 186269505 : SN C19SM7166A 17 Apr : WPPL send the drawing to Milfrod for update 24 Apr : Gary to send PACK and Go Package in Solidworks 21 May 2019 : Completed and close | Gary | 2019-01-16 | 2019-04-30 | Close | High
7 | Waters Server installation 28Nov18: Currently 2x servers in rack. Recieving 3rd server. - Collection of 2 server on 03Dec18 from WPPL. (Kallang and Senai) - 1 server to setup at Kallang on 04Dec18. - 1 server to setup at Senai on 05Dec18. - WPPL plan to decommission Kallang old Waters server on Jan19 05Dec18: Will update the status by this week - New Server (ST: FVDCBS2) installed 04 Dec 2018 - utilized existing UPS and to ship new UPS back to WPPL as SPARE - Request to SCRAP the OLD server and defective UPS 09Sep19: JMD send follow-up Whats app message to HZL for update from Waters IT Team 18Sep19: HZL advised to witness the decomissioning of the old sever by 4:00 PM 25Sep19: Pls. refer to previous minutes for the deleted discussions: Server is working and no issue observed for the whole week (Laber printing and test reports) HZL to check on scrapping the Server through local IT Support 27Sep19: HZL Updated that Flex to send the Server HDD to Waters and Scrap Certificate to WPPL. 09Oct19: JMD to get vendor that will scrap and issue certificate to vendor/Server HDD to return to WPPL 16Oct19: JMD to send to scrap today and vendor mentioned that Certificate Of Destruction available after 3 weeks only (to recieve by 30Oct19) 13Nov19: JMD emailed to HZL the Certificate Of Destruction and Scrap Form 3291 | Shamsul HZL | 2018-11-21 | 31-Dec-2018 31-Mar-2019 30-Jun-2019 16-Oct-2019 30-Oct-2019 | Close | Low
8 | Competency Tracking - Alliance 10 Apr 2019 : Need to provide objective evidence of closure, eg photo - Open item to provide target date 24 Apr 2019: to update the progress 21 May 2019 : 8X still need to process, ionized gun and fixturing in-progress target by - to send mark up ASP to WPPL 23 May 2019 (Tan CK) - Fixture ETA 24 May 2019 12 Jun 2019 : CK Tan to update the tracker and send out by end of today - Air-gun PO released to vendor TBD - Torque driver to have a two way driver and RFQ in progress / approx S$800 - 19 Jun 2019 - pending ECR 9229947 Approval stage / Airgun waiting for E.T.A. date by today (19 Jun 2019) tooling PN 999025154 by 21 Jun 2019 (CK Tan) Torque Wrench - use current torque wrench and no issue with operation purchase of new torque wrench for improvement will be under separate item for review - 26 Jun 2019 Next week to have 1x Airgun and 1x tooling PN 999025154 (Tan CK) - 03 Jul 2019 Delivery confirmation to follow-up for Airgun (CK Tan 07 Jul 2019) - 10Jul 2019, Air gun received on 04Jul19 and installed on 05Jul19. Tooling 999025154 received on 08Jul19. | Tan CK | 2019-04-10 | 30-Apr-2019 30-May-2019 30-Jun-2019 05-Jul-2019 | Close | High
9 | 210000415 Shutter Home Failure Investigation - Install to running model 186015032 and observe repeatability of failure by 14 Jun 2019 - Capture and share hyperterminal data to team 19 Jun 2019 No build PDA and eLambda from 14 to 19 Jun 2019, Golden Unit is not able to capture the reject , will check next PDA build (19 Jun 2019 Joel) - will try using elambda and request for build 1 PDA unit to use for verfication (19 Aug 2019) 26 Jun 2019 Verified using eLambda and PDA Good Unit - Found to be NTF / release the board to the line | Joel | 2019-06-12 | 2019-06-26 | Close | Medium
10 | Detector PCBA Fixture Implementation - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08-Jan-2020 : ECR 9233596 waiting for submission 05-Feb-2020 : Implement PCB fixture for Detector Family as per : ECO 9239291 (289024872 : 279001743) ECO 9232732 (279000805 : 279001961 : 279002224 : 279002897 : 279003287) ECR 9233596 cancelled submission due to new PCBA design | GLY SFC | 2019-06-12 | 30 Sep 19 30 Dec 19 07 Feb 20 | Close 05-Feb-20 | Medium
11 | TE82969 and TE82882 (7x testers) Validation to accomodate new spares (OESR 193421) -12 Jun 2019 : To try the new Linearity Test Software ver 2.2 \ Final Test ver 4.2 -19 Jun 2019: Test in Progress, refer to daily update to team / Get VAF Approval Status (Joel) -26 Jun 2019 HZ preparing IQ for submission / JMD to send test report today -03 Jul 2019 Eric Turner increase empower database to remove loss comm (Joel to try today) - To check with Eric on how to increase empower data base (HZ) 10 Jul 2019 - In procees. HZ have increase database size from 50mb to 1500mb for all PDA tester. Joel to monitor if there is any loss comm issue arise from today onwards. | Joel / Hua Zhen | 2019-06-12 | 6/19/2019 6/26/2019 7/05/2019 | Close | High
12 | TE82943 (8x testers) Qualification to update Firmware & driver Pack - 12 Jun 2019 TMD010661 v01 release and effective 03 Jun 2019 - 19 Jun 2019 TE82943-11 Test Station updated to version 1.8 and ver 2.3, Test Report submitted, pending approval from Milford. (Yee Hwa) - 26 Jun 2019 TE82943-12 Test Station updated to version 1.8 and ver 2.3 last night, Test report submitted to YHC - 03 Jul 2019 TE82943-11 Approved / TE82943-12 IQ Report approval in progress (YH) - 10 Jul 2019: 2 testers are routing for approval - In process - 17Jul 2019: 14 - 18 testers are still in process - 24 Jul 2019: In-process. By this month to complete. (Yee Hwa) - 31 Jul 2019: TE82943-16,17 and 18 on-hold for a week as E. Jackson is onleave. - 07 Aug 2019:TE82943-16,17 in-progress of testing (Joel) - 14 Aug 2019 TE82943-16 and 17 completed and submit to YH (TCK Approved),TE82943-18 Retest is progress (Joel) - 21 Aug 2019 TE82943-18 Test Report, Calibration Report with Evaluation Record for EQID 0497 has been submitted to CYH - 28 Aug 2019 To check with ITS if able to do the PM and verification before 30 Aug 2019 (JMD) Jia Chen Lim (ITS) confirmed schedule balance PM and verification by 02 Sep 2019, due to tight schedule -02 Sep 2019 - ITS Jia Chen verfied balance and found to be within specification -04 Sep 2019 Flex to arrange external calibration for Balance (JMD - arranged 13 Sep 2019) -11 Sep 2019 Ming Deng calibrated AND HR 200 Balance and found to be within spesc (waiting for official report) -24 Sep 2019 EQID 0497 SN 12326186 Calibrated and within specs released to production. VAF not required as per HZ advise (to follow 730001835) | JMD YHC HZL | 2019-06-03 | 15-07-2019 31 Jul 2019 30 Aug 2019 27 Sep 2019 | Close | High
13 | Carryover FPY Tabulation for Alliance - 12 Jun 2019 Continue the tracking and send to WPPL - every Friday Morning (9:30am) - 19 Jun 2019 Routine submission as is. - 26 Jun 2019 Routine submission as is. - 03 Jul 2019 Understand the 5% on WW26 and set trigger point - 17 Jul 2019 Continue monitoring - 24 Jul 2019 Continue monitoring - 11Dec 2019: To close this activity and go by routine monitoring by production line. Any spike to trigger WPPL >95% only for Carryover | Joel / Jing Xuan | 2019-06-12 | 2019-07-31 | Routine | Low
14 | JBS For Alliance Needle Alignment - 12 Jun 2019 CK Tan to start drafting information - 19 Jun 19 : JX will share template to CK - 21 Jun 19 CK Tan to start drafting information and share with WPPL - 03 Jul 19 - 26 Jun 2019 JBS preparation of visual aid (Tan CK) - 03 Jul 2019 CK to send out Draft copy for review (CK) - 10 Jul 2019 Revisit on the Draft JBS and enhance on the KeyPoints and Reason section. - 17 Jul 2019: CK to revised and send out to WPPL for review - 24 Jul 2019: Completed reviewing with Desmond - Send out to Edwin and Desmond by this week. - 31 Jul 2019: Desmond commented the JBS and CK to update accordingly - 07 Aug 2019: JBS is send to Desmond dated 31Jul19. Pending for Desmond's advice on the next course of action. - 14 Aug 2019: Pending for Desmond's advice on the next course of action. 28 Aug 2019 : JBS agreed, Desmond to pass the JBS to CK and insert as per formatting 04 Sep 2019 : CKT to verify ASP. 11 Sep 2019 ECR 9231435 - Review and approval in progress in Waters 18 Sep 2019 Track as per ECO process | CKT JXL | 2019-06-12 | 12-Jul-2019 17-Jul-2019 30-Sep-2019 | Close | Medium
15 | Consign PCBA Debugging (Alliance Instrument) at Senai 26Jun19 to 06 Nov19 refer to previous minutes 13Nov19: Training reschedule due to QBR Preparation. Reference to used ASP 279002695 and TSP WAT271129 and event log recording. 20Nov19: Training completed on the ASP 279002695 and TSP WAT271129 and event log recording.(to send event log sample to WPPL) 27Nov19: Training conducted last 15Nov19 for event log and TSP. | JMD DA | 2019-06-26 | 06 Nov 2019 30 Nov 2019 | Close | Medium
16 | OFF Line Verification 1. Known Good Board Verification must have SN of the UUT as well as failure symptom 2. Action taken to verity the unit 3. Timeline of the verification process workflow by next week (JMD) | Joel | 2019-06-26 | 2019-07-03 | Close | Medium
17 | TUV Flow Cell Low Energy Level. - Need to check the TUV Shunt Tranmittance value. (5 data points each shunt) - review the ASP and TSP - PM to submit (Joel) 21 Aug 2019 - Joel reviewed TSP 735001356 TUV Detector, Final Test Procedure and it does not specified transmittance value TUV TSP's . WPPL to advise after verification from Milford 28 Aug 2019 - Hua Zhen to follow-up transmittance value to Milford (JMD to HZ) 04 Sep 2019 - HZL send email inquiry to MJ Milford (pending advise) 16 Sep 2019 - Mark Jussaume update no limit and sees reading 0f 80, he put Eric in list hoping may have some add'l knowledge of the shunts 18 Sep 2109 WPPL to follow-up in knowledge base info. | JMD HZL | 2019-08-14 | 2019-08-30 | Close | Medium
18 | 186003031 - X-bridge Column Replacement Frequency - to check with HZ the criterial for Column Replacement Frequency 21Aug19 JMD emailed HZ to seek advise for frequency of replacement 28Aug19: Avg per port is 23.7 times (round up to 24) Per test, number of injection is 53 times. Avg per column injection every mth is 24 x 53 = 1272 Pls. advise what is the max. injection before we can replace column (Hua Zhen to follow-up Milford or Mr Wong) 04Sep19: Mark Jussaume advised that 6 months replacement of column is not applicable to Flex at the moment. Its only applicable to new Alliance tester. And the replacement interval hasn’t been finalized yet. So please keep monitoring the usage of column and record the frequency of replacement. 11Sep19: Flex replaced 1X Column Alliance Final Test 3.4 dated 08 Jul 2019 (Chemistry Test consistent fail) 13Nov19: Flex replaced 1X Column Alliance Final Test 4.4 dated 07 Nov 2019 (System Overpressure) 27Nov19: (move to monitoring phase) 11Dec19: Started monitoring for 6 months and result only 2 replacements. To monitor for anotther 6 months. Status change from Monitoring to Routine. 17Jul20: to set the date end of this year | David HZL | 2019-08-14 | 20 Jun 19 31 Dec 19 30 Jun 20 31 Dec 20 | Routine | Low
19 | OESR # 193460 Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 2020-01-31 | Close | Medium
20 | OESR # 193461 Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 31-Jan-20 31-Jul-20 | Close | Medium
21 | OESR # 193124 Additional Capacity for Alliance Pump Tray Tester (TE82526) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update (Closed @ Minutes of Meeting Tab and Open @ OESR_v01 tab) | JMD HZL | 2019-08-14 | 2020-01-31 | Close | Medium
22 | OESR 190059 Alliance Final Test 3 Frequently Freeze During Final Test OESR 193244 Milford OESR - 03 Apr 2019 04 Apr 2019 - Matt Brule request Flex to provide detail on Hang issue 05 Apr 2019 - Joel provide detail to Matt Brule for detail of Hang issue 04 Sep 2019 - Joel to monitor and update by 06 Sep 2019 06 Sep 2019 - JMD sent monitoring data to HZL (refer rto attach Alliance FT3 System Freeze Tracking Record) 09 Sep 2019 : HZL updated Matt Brule with attached Alliance FT3 System Freeze Tracking Record 18 Sep 2019: HZL last occurrence is 3 Sep 19, slow build for Alliance follows thus need to extend monitoring / with respective volume (min 100 units), HZL to inform Matt that Flex to further monitor this issue. 25 Sep 2019: JMD issue the tracket for 100 units to Prod'n and monitoring in progress 31 Oct 2019: Monitoring for 100 instrument competed> 23 units freeze / 100 units | JMD HZL | 2019-09-04 | 13 Sep 19 27 Sep 19 30 Oct 19 | Close | High
23 | OESR 190079 TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. 15May2019 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04Sep2019 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11Sep2019 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18Sep2019 : JMD will inform YHC once EQID 0497 is available 19Sep2019: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25Sep2019 VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. | JMD HZL | 2019-09-04 | 2019-09-27 | Close | High
24 | 190102 - TE82988-2 Power Supply 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
25 | 190104 - TE82988-2 Spare Parts holding at Senai - to request return back to WPPL 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10 Sep 2019 - Erzad update ETA back to WPPL by 23 Sep 2019 | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
26 | 190108 - TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep 2019 FLX Kallang JMD to follow-up to Edwin 05 Sep 2019 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
27 | OESR S190109 - TSP having 735002531 with 735001613 does not have consistent content 04 Sep19: reviewed with HZL/EK/DA and JMD and agreed to update the OESR content and its Model PN. 05-Sep-2019: JMD updated the OESR 190109 and send to HZL 18-Sep-2019: HZL to checked for update 24-Sep-2019: HZL to follow-up with E Turner 27-Sep-2019: HZL to follow-up with E Turner reminder due as Waters Annual Audit for Flex 08-Oct-2019 : HZL send reminder to E.T 08Oct19 - HZL ask ET if need to revise procedure/monitoring or uprev the next revision. 13-Nov-2019: HZL to check for update 11-Mar-2020: JMD to send email update to HZL that QDMS is not allowed to have consistent TSP applicable to particular model PN 18-Mar-2020: JMD send email to HZL but need to get more justification (will send another update today) 01-Apr-2020 : ECO 9234775 last 25-Mar-2020, revision of Flex document in-proress. 08-Apr-2020 : Technicians trained and released document WAT-RW-SKI01718 Rev L \ 735001613 Rev YY (03 Apr 2020) | JMD HZL | 2019-09-04 | 06-Sep-2019 10-Oct-2019 30-Dec-2019 08-Apr-2020 | Close (08-Apr-2020) | Medium
28 | 190117 - TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05 Sep 2019 Start 20 Aug 2019 used in 210000411 SN 010919110M 11 Sep 2019 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
29 | OESR #190118/ #190119 Reticule Plate Broken Target Mirror 19Aug19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30Aug19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04Sep19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23Sep19: SFC relelase PO to vendor 10Oct19 PM documentation approval in-progress 16Oct19: Part lead time is 4week from now (16Nov19) 20Oct19: GLY assembled and release to Production. | JMD SFC | 2019-09-04 | 2019-12-30 | Close | Medium
31 | 190124 - 193544 Request for TE# for ICT Fixture 210001344 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05 Sep 2019 Erzad to send the documentation to JMD 06 Sep 2019 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11 Sep 2019 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
32 | OESR S190130 -193550 TE82882 PDA Final Test System Missing Spare OB PN 700011652 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18Sep2019 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25Sep2019 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26Sep2019 - TE82882-10SEP19-VAF_Approved and released tester for usage. | JMD HZL | 2019-09-04 | 06 Sep 19 30 Sep 19 | Close | High
33 | OESR S190132 - TE82943-13 USB Device Error 03Sep19 JMD raised OESR for TE82943-13 USB Device Error 04Sep19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06Sep19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10Sep19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11Sep19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18Sep19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25Sep19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09Oct19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 10Oct19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10Oct19: Informed purchaser to release pay using Paypal. 16Oct19: JMD to follow-up E.T.A from buyer 23Oct19: Bliss target ETA to Flex is on 29Oct19 30Oct19: To install today (Shamsul) and update status 06Nov19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | JMD HZL | 2019-09-04 | 06 Sep 19 18 Sep 19 25 Sep 19 30 Oct 19 | Close | High
34 | 190133 Label Printing 04 Sep 2019 GLY to follow-up with DA | GLY / DA | 2019-09-04 | 2019-09-06 | Close | High
35 | AND HR 200 Balance Misssing Segment Activity 25Sep2019: HZL to discuss with Edwin if scrap or repair .New Model GR 200 S$1765.5 Repair :TE82943-11 AND HR200 09Oct2019: To engage vendor to scrap with certification 16Oct19: JMD to send to scrap today and vendor mentioned that Certificate Of Destruction available after 3 weeks only (to recieve by 30Oct19) 13Nov19: JMD emailel to HZL the Certificate Of Destruction and Scrap Form 3291 | Shamsul HZL | 2019-09-25 | 02 Oct 19 16 Oct 19 30 Oct 19 | Close | Low
36 | Alliance Test Station Study Upgrade to use for common use to other model (Sample Manager) 09Oct19: HZL to study the upgrading for the test station to save space 16Oct19: Milford is studying the proposal 30Oct19: Result from Milford is not feasible to have a common test station. | JMD HZL | 2019-10-10 | 2021-01-31 | Close | Low
37 | OESR S190156 TE82503-5 Injector Seal Pack Ramp and Decay Test Fail 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | JMD HZL | 2019-10-03 | 2019-10-21 | Close | High
38 | Flex-WPPL-CEI Tripartite meeting (David) - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08Jan2020: last update: Still under review by Waters Milford on the program proposed by CEI, information shared to WPPL Team Ops, any new information will update WPPL OPs Team and close this action item. | David HZL | 2019-10-30 | 2020-01-08 | Close | Low
39 | Cross Thread Analysis - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: ECO 9233201 to send to Flex today 15 Jan 2020: Start impelementation today | GLY GYC | 2019-11-06 | 2020-01-17 | Close | Low
40 | Hot Part - Return to Flex on 30 Oct 2019 - 700001031 removed from Alliance unit # L18SM4424A - RMA-SKW-00000136 - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: awaiting official instruction from WPPL SQE / Investigation from OPS Team already close as of 09 Dec 2019 | AP JMD GYC | 2019-11-06 | 2019-12-09 | Close | High
41 | Lean Project Update - 196015031 Failed Noise Test - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update | AP JMD GYC | 2019-11-06 | 2020-01-03 | Close | High
42 | Additional tacker 11Dec19: Add OESR and TMD in this Engineering Meeting for the purpose of Status update and transparency from both Team (WPPL and Flex) | TEAM | 2019-12-11 | 2020-12-31 | Routine | Low
42 | Asset Manual Update - REFER HISTORY from 24th Dec 2019 Weekly Eng'g Update 08 Jan 2020: JMD presented to HZL and to email the Asset List by today | JMD HZL | 2019-12-24 | 2020-01-31 | Close | Low
43 | 210000415 Baking and Retest Evalutaion 15-Jan-2020 - Flex Senai to bake today and once finisk, will be cool down and zip lock before sending to Kallang 22-Jan-2020 - Baked 210000415 arrived in sealed ESD Bag with desiccant and good humidity indicator. We retest and the failure still the same : Shutter Fail to Home.(16 Jan 2020) Pending advise from GYC 05-Feb-2020 : Shaynee to revise a complete unit PO (Shipment ONHOLD) and Guan Yee to create PO for PCBA only. 12-Feb-2020 : PO received from GYC (06 Feb 20), shipped the PCBA attention to Bill Bates (07 Feb 20), FEDEX 7777-1061-9174 (received date: 11 Feb 2020) | JMD GYC | 14-Jan-2020 05-Feb-2020 | 22-Jan-20 31-Mar-20 | Close | Medium
45 | Consumables to include/update SPARE Parts for PM 15-Jan-2020 - Update in the tab for consumable. HZL will share the MRO form to be fill-in by Flex for ordering PM parts as well as the required consumables for Final Testers. Part to be direct ship to Flex. 05-Feb-2020 to send today 12-Feb-2020 Consumable updated and order must be in quarterly usage (as per discussion with HZL 11 Feb 2020) HZL to check and advise E.T.A from Waters logistics once MRO approve JMD to raise MRO by mid of March 2020 for Q2 19-Feb-2020 receieved the item as per MRO 17 Feb 2020 and to raise MRO in quarterly base on loading (Close) | JMD HZL | 2020-01-15 | 29-Jan-20 05-Feb-20 | Close 19-Feb-2020 | Low
46 | QSP Interface Tubing (430001150) Frequency of Replacement due to 1. Test flowrate accuray 0.5ml/min , 2. Test Flow Rate Accuracy 2.0ml/minTest Refer to previous minutes for deleted items 22-Jul-2020 - Flex experienced low yield for the last 2 weeks and the troubleshooting result was retightening . We are not sure if our tech’s are retightening the fittings on the UUT or the interface tubing. We interviewed them and found out that most of the time was retightening of interface tubing. 29-Jul-2020 - Plugged TE82943-11, 12,13 and 14 and unplug and installed new Interface Tubing on TE82943-14, 15, 16, 17 and 18 (27 Jul 2020) 05-Aug-2020 - TE82943-14, 15, 16, 17 and 18 and swap by 10 Aug 2020 The team concuded today to stop the evaluation of the average interface tubing replacement. Flex to continue monitor the yield. 12-Aug-2020- TE82943-11, 12,13 and 14 activated and installed new Interface Tubing. Tester #15, #16 #17 and #18 deactivated and plugged 19-Aug-2020- Swapping of 4x tester is being carried out every 2 week to make sure all testers are being used and replaced the Interface tubing on time 26-Aug-2020- Continue the practice and monitor the yield for update (to observed build plan as 26 Aug, Production request to activate 8x tester.) - to look into reseating of components for the test flow rate accuracy | JMD HZL | 2020-01-02 | 02/29/2020 03/30/2020 06/30/2020 | Open | Medium
47 | Loss Comm to align TSP to Tester TE82882 PDA Final Tester Refer to previous minutes for deleted items 22-Jul-2020 Refer to Manual Data Collection (502/600 Tested as of 20 Jul 2020) 29-Jul-2020 Refer to Manual Data Collection (600/600 Tested as of 28 Jul 2020) 1. Check tester cables and accessories wiring and see if can find differences (Completed 09 Jul 2020 JMD) 2. Swapping CPU between high and low FPY stations to see if it’s caused by CPU (Team to advise) 3. Clone high FPY station (-9) image and test it on low FPY station; (Team to advise) 05-Aug-2020 Team to discuss next action. HZL to share the 600 data to Milford (Eric Turner) and propose for the next plan. JMD to add the 600 data points summary in the CIP 2020 slide. 12-Aug-2020 HZL to share update from Milford (Eric Turner) and advice for the next steps. To continue discuss next week HZL to advise. 19-Aug-2020- Swapping of 4x tester is being carried out every 2 week to make sure all testers are being used and replaced the Interface tubing on time 26-Aug-2020 Eric to update HZL after return to office - HZL shared update from Eric Truner. - JMD to ensure all PDA tester stations Eric Truner is able to access. Remote desktop should be active. - JMD to reply that clearing of rogue IP is done every day and low yield Testers are -8, -11 & -13. - HZL to ask Eric Turner to share the actual change performed on Tester -8 and -11 (e.g. database space) for team understanding, how to check? if the rest can be done? any frequency to perform database space aavailable? - Start monitoring of Tester -8 and -11 yield. 09-Sep-2020 Monitoring in progress for TE82882-8,11 and 13 for 50x UUT | JMD HZL | 2020-01-22 | 2020-05-30 | Open | Medium
48 | Production utilization on tester Refer to previous minutes for deleted items 13-May-2020: Shutdown TE82882-8,9 amd 14 and see balance loading on next week - To check Tester utilization for BSM and QSM 20-May-2020 To publish on 22 May 2020 27 May 2020 -refer to tester loading (published 22 May-20), add-in the TUV loading 03 Jun 2020 - TUV Loading added to Dual Noise CI Slide 10-Jun-2020 Tracking of tester utilization in-progress for all the 3 model/family. (refer to updated slides) (need to check volume as per build plan.) - in-progress, continue monitor per build plan 24-Jun-2020 Prod'n still having low build as per build plan (KIV in Q4'2020) | JMD HZL | 2020-01-22 | 02/05/2020 04/11/2020 04/30/2020 05/30/2020 | Open | Medium
49 | 2690 Injector Sensor Adjust - that biasing the sensor adjustment so top count was closer to upper limit would produce good test results 05-Feb-2020 : To look into DHR/Sub Assembly Test and provide details to Mark Jussaume 12-Feb-2020 : Submitted 24 data sample to team for review and advise (06 Feb 2020) HZL to follow-up with MJ 19-Feb-2020 : MJ receieved all the update and information he need and advise to close the activity 14-Feb-2020) | JMD HZL | 2020-02-05 | 2020-02-07 | Close 19-Feb-2020 | High
50 | PC BACK UP (Quarterly Backup) - 05-Feb-2020 To backup PC quaterly | JMD HZL | 2020-02-05 | 2020-02-19 | Close 19-Feb-2020 | Medium
51 | Dual Noise Due to C-Clip 12 Feb 2020 : Observed that the C-Clip is touching the screw refer to MTR Slide) 19-Feb-2020 Found one unit failing in dual-noise (85 counts), removed OB and removed grating drive observed C-Clip, not able to see screw head touching the C-Clip, re-assembled and retest. Retest result result Fail, GYC request to replaced screw. To update result after eng'g review. 17:30 PM GLY updated that the UUT still failing for Dual Noise (Investigation to be continue) 26-Feb-2020 : Tried 2x units and still failing, last replaced is motor (PN-GLY TBA) and able to pass. 04-Mar-2020 : Gary send RCA for GLY, for review, not ble to conclude if related to C-Clip, close the item but still need to open another action item for Dual Noise under CONTINOUS IMPROVEMENT PLAN 2020 | GLY GYC | 2020-02-12 | 2020-03-11 | Close 04-Mar-2020 | High
52 | EVALS20002 : 279001810-L- ASSY, GPV, DUAL CHANNEL, QSM Test Input Prime MeOH Fail 19-Feb-2020 : To update the latest activity as per GYC instruction (submitted the updated excel file to KS Chua 25-Feb-2020 : Concall with WPPL, Milford and Flex 04-Mar-2020 : To raise evaluation for sa per Randy input / JMD to GLY to decide the evaluation plan 11- Mar-2020 : Completed the 8 pcs evaluation as per EVALS20002, data compilation in progress 18-Mar-2020 : Miss out the step 4 verification of tubing for valve to waste backpressure if able to produce 10-85 psi, today to check availability of 515 Pump and perform back pressure test. (JMD) 25-Mar-2020 : Submitted to Team the EVALS20002 - 3 Fail / 5 Pass, for team disposition 01-Apr-2020 : WPPL QA Tony and Guan Yee to witness the retest of 3x fail from evaluation by 2 PM , must base on the same report submitted. 08-Apr-2020 : Some evidence of imperfections with 2 GPV manifolds (S/N’s 76465 and 76485). These 2 GPV assemblies were assembled into the Waters Instrument and failed during a diagnostic test. The diagnostic test ran Valves D, C, B and finally A 4 different times and the results of this test are found in the attached presentation. | JMD GYC | 2020-02-19 | 08Apr 2020 | Close (08-Apr-20) | High
53 | Introduce usage of Engineering Evaluation Form | Team | 2020-02-16 | TBD | Close 11-Mar-2020 | Routine
54 | EVALS20003 : To evaluate the transducer PN 279010881 using PN 186015080 1. Refer to TMD 10782 v01 for details 2. Measure 100 pcs of transducer voltage criterial as per drawing 0.0750 ± 0.0125 VDC @ zero balance 3. Publish the result to the team. 18-Mar-2020 JMD to publish later result for today 25-Mar-2020 Completed build and test for 50 units. All PASS Final Test 01-Apr-2020 Engineering activity, QA complete buy-off | GYC/ GLY/ JMD/HZL | 2020-03-04 | 2020-04-01 | Close 01-Apr-2020 | High
55 | Sigma Sure 7 Computer for Detector 11-Mar-2020 JMD to check with HZL if able to use the old CPU from pump tray, no changes on settings, just for Sigma Sure 7 printing 18-Mar-2020 : To use removed Pump Tray Tester CPU and used as SS& access for detector line, update the PC List and assett manual. z 25-Mar-2020: CPU Monitor and setup completed and need advise from HZL for the access SS7 portal. 01-Apr-2020: JMD used the image of Genealogy-SDS and rename to GENELOGY-DET and change IP address | JMD HZL | 2020-03-11 | TBD | Close | Low
56 | Consumable Data for Wear and and Tear 11-Mar-2020 HZL and JMD to categories if will go to OESR, adhoc or repair 18-Mar-2020 Continue the discussion to avail the SS7 computer for detector line 25-Mar-2020 Remove the Wear and Tear / adhoc 01-Apr-2020 Adhoc to cover by OESR / keep those related to MRO 08-Apr-2020 MRO form to raise and approve. | JMD HZL | 2020-03-11 | 2020-03-25 | Close 25-Mar-2020 | Low
57 | DUAL-NOISE ACTIVITY Refer to prevoius minutes the for deleted information: 13-May- 2020: - Aldwin send EVAL for 30 mins. Power ON before test. - SS7 confirmed 221 units no dual noise. - to investigate what change on the line. - to get sampling test report and verify the dual noise (find the small change) that make difference. 20-May-2020 - No changes on the line. Test procedure was followed as intended usage and specification, sampling report does not observed changes on test result. - to continue investigate the improvment of Dual Noise 27-May-2020: No changes on the line. - Major Change : Briefing conducted to utilized TSP 735001356 and 735003877 that specs to run test at 99 cycles. 03-Jun-2020. Briefing conducted to Production personnel to run dual noise test till 99 cycles as per TSP 735001356 and 735003877. - JMD to collect data for Dual Noise Run Count for 1 week data 10-Jun-2020. Failure analysis for Jan and Feb 2020 includes 1.retesting without doing anything, 2.realignment the grating drive, 3.replacement of screw 1st WEEK Jun 2020 update - no dual noise failure. - to conclude that JAN-FEB'20 - due to re-seating; MAR'20 briefing to production allow 99 cycles, for MAY-JUN continue to monitor - 5.22% need to be corrected 24-Jun-2020. Correction has been completed for team review, to wait for one more week data 01-Jul-2020 No Dual Noise reject for month of May 2020. Monitor for final data Jun 2020 data 14-Jul-2020 No Dual Noise observed for 1st week of July 2020 15-Jul-202 Monitoring KIV until Q4 26-Aug-2020 No Dual Noise observed to date 09-Sep-2020 Request for Dual Noise CIP in-progress | GLY | 2020-03-18 | 2020-06-30 | Open | Low
58 | Validation of TE82943 - To update the TE82943 Test Software Refer to previous minutes for deleted items 13-May-2020: Approved today for TE82943-14 and 15, inform Eugene to relay message to update TE82943-13 and TE82943-16 tonight 20-May-2020: TE82943-13 and TE82943-16 Approved 19-May-2020, Flex to transfer the UUT's to TE82943-17 and 18 for Edward to upgrade. 27-May-2020: Test record submitted to WPPL on 22 May 2020 , TE82943-17 and TE82943-18 has been approved | JMD GLY Eugene Lim | 2020-03-25 | 5/19/2020 5/27/2020 | Close (27-May-2020) | Medium
59 | Calibration Master List 03-Apr-2020 JMD to add tab for "Calibration Master List" in the Eng'g Minutes 08-Apr-2020 JMD to added tab for "Calibration Master List" in the Eng'g Minutes | JMD HZL | 2020-04-03 | 2020-04-15 | Close 15-Apr-2020 | Medium
60 | 210000416 - Activity to compile requirement for ICT Test Sequence,to cover test U7 and U13 29-Apr-2020 : JMD checked with Senai, review is in progress in Senai 06-May-2020: Senai containment action is Bench test verification on this 2 location, U13 and U7. - HZL communicated to Milford peding reply from them 13-May-2020: To transfer to OESR | JMD HZL | 2020-04-22 | Transfer | Close | Medium
61 | BSM\QSP Manual Priming Refer to prevoius minutes the for deleted information: 22-Jul-2020 11/16 PN186015001 UUT solvent line was primed prior to Final Test 3x fail due to Test Seal Wash Pump Prime/Flow 1x fail due to Test Pump A Input 2 Prime with IPA 29-Jul-2020 16/21 (76.19%) PN186015001 UUT solvent line was primed prior to Final Test 4x fail due to Test Seal Wash Pump Prime/Flow 1x fail due to Test Pump A Input 2 Prime with IPA UUT able to pass after retest 12-Aug-2020 22/27 (81%) 19-Aug-2020 No build of BSM until 10 Oct 2020 26-Aug-2020 - JMD to review DHR of the 9 units QSM Plus that cause Manual Priming yield low 09-Sep-2020 - Refr to attached result for Before priming and after priming result | JMD | 2020-04-29 | TBD | Open | High
62 | To reverify the cycle time for the new pump tray. (Joel) 06-May-200: To verify the before and after yield and CT 13-May-2020: - new and Upgraded Pump Tray Tester Total Cycle Time is 25 mins. (19 mins Auto test and 6 mins for Handling ) - Old Tester is 39 mins. - Breakdown of cycle time of 25 mins. 20-May-2020: Refer to "New Pump Tray CT" tab, FPY data collection in-progress 22 May 2020 03-Jun-2020:-Old Yield as per Dashboard is 80% and Present yield is 83%, (Dashboard Updated) - Update the Asset List with new photo and quantity (Completed) - Update the capacity calculator and dashboard (request Amy to send the capacity calculator) 24-Jun-2020 - JMD to provide CT breakdown to Amy. 01 Jul 2020 Flex Senai received the SMU. Team agree ro close this activity . | JMD | 2020-05-13 | 2020-07-01 | Close 01-Jul-2020 | High
63 | To ensure checklist follow by Production - JMD to check line people for following checklist 10-Jun-2020. Conducted briefing to Production personnel (10-Jun-2020) Verification by Eng'g on following the designated checklist (11-Jun-2020) - to review if information in checklist e.g. sincker if can be captured into TSP/PM 24-Jun-2020 To be discuss wit the team with briefirn and closure date Add'l checkpoint into the test record (Alliance and DHR for BSM and QSP for Sinkers) Sinker 01-Jul-2020 Checklist OE presented and team agree to close | JMD | 2020-05-20 | 2020-07-01 | Close 01-Jul-2020 | High
64 | 210000416 Fluidics PCBA 24-Jun-2020 To understand and gather information for 416 failure related to U7, U3 and U13, JMD to provide update on 26-Jun-2020 01-Jul-2020 Joel submitted preliminary report to team More detail need to learn. | JMD | 2020-07-24 | 2020-07-01 | Close 01-Jul-2020 | High
65 | 2489 Lamp reliability Test Units - Scope of work 186248900 01-Jun-2020 Team to advise (KIV) 15-Jul-2020 - Quote to WPPL to Refurb to current revision. 29-Jul-2020 Reliability unit need to refurbished by : 1. Replace Ellipsoidal DT Mirror P/n: 402000170 2. Replace 402000451 Deutium Lamp 3. Reoptimize Optic Bench 4. Full Final Test 05-Aug-2020 - to get qoutation from Program 19-Aug-2020 - Quotation sub,itted to WPPL 26-Aug-2020 Awaiting for PO for processing (4 units), in progress pending Edwin to approve. 09-Sep-2020 Completed the 4x Detector for Tester | JMD | 2020-07-01 | TBD | Open | Low
66 | Solvent Container Cut-out window - GLY to start preparation of cut-out for the window 14-Jul-2020 Solvent container window has been cut-out and now placing the plastic cover (in-progress) 15-Jul-2020 placing the plastic cover (in-progress) target end of Jul '20 29-Jul2020 In-progress actvity until end Sep 2020 | GLY | 2020-07-01 | 2020-08-12 | Close 12-Aug-20 | Medium
67 | Forecast Number reviewed for consumables increase in Q3 15-Jul-20: was the consumables reviewed and everything is sufficient | JMD | 2020-07-15 | 2020-07-22 | Close 05-Aug-20 | High
68 | CCD Inquiry 05-Aug-2020 - Flex team to share informaions gathered 1)How many failures of CCD instance of “Carousel Malfunction” (mostly) and “Carousel Z-axis 5 malfunction in the past 6 months and percentage failed. 2) Daily verification record of carousel in Final test station. 3) Last replacement and avg frequency of replace of carousel in Final test station 4) Total number of sets of carousel in Alliance final test station. 5) PM record of both CCD alignment station. 19-Aug-2020 - 4M has been shared to the team last 14 Aug 2020 | JMD | 2020-08-05 | 2020-08-07 | Close 26-Aug-20 | High
69 | 515 PUMP Competency Closure 05-Aug-2020 - to update by GLY after review - WPPL Eugene request for Monthly schedule for competency check. - Monthly from WPPL will be coming down for competency check. - Flex requesting WPPL to share the plan to know the date and model 19-Aug-2020 Team to advise for closure | JMD | 2020-08-05 | 2020-08-31 | Close 26-Aug-20 | Low
70 | Golden Unit List 05-Aug-2020 - JMD to update the golden unit list and send to HZL - JMD to perform retest for PUMP family 19-Aug-2020 515 Golden Unit - retestand pass testing 1515 - to retest after premier detector activity 1525 - to retest after premier detector activity 26-Aug-2020 515 Golden Unit - retest and pass testing - 1515 - to retest after Prod'n complete their unit , not able to communicate with tester - 1525 - No CPU board to retest / different PN - Team agree to decare not Golden. To scrap this unit if no need. | JMD | 2020-08-05 | 2020-08-31 | Open | Low
71 | CCD Capacity Review 05-Aug-2020 - JDM & HZL to align with Program on 12 months loading number until 2021 | JMD | 2020-08-05 | 2020-08-31 | Close 12-Aug-2020 | Low
71 | PM For CCD - for discussion of Plan 09-Sep-2020 - for internal discussion prior sending update to OPS Team | HZL/JMD | 2020-08-26 | TBD | Open | High

## Venture
Waters Engineering Weekly Meeting | [7 empty] | Work Week 22-2020
Item | Priority | Action and Updated Status | OESR/Owner | Status | Required Completion Date | Actual Completion Date | Priority | Reference document | WPP Discussion
1 | High | BSM RSC packaging [04Aug20] The TUV test date still in the plan on 13-Aug. [12Aug20] The TUV test date still in the plan on 13-Aug. [19Aug20] The environmental storage tests to be completed by 24-Aug, 7 units to be returned to Venture for inspection [26Aug20] The palletized units had returned on 25-Aug, there is no observation of carton bulging, will check the carton flags crack lines on 26-Aug. [27Aug20]All 7 units had been completed the visual inspection together with Guan Yee and no issue found. One unit will go through functional test. [02Sep20] The 7 units had just back on 1-Sep for full functional tests, they are under visual inspection, the results are good so far. Target to complete and ship the palletized units for final round of tests on 8-Sep. [09Sep20] All the 7 units passed the visual and functional tests except 1 of the palletized unit encountered Dynamic Leak due to I2V check valve faulty. As this is a known failure, the I2V check valuve is classified as material faulty, not induced by TUV tests. [16Sep20] The palletized units had passed the final round of TUV tests and verification tests. The overall results are good and the qualification was completed. | GES CK / WPPL Guan Yee | Closed (1st test) Open (2nd Test) | 31/3/2020 End of April'20 7-Aug-20 | 2020-04-28
2 | High | FLR 2475-- Raleigh wavelength failure [05Aug20] Pending to hear comments from Matt after he reviewed the first set of data with new alignment method. [19Aug20] Matt reviewed the data and he said new alignment tools do not look ready to implement and stop testing/aligning benches with them. Continue production using the existing alignment method with the Golden Flow cell. He will discuss with the design team for the next steps and most likely he will request to send back both Photodiode and LED tool back to him. High backpressure of Golden flow cell issue highlighed to Matt and following his advice on flushing of flow cell partial blockage. [27Aug20]Venture to rasie another OESR to request golden flowcell ( S20070 ) [02Sep20] Running some experiment requested by Matt before sending back the LED flow cell fixture. [09Sep20] Sent experiment data to Matt to review. | GES Mi Mi / WPPL Yee Hwa | Open | TBA
3 | Low | QSM -- Test flow rate accuracy 0.3ml/min -- Empty breaker and zeroing scale [04Aug20]Re-route waste tubing at TE82943-1 and tested 3/5 and 0% failure. To continue to monitor [11Aug20]5/5 tested and 0% failure of flowrate test. [19Aug20]will continue to test few more units and update the result accordingly [26 Aug20] Tested 3 more units and no failure observed. (Total: 8/8 passed. 0% failure) [27Aug20] To collect 7 more units and to move this item to monitoring if there is improvement [09Sep20] To collect 5 more units and to move this item to monitoring if there is improvement [16Sep20]To collect 2 more units. Result (2/13= 15% failed) | GES Chan / WPPL HuaZhen | Open | TBA
4 | Low | QDa ICT fixture [01Jul20]Transfer ICT fixture for DCC board from TT to Venture. [08Jul20] Edwin to work with Ray for the schedule of transfer. [25Aug20] TT is still depleting the existing stock, will advise the transfer date | GES Ricky / WPPL Edwin / Sadiq | Open | TBA |  |  | * This tester is to be tranferred from TT to Venture as a spare.
5 | Medium | DSOP - Cioplenu Pilot Run [04Aug20] To start the laptop setup for the trial run on 14-Aug. [12Aug20] Received laptop on 6-Aug, still in the setup process. [19Aug20] The DSOP Pilot run had started on FTN 5085 on 19-Aug, target to complete by 18-Sep. [26Aug20] 40 sets of data completed. [02Sep20] To continue the remaining 40 sets of data, total : 80 sets of data, target by 11-Sep. [09Sep20] 80 sets of data was completed on 8-Sep, to review the details with WPPL by this Friday. | GES CK / Alan / YT WPPL Sid / Edwin | Open | 2020-09-04
6 | High | ECO9235548 - 735003612 (REV G) for QSM Premier (S.W : 1.4) & ECO9236932 for SW 1.5 [11Aug20]Received updated TSP for software 1.4 but the current s.w in GES is 1.3 [27Aug20] Target to update new test software in GES on 2nd week of Sep. This sw is also to improve priming. Once software is upgraded, Chan to provide the FPY comparision before and after and troubleshooting time [16Sep20]Received another TSP rev H for sw 1.5 | GES Chan WPPL Ernest | Open | TBA
7 | High | Test Software Update and Validation for Premier [11Aug20]New test software for Premier is installed into 4 test stations and pending VAF to start validation [19Aug20]CMA test reports and the supporting documents are provided to Ernest . CHA is still in progress and update accordingly [26Aug20]Had provided all the supporting documents for CMP, CM-Aux,CHA and FLR. Test could not run and hang after initialization at FLR TE82760-7. This tester validation will resume only at next premier build [16Sep20]Received IQ approval of CHA, CM-Aux/CMP test software. | GES Chan WPPL Ernest/GY | Open | 2020-10-05 |  |  | Validation must be completed before MR
8 | Medium | SIT - Process and Plan to handle the Quarter End (WK13) upside [26Aug20] To brainstorm on process and control plan for materials and testing to handle the quarter end upside [02Sep20] Refer to the ppt for proposals. | GES Alan / Lwin WPPL Yee Hwa | Open | 2020-09-30
9 | Medium | VAVE - Cables & Power Entry Modules [02Sep20] Value Added Value Engineering for Cables and Power Modules. [09Sep20] Correct the pricing and add the update for the Tg of raw PCB. [16sEP20] Add the study for Electrical Shocks & High Temperature Labels and Metal Bezel "overspray permissible". | GES Alan / CK / YT WPPL Edwin / Yee Hwa | Open | 2020-10-09
10 | Low | BestLabs - Wrong wordings in the Calibration Report for Alarm Hygrometer [09Sep20] BestLabs to customized report template for Venture GES and let the Head of Lab for approval prior to start the calibration. | GES Kwang Quek / YT WPPL Ernet Chua | Closed | 2020-09-10

## Premier Detector (hidden)
ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | TSP 735005108 1. On page 3 , speficifc file name should be stated in TSP to give a clear instruciton for operator to select. Step 3.7 3.7. Select the appropriate firmware file (ACQUITY TUVII Detector.sft for the instrument type then click OK. 2. On page 3, speficifc file name should be stated in TSP to give a clear instruciton for operator to select. Step 4.1 Install analytical cell 289002450 Medium Shunt into bench . 3. On page 5. Leak Detect Test result should be added into Test Record. | HZL/ET | 2020-08-19 | TBD | Open | High
2 | TSP 735001613 Header of test record for TE83507 mismatch Product to TE82882-8 Download-Configure System - eLambda Detector PDA Final Test System - Acquity Premier eLambda PDA | HZL/ET | 2020-08-19 | TBD | Open | High

## ASP TSP Summary (hidden)
Pumps | ASP | 15 | 0 | 15 | 4 | 11 | 26.67% | 7x ASP belong to sub assy & Main Line 4x ASP belong to Packing
| TSP | 2 | 0 | 2 | 1 | 1 | 50% | 1x TSP belong to 515 pump
Last update: 26 Sep 2018
Product Family | ASP/ TSP | ASP/TSP Compliance Check by Flex/Waters | [6 empty] | ASP/TSP Compliance Implementation by Flex/Waters |  |  |  | PKI | Arc Bio
|  | Total | Not in scope -Low runners (Velox/5034/5053/Arc Bio) | In scope (High Runners) | Completed Review | Pending Review | % Status | Pending Document Remarks | Completed Implementation | Pending Implementation | % Status | Remarks
eAlliance | ASP | 30 | 4 | 26 | 26 | 0 | 100% | NA | 26 | 0 | 100% | completed
| TSP | 10 | 0 | 10 | 10 | 0 | 100% | NA | 10 | 0 | 100% | completed
BSM/QSM | ASP | 15 | 3 | 12 | 12 | 0 | 100% | NA | 11 | 1 | 91.67% | 7x Red line documents (word format)submitted to WPPL for ECO. 4x competed review. No ECO required. 2x pending ECO. - 19 Sep 18 - 1X ECO release today - 31 Oct 18 - 1X ECO release today 16Jan19 - In process for ECR Target closure: 31Jan19
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed
TUV Detectors | ASP | 26 | 7 | 19 | 19 | 0 | 100% | NA | 11 | 8 | 57.89% | 8x Red line documents (pdf format)submitted to WPPL for ECO. Target 30 Nov 18 | 8 | 2
| TSP | 6 | 0 | 6 | 6 | 0 | 100% | NA | 6 | 0 | 100% | completed | 2
PDA Detectors | ASP | 39 | 13 | 26 | 26 | 0 | 100% | NA | 7 | 19 | 26.92% | 19x Red line documents (pdf format)submitted to WPPL for ECO. Target 31 Oct 18, 31 Dec 2018 | 8 | 2
| TSP | 2 | 1 | 1 | 1 | 0 | 100% | NA | 1 | 0 | 100% | completed
Pump | ASP | 16 | 0 | 16 | 14 | 2 | 87.5% | NA | 6 | 8 | 42.86% | 5x Redline done (PDF format), 3x Redline done (Word format) 7 - In process of ECR 1 - In process of verification with Waters Target closure:31 Jan 2019
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed

## DUAL NOISE (hidden)
2489/TUV Tester Family | 2020-01-01 | 2020-02-01 | 2020-03-01 | 2020-04-01
TE82903-6 | 127 | 37 | 95 | 0
186017002 | 14 | 10 | 11
186248900 | 112 | 2 | 77
186002489 | 0 | 0 | 0
186015031 | 1 | 25 | 7
TE82903-8 | 51 | 75 | 25 | 0
186017002 | 16 | 11 | 0
186248900 | 23 | 34 | 14
186002489 | 0 | 0 | 0
186015031 | 12 | 30 | 11
TE82903-9 | 137 | 116 | 52 | 0
186017002 | 7 | 8 | 4
186248900 | 96 | 83 | 47
186002489 | 0 | 0 | 0
186015031 | 34 | 25 | 1
TE82903-10 | 8 | 10 | 7 | 0
186017002 | 0 | 1 | 0
186248900 | 8 | 7 | 5
186002489 | 0 | 0 | 0
186015031 | 0 | 2 | 2
TE82903-6 | 6.299% | 13.51% | 1.053% | #DIV/0!
Dual Noise | 8 | 5 | 1
TE82903-8 | 0 | 4% | 0% | #DIV/0!
Dual Noise | 0 | 3 | 0
TE82903-9 | 3.65% | 5.172% | 0% | #DIV/0!
Dual Noise | 5 | 6 | 0
TE82903-10 | 37.5% | 0% | 14.29% | #DIV/0!
Dual Noise | 3 | 0 | 1
Total Test | 323 | 238 | 179 | 0
Dual Noise | 16 | 14 | 2 | 0
FPY | 95.05% | 94.12% | 98.88% | #DIV/0!
% Contribution | 4.954% | 5.882% | 1.117% | #DIV/0!

## TE82526 Setup Schedule(CLOSE) (hidden)
UPDATE : | 2020-03-25
Pump Tray Test Capacity Enhancement |  |  | WW 0320 | WW 0420 | WW 0520 | WW 0620 | WW 0720 | WW 0820 | WW 0920 | WW 10'20 | WW 11'20 | WW 12'20 | WW 13'20 | WW 14'20
(OESR 193124-TRF MJUE-BJYJL4) |  |  | 13-17 Jan'20 | 20-24 Jan'20 | 27-31Jan'20 | 03-07 Feb'20 | 10-14 Feb'20 | 17-21 Feb'20 | 24-28 Feb'20 | 02-06 Mar'20 | 09-13 Mar '20 | 16-20 Mar'20 | 23-27 Mar'20 | 30 Mar-03 Apr '20
TE 82526-2 | Plan | Tester E.T.A. to Flex and | [5 empty] | FLEX PHYSICAL INVENTORY
verified complete
| Actual |  | 2020-01-16
Do Not Open
| Plan | Build Creform |  | Creform Table Setup
Actual
| Plan | Mount and Install the Pump Tray Tester
Actual
| Plan | Validation
| Actual | [5 empty] | 14 Feb 2020 Cal Team had inquiry about report |  | 2020-02-24
MP release the Balance Calibration report
2020-02-26
IQR Approved
TE 82526 Upgrade of CPU | Plan | Tester E.T.A. to Flex and
verified complete
| Actual |  | 2020-01-16
Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade
| Actual | [8 empty] | Upgraded Tester
| Plan | Validation
| Actual | [8 empty] | 05 Mar 2020 IQR Report submitted to WPPL | 12-Mar-2020 MJ IQR Approved
TE 82526-1 Upgrade of CPU | Plan | Tester E.T.A. to Flex and
verified complete
| Actual |  | 2020-01-16
Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade
| Actual | [9 empty] | 13-Mar-2020 Upgrade Complete
| Plan | Validation
| Actual | [9 empty] | 13-Mar-2020 Test completed | 18-Mar-2020 Report submitted by HZL to MJ for approval | 24-Mar-2020 MJ Approved IQR
TE 82526-4 Upgrade of CPU | Plan | Tester E.T.A. to Flex and
verified complete
| Actual |  | 2020-01-16
Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade | [10 empty] | 25-Mar-2020 Start Upgrading TE82826-4
| Actual | [11 empty] | 27-Mar-2020 Test complete and compiled report submitted to WPPL
| Plan | Validation
| Actual | [12 empty] | 30-Mar-2020 resubmit for correction of wrong entry 01-Apr-2020 IQR Approved by MJ

## OESR_v00 (hidden)
OESR ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | OESR # 193460 Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4: 74.1% Prime / 89.3 with Prime and Spare Q1-Q2 2020:Prime 79.3% / Prime + Spare 91.56% (HZ to update) 04Sep19 - HZL send reminder to MJ Milford 16Sep19 - MJ update the Milford is working concept to modernize this tester as duplicating it is impractical due to parts obsolecence 10Oct19: HZL to check for update from MJ for the concept of modernization of this tester 13Nov19:HZL to checked for update from Milford 27Nov19: HZL updated flowcell PN to MJ and prepare CEA for review. 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
2 | OESR # 193461 Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4 66.1% Prime / 74.7 with Prime and Spare Q1-Q2 2020: Prime 81.1% / Prime + Spare 85.5% (HZ to update) 04Sep19 : TE82520-3 CEA Approval loop 09Sep19: TE82520-3 CEA Approved/ HZL sent the CEA to Milford 10Oct19: HZL to check and advise the progress from Milford and respective completion date 16Oct19: ETD on Q2'20 advice from Chau Nguyen 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 31-Jan-20 31-Jul-20 | Open | Medium
3 | OESR # 193124 Additional Capacity for Alliance Pump Tray Tester (TE82526) 14Aug19: to put the TE number (JMD) (TE82526) 28Aug19: Milford in progress in building and updating the CPU to Windows 10 (HZ to update) 03Sep19: HZL check with Matt Brule for the tester BOM and pricing / CEA 11Sep19: Finance AT mentioned that able to use existing CEA 10Oct19: HZL to check and advise the progress from Milford 13Nov19: HZL to check for update 04Dec19: Pending for HZL comfirmation 11Dec19: On-track as per HZL feedback from Milford | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
4 | OESR # 190059 Alliance Final Test 3 Frequently Freeze During Final Test OESR # 193244 Milford OESR - 03 Apr 2019 04 Apr 2019 - Matt Brule request Flex to provide detail on Hang issue 05 Apr 2019 - Joel provide detail to Matt Brule for detail of Hang issue 04 Sep 2019 - Joel to monitor and update by 06 Sep 2019 06 Sep 2019 - JMD sent monitoring data to HZL (refer rto attach Alliance FT3 System Freeze Tracking Record) 09 Sep 2019 : HZL updated Matt Brule with attached Alliance FT3 System Freeze Tracking Record 18 Sep 2019: HZL last occurrence is 3 Sep 19, slow build for Alliance follows thus need to extend monitoring / with respective volume (min 100 units), HZL to inform Matt that Flex to further monitor this issue. 25 Sep 2019: JMD issue the tracker for 100 units to Prod'n and monitoring in progress | JMD HZL | 2019-09-04 | 13 Sep 19 27 Sep 19 30 Oct 19 | Monitoring | High
5 | OESR # 190079 TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. 15May2019 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04Sep2019 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11Sep2019 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18Sep2019 : JMD will inform YHC once EQID 0497 is available 19Sep2019: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25Sep2019 VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. | JMD HZL | 2019-09-04 | 2019-09-27 | Close | High
6 | OESR # 190102 - TE82988-2 Power Supply 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
7 | OESR 190104 - TE82988-2 Spare Parts holding at Senai - to request return back to WPPL 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10 Sep 2019 - Erzad update ETA back to WPPL by 23 Sep 2019 | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
8 | OESR # 190108 - TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep 2019 FLX Kallang JMD to follow-up to Edwin 05 Sep 2019 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
9 | OESR # S190109 - TSP having 735002531 with 735001613 does not have consistent content 04 Sep19: reviewed with HZL/EK/DA and JMD and agreed to update the OESR content and its Model PN. 05Sep19: JMD updated the OESR 190109 and send to HZL 18Sep19: HZL to checked for update 24Sep19: HZL to follow-up with E Turner 27Sep19: HZL to follow-up with E Turner reminder due as Waters Annual Audit for Flex 08Oct19 : HZL send reminder to E.T 08Oct19 - HZL ask ET if need to revise procedure/monitoring or uprev the next revision. 13Nov19: HZL to check for update | JMD HZL | 2019-09-04 | 06-Sep-19 10-Oct-19 30-Dec-19 | Monitoring | Medium
10 | OESR # 190117 - TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05 Sep 2019 Start 20 Aug 2019 used in 210000411 SN 010919110M 11 Sep 2019 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
11 | OESR #190118/ #190119 Reticule Plate Broken Target Mirror 19Aug19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30Aug19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04Sep19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23Sep19: SFC relelase PO to vendor 10Oct19 PM documentation approval in-progress 16Oct19: Part lead time is 4week from now (16Nov19) 20Oct19: GLY assembled and release to Production. | JMD SFC | 2019-09-04 | 2019-12-30 | Close | Medium
12 | OESR #190124 - 193544 Request for TE# for ICT Fixture 210001344 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05 Sep 2019 Erzad to send the documentation to JMD 06 Sep 2019 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11 Sep 2019 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
13 | OESR S190130 -193550 TE82882 PDA Final Test System Missing Spare OB PN 700011652 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18Sep2019 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25Sep2019 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26Sep2019 - TE82882-10SEP19-VAF_Approved and released tester for usage. | JMD HZL | 2019-09-04 | 06 Sep 19 30 Sep 19 | Close | High
14 | OESR # S190132 - TE82943-13 USB Device Error 03Sep19 JMD raised OESR for TE82943-13 USB Device Error 04Sep19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06Sep19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10Sep19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11Sep19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18Sep19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25Sep19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09Oct19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 10Oct19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10Oct19: Informed purchaser to release pay using Paypal. 16Oct19: JMD to follow-up E.T.A from buyer 23Oct19: Bliss target ETA to Flex is on 29Oct19 30Oct19: To install today (Shamsul) and update status 06Nov19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | JMD HZL | 2019-09-04 | 06 Sep 19 18 Sep 19 25 Sep 19 30 Oct 19 | Close | High
15 | OESR # 190133 Label Printing 04 Sep 2019 GLY to follow-up with DA | GLY / DA | 2019-09-04 | 2019-09-06 | Close | High
16 | OESR # S190156 TE82503-5 Injector Seal Pack Ramp and Decay Test Fail 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | JMD HZL | 2019-10-03 | 2019-10-21 | Close | High
17 | OESR # S190194 TE82943 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium
18 | OESR # S190195 TE82882 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium

## TE82943 VAL (Close 27May20) (hidden)
|  | WW'14 | [6 empty] | WW15 | [6 empty] | WW16 | [6 empty] | WW17 | [6 empty] | WW18 | [6 empty] | WW19 | [6 empty] | WW20 | [6 empty] | WW21
ACQUITY SDS TESTER |  | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun | Mon | Tue | Wed | Thu | Fri | Sat | Sun
Validation Plan | time axis 2020-03-30 to 2020-05-31
TE82882-11 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | bar: 2020-03-30; 2020-05-25
Actual | bar: 2020-03-30; 2020-05-25
Plan | Flex test and submit | bar: 2020-03-31; 2020-05-25
Actual | bar: 2020-03-31; 2020-05-25
Plan | Waters Approve /Release | 2020-04-02: GLY submitted to KCT; 2020-04-03: Pending Approval; 2020-04-06: Pending Approval; 2020-05-25: PH
Actual |  | 2020-04-07: IQR Approved; 2020-05-25: PH
TE82882-12 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | bar: 2020-03-31; 2020-05-25
Actual | bar: 2020-03-31; 2020-05-25
Plan | Flex test and submit | bar: 2020-04-01; 2020-05-25
Actual | bar: 2020-04-01; 2020-05-25
Plan | Waters Approve /Release | 2020-04-02: GLY submitted to KCY; 2020-04-03: Pending Approval; 2020-04-06: Pending Approval; 2020-05-25: PH
Actual |  | 2020-04-07: IQR Approved; 2020-05-25: PH
TE82882-13 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | 2020-05-11: GYC Plan; 2020-05-25: PH
Actual | bar: 2020-05-13; 2020-05-25
Plan | Flex test and submit | bar: 2020-04-07; 2020-05-14; 2020-05-25
Actual | bar: 2020-05-14; 2020-05-25
Plan | Waters Approve /Release | 2020-04-23: TUV under calibration; 2020-04-27: YCH Plan; 2020-05-25: PH
Actual |  | 2020-05-19: Approve and Release; 2020-05-25: PH
TE82882-14 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-30
Plan | Waters upgrade FCT Test to 1.9 | bar: 2020-04-07; 2020-05-25
Actual | bar: 2020-05-25
Plan | Flex test and submit | bar: 2020-04-08; 2020-05-25
Actual | bar: 2020-05-25
Plan | Waters Approve /Release | 2020-04-22: QSP release by Prod'n; 2020-04-23: Restart Validaton; 2020-04-24: Final Test PASS; 2020-04-27: FT Report Submitted; 2020-05-12: Pending Approval (WPPL); 2020-05-25: PH
Actual |  | 2020-05-13: Approved; 2020-05-25: PH
TE82882-15 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | 2020-04-13: On-Hold; 2020-05-25: PH
Actual | bar: 2020-05-25
Plan | Flex test and submit | bar: 2020-04-14; 2020-05-25
Actual | bar: 2020-05-25
Plan | Waters Approve /Release | 2020-04-22: QSP release by Prod'n; 2020-04-23: Restart Validaton; 2020-04-24: Final Test PASS; 2020-04-27: FT Report Submitted; 2020-05-12: Pending Approval (WPPL); 2020-05-25: PH
Actual |  | 2020-05-13: Approved; 2020-05-25: PH
TE82882-16 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | 2020-04-14: On-Hold; 2020-05-11: GYC Plan; 2020-05-25: PH
Actual | bar: 2020-05-13; 2020-05-25
Plan | Flex test and submit | bar: 2020-04-15; 2020-05-14; 2020-05-25
Actual | bar: 2020-05-14; 2020-05-25
Plan | Waters Approve /Release | 2020-04-27: YCH plan; 2020-05-25: PH
Actual |  | 2020-05-19: Approve and Release; 2020-05-25: PH
TE82882-17 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | 2020-04-20: On-Hold; 2020-05-18: GYC Plan; 2020-05-25: PH
Actual | bar: 2020-05-20; 2020-05-25
Plan | Flex test and submit | bar: 2020-04-21; 2020-05-20; 2020-05-25
Actual |  | 2020-05-22: Test record has been submitted to WPPL; 2020-05-25: PH
Plan | Waters Approve /Release | 2020-05-04: YCH Plan; 2020-05-25: PH; 2020-05-27: Approved
Actual | bar: 2020-05-25
TE82882-18 | bar: 2020-03-30 to 2020-04-03; 2020-04-06 to 2020-04-10; 2020-04-13 to 2020-04-17; 2020-04-20 to 2020-04-24; 2020-04-27 to 2020-05-31
Plan | Waters upgrade FCT Test to 1.9 | 2020-04-21: On-Hold; 2020-05-18: GYC Plan; 2020-05-25: PH
Actual | bar: 2020-05-20; 2020-05-25
Plan | Flex test and submit | bar: 2020-04-22; 2020-05-20; 2020-05-25
Actual |  | 2020-05-22: Test record has been submitted to WPPL; 2020-05-25: PH
Plan | Waters Approve /Release | 2020-05-04: YCH Plan; 2020-05-25: PH
Actual |  | 2020-05-25: PH; 2020-05-27: Approved
GYC PLAN 12 May 2020
[14 empty] | WW14 | WW17 | WW20 | WW21 | Remarks
[14 empty] | 30.Mar | 20.Apr | 11.May | 18.May
[12 empty] | Flex | TE-11 | IQ |  |  |  | IQR approved.
[13 empty] | TE-12 | IQ |  |  |  | IQR approved.
[13 empty] | TE-13 |  |  | IQ |  | Ed to update tester on 13May2020.
[13 empty] | TE-14 |  | IQ |  |  | IQR approved.
[13 empty] | TE-15 |  | IQ |  |  | IQR approved.
[13 empty] | TE-16 |  |  | IQ |  | Ed to update tester on 13May2020.
[13 empty] | TE-17 |  |  |  | IQ
[13 empty] | TE-18 |  |  |  | IQ

## Alliance Forecast (hidden)
PART NUMBER | Material Description | Q1-20 | Q2-20 | Q3-20 | Q4-20 | 2020 | Q1-21 | Q2-21 | Q3-21 | Q4-21
186269505 | e2695 Separations Module w/o HC | 336 | 150 | 112 | 189 | 787 | 183 | 211 | 210 | 254
186269506 | e2695 Separations Module w/HC | 820 | 450 | 714 | 768 | 2752 | 694 | 754 | 783 | 959
| Alliance Model | 1156 | 600 | 826 | 957 | 3539 | 877 | 965 | 993 | 1213
| Monthly | 385.3333 | 200 | 275.3333 | 319 | 1179.6667 | 292.3333 | 321.6667 | 331 | 404.3333
| Weekly | 88.9231 | 46.1538 | 63.5385 | 73.6154 | 272.2308 | 67.4615 | 74.2308 | 76.3846 | 93.3077
| Daily | 17.7846 | 9.2308 | 12.7077 | 14.7231 | 54.4462 | 13.4923 | 14.8462 | 15.2769 | 18.6615
| CCD Capacity per day with 2 stations | 32 | 32 | 32 | 32 | NA | 32 | 32 | 32 | 32
| Utilization | 55.58% | 28.85% | 39.71% | 46.01% | NA | 42.16% | 46.39% | 47.74% | 58.32%

## Waters PC List (hidden)
Waters PC List |  |  | Image Created | Update Image Date | Remarks
IP Address | Name | Description | Date
10.249.26.2 | Waters SISNL1 | Solectron Instrument Serial Number 1 | 2019-03-12
10.249.26.3 | Waters SPNL1 | Solectron Part Number Labeling 1 | 2019-03-12
10.249.26.4 | WATERSCOFP1 | WATERS Certificate Of Funtional Performance 1 | 2019-03-12
10.249.26.9 | TE 82800 - 2 | Acquity Actuator Tester #2 | 2019-02-13
10.249.26.10 | TE 82800 - 3 | Acquity Actuator Tester #3 | 2019-03-12
10.249.26.11 | TE 82943 - 11 | Acquity SDS Tester #11 | 2019-02-13 | 2019-12-13
10.249.26.12 | TE 82943 - 12 | Acquity SDS Tester #12 | 2019-03-12 | 2019-12-13
10.249.26.13 | TE 82943 - 13 | Acquity SDS Tester #13 | 2019-02-13 | 2019-12-13
10.249.26.14 | TE 82943 - 14 | Acquity SDS Tester #14 | 2019-02-13 | 2019-12-13
10.249.26.15 | TE 82943 - 15 | Acquity SDS Tester #15 | 2019-02-13 | 2019-12-17
10.249.26.16 | TE 82943 - 16 | Acquity SDS Tester #16 | 2019-02-13 | 2019-12-18
10.249.26.17 | TE 82943 - 17 | Acquity SDS Tester #17 | 2019-03-12 | 2019-12-17
10.249.26.18 | TE 82943 - 18 | Acquity SDS Tester #18 | 2019-03-13 | 2019-12-17
10.249.26.55 | Genealogy SDS | Acquity SDS Genealogy Stn (BSM/QSM) | 2019-03-12
10.249.26.40 | TE 82903 - 6 | TUV Final Test Station | NA |  | Unable to activate Acronis
10.249.26.42 | TE 82903 - 8 | TUV Final Test Station | 2019-03-14
10.249.26.46 | TE 82903 - 9 | TUV Final Test Station | 2019-03-14
10.249.26.41 | TE 82903 - 10 | TUV Final Test Station | 2019-03-14
10.249.26.45 | TE 82964 - 1 | TUV Pretest System | 2019-03-13
10.249.26.75 | TE 82969 - 1 | PDA Linearity Tester | 2019-03-13 | 2019-10-11
10.249.26.76 | TE 83057-4 | PDA Download Station | 2019-03-13
10.249.26.79 | TE 82882 - 8 | 2998/PDA Final Test 8 | 2019-03-13 | 2019-08-22
10.249.26.80 | TE 82882 - 9 | 2998/PDA Final Test 9 | 2019-03-13 | 2019-08-22
10.249.26.81 | TE 82882 - 10 | 2998/PDA Final Test 10 | 2019-03-13 | 2019-08-26
10.249.26.82 | TE 82882 - 11 | 2998/PDA Final Test 11 | 2019-03-13 | 2020-03-06
10.249.26.86 | TE 82882 -13 | PDA Final Test System 13 | 2019-03-14 | 2020-03-06 | Backup thru Windows OS
10.249.26.85 | TE 82882 -14 | PDA Final Test System 14 | 2019-03-14 | 2020-03-06 | Backup thru Windows OS
10.249.26.63 | TE 82918 - 2 | 2998 Fine Alignment Sys | 2019-03-14
10.249.26.64 | TE 82868 - 2 | 2998 Aperture Alignment System | 2019-03-14
10.249.26.65 | TE 82951 - 2 | 2998 Mask Alignment Sys | 2019-03-14
10.249.26.66 | TE 82916 - 2 | Flowcell Alignment Fixture | 2019-03-14
10.249.26.68 | TE 82973 - 1 | 2998 Flowcell Noise & Drift System | NA |  | CD-ROM unable to read
10.249.26.72 | TE 82917 - 1 | 2998 Rough Alignment | 2019-03-14
10.249.26.73 | TE 82880 - 2 | PDA Rough Alignment Sys | 2019-03-14
10.249.26.74 | TE 82881 - 2 | PDA Fine Alignment Stn | 2019-03-14
10.249.26.56 | Genealogy SDS | Acquity SDS Genealogy Stn (Detector) | 2019-03-14
10.249.26.58 | TE 83096 - 2 | PDA Universal Optic Bench Alignment | 2019-03-15 |  | Backup thru Windows OS
10.249.26.59 |  | Universal PDA Alignment | 2019-03-16 |  | Backup thru Windows OS
10.249.26.50 | TE 82534 - 2 | 515 Pump Final Test | 2019-02-12
10.249.26.51 | TE 82705 - 2 | 1500 pump Final Test | 2019-03-12
10.249.26.20 | Final Test - 1 | Waters Alliance Test Station | 2019-02-12
10.249.26.21 | Final Test - 2 | Waters Alliance Test Station | 2019-03-13
10.249.26.22 | Final Test - 3 | Waters Alliance Test Station | 2019-02-11
10.249.26.23 | Final Test - 4 | Waters Alliance Test Station | 2019-02-12
10.249.26.24 | Final Test - 5 | Waters Alliance Test Station | 2019-03-13
10.249.26.25 | Final Test - 6 | Waters Alliance Test Station | 2019-02-11
10.249.26.26 | TE 82526 - 1 | Waters Alliance Test Station (TE82526-1) | 2019-03-13 |  | New PC
10.249.26.27 | TE 82526 | Waters Alliance Test Station (TE82526) | 2019-03-13 |  | New PC
10.249.26.28 | TE 82526 - 4 | Waters Alliance Test Station (TE82526-4) | 2019-03-13 |  | New PC
10.249.26.39 | TE 82526 - 2 | Waters Alliance Test Station (TE82526-2) | New PC
10.249.26.29 | TE 82519 | Waters Alliance Test Station ( Pump Test - 1 ) | Computer auto shutdown
10.249.26.32 | TE 82573 - 1 | E-Alliance Rear Panel Test | 2019-03-12
10.249.26.34 | AutoLoader - 2 | Alliance AutoLoader Station 2 | 2019-03-12
10.249.26.33 | AutoLoader - 3 | Alliance AutoLoader Station 3 | 2019-03-12
10.249.26.35 | TE 82520 - 1 | Alliance Carousel/Carrier Alignment Station -1 | Computer auto shutdown
10.249.26.36 | TE 82520 - 2 | Alliance Carousel/Carrier Alignment Station -2 | 2019-03-13 |  | Backup thru Windows OS
10.249.26.37 | TE 82546 | Front Panel Station | 2019-03-13
10.249.26.38 | TE 0084018-1 | Heater Cooler Tester | 2019-12-17 |  | Backup thru Windows OS
10.249.26.90 | QA Workstation | QA Workstation/SQT.ORT | 2019-02-11
10.249.26.9 | TE82800-2 | Acquity Solvent Manager Pump Test Station | 2019-05-29
10.249.26.10 | TE82800-3 | Acquity Solvent Manager Pump Test Station | 2019-05-29
10.249.26.11 | TE82943-11 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.12 | TE82943-12 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.13 | TE82943-13 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.14 | TE82943-14 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.15 | TE82943-15 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.16 | TE82943-16 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.17 | TE82943-17 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.18 | TE82943-18 | Acquity Solvent Manager Test Station | 2019-05-29
10.249.26.20 | Final-Test-1 | Alliance Final Test Station | 2019-05-29
10.249.26.21 | Final-Test-2 | Alliance Final Test Station | SERVICE PACK 2
10.249.26.22 | Final-Test-3 | Alliance Final Test Station | 2019-05-29
10.249.26.23 | Final-Test-4 | Alliance Final Test Station | 2019-05-29
10.249.26.24 | Final-Test-5 | Alliance Final Test Station | 2019-05-29
10.249.26.26 | Pumptray-Test-1 | Alliance Pumptray Tester | WINDOWS 7
10.249.26.27 | Pumptray-Test-2 | Alliance Pumptray Tester | 2019-05-29
10.249.26.28 | Pumptray-Test-3 | Alliance Pumptray Tester | 2019-05-29
10.249.26.29 | Pump-Test-1 | Alliance Pump Tester | SERVICE PACK 2
10.249.26.32 | RearPanel-Tester | Autoloader and Rear Panel Tester | 2019-05-29
10.249.26.33 | AutoLoader-3 | Alliance AutoLoader 3 | 2019-05-29
10.249.26.34 | AutoLoader-2 | Alliance AutoLoader 2 | 2019-05-29
10.249.26.35 | TE82520-1 | Alliance CCD Test Station | 2019-05-29
10.249.26.40 | TE82903-6 | TUV Final Test System | 2019-05-29
10.249.26.42 | TE82903-8 | TUV Final Test System | 2019-05-29
10.249.26.43 | TE82903-9 | TUV Final Test System | 2019-05-29
10.249.26.44 | TE82903-10 | TUV Final Test System | 2019-05-29
10.249.26.45 | TE82964-1 | Detector Pretest System | 2019-05-29
10.249.26.50 | TE82534-(1)2 | 515 Test | 2019-05-29
10.249.26.51 | TE82705-2 | Breeze Test | 2019-05-29
10.249.26.55 | Genealogy-SDS | Acquity BSM/QSM Genealogy Input Station | 2019-05-29
10.249.26.56 | Genealogy-DET | Detector Genealogy Input Station | 2019-05-29
10.249.26.63 | TE82918-2 | 2998 Fine Alignment System | 2019-05-29
10.249.26.64 | TE82868-2 | 2998 Aperture Alignment System | 2019-05-29
10.249.26.65 | TE82951-2 | 2998 Mask Alignment System | 2019-05-29
10.249.26.68 | TE82973-1 | 2998 Flowcell Noise and Drift System | 2019-05-29
10.249.26.71 | TE82916-(2)1 | 2998 Flowcell Alignment System | 2019-05-29
10.249.26.72 | TE82917-1 | 2998 Rough Alignment System | 2019-05-29
10.249.26.73 | TE82880-2 | PDA Rough Alignment System | 2019-05-29
10.249.26.74 | TE82881-2 | PDA Fine Alignment System | 2019-05-29
10.249.26.75 | TE82969-1 | PDA Linearity Test System | 2019-05-29
10.249.26.76 | TE83057-4 | PDA Download Station | 2019-05-29
10.249.26.30 | - | HP PC with RS232 Port | NOT IN USED
10.249.26.78 | TE82882-2 | 2998/PDA Final Test System | NOT IN USED
10.249.26.1 | solectronmfgnew | Windows Server 2012 R2 | Intel Xeon
10.249.26.2 | SOLECTRONISNL1 | Solectron Instrument Serial Number Label 1
10.249.26.3 | SOLECTRONPNL1 | Solectron Part Number Label 1
10.249.26.4 | SOLECTRONCOFP1 | Solectron Certificate Of Functional Performance 1

## Calibration Master List (hidden)
S/N | Equipment ID | Equipment Description | Manufacturer | Model | Equipment Serial Number | Interval | Calibration Date | Due Date | Vendor | Equipment-Owner | Customized Equip. | Customer ID | Customer | Location | Active | Remarks
1 | 0196 | Digital Thermometer with Probe (K type) | Fluke | 51 II | 89280141 | 12 Months | 2019-07-30 | 2020-07-30 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters | Waters - Production | Y | N/A
2 | 0206 | Electronic Weigh Balance | AND | HR200 | 12305757 | 06 Months | 2020-01-13 | 2020-07-13 | Micro Precision | Joel Mamalateo Deduyo | N | 90891 | Waters | Waters - Production | Y | N/A
3 | 0207 | Column Heater | Waters | WAT091727 | M03CHM015R | 12 Months | 2020-04-03 | 2021-04-03 | Internal | Joel Mamalateo Deduyo | Y | 81008 | Waters | Waters 5th Floor | Y | N/A
4 | 0208 | Column Heater | Waters | WAT091727 | E01CHM203R | 12 Months | 2020-04-03 | 2021-04-03 | Internal | Joel Mamalateo Deduyo | Y | 81009 | Waters | Waters 5th Floor | Y | N/A
5 | 0209 | Column Heater | Waters | WAT091727 | D01CHM128R | 12 Months | 2020-01-23 | 2021-01-23 | Internal | Joel Mamalateo Deduyo | Y | 81010 | Waters | Waters - Production | Y | N/A
6 | 0210 | Column Heater | Waters | WAT091727 | K03CHM938R | 12 Months | 2020-01-23 | 2021-01-23 | Internal | Joel Mamalateo Deduyo | Y | 81011 | Waters | Waters - Production | Y | N/A
7 | 0211 | Column Heater | Waters | WAT091727 | B03CHM551R | 12 Months | 2020-01-24 | 2021-01-24 | Internal | Joel Mamalateo Deduyo | Y | 81012 | Waters | Waters - Production | Y | N/A
8 | 0212 | Column Heater | Waters | WAT091727 | F00CHM528R | 12 Months | 2020-01-24 | 2021-01-24 | Internal | Joel Mamalateo Deduyo | Y | 81013 | Waters | Waters - Production | Y | N/A
9 | 0213 | Column Heater | Waters | WAT091727 | F00CHM531R | 12 Months | 2020-01-29 | 2021-01-29 | Internal | Joel Mamalateo Deduyo | Y | 81014 | Waters | Waters - Production | Y | N/A
10 | 0216 | Electronic Weigh Balance | AND | HR200 | 12200404 | 06 Months | 2020-01-13 | 2020-07-13 | Micro Precision | Joel Mamalateo Deduyo | N | 90038 | Waters | Waters - Production | Y | N/A
11 | 0217 | Electronic Weigh Balance | AND | HR200 | 12303279 | 06 Months | 2020-01-13 | 2020-07-13 | Micro Precision | Joel Mamalateo Deduyo | N | 90018 | Waters | Waters - Production | Y | N/A
12 | 0221 | 1500 Series Pump Final Test Station | Waters | Pump | TE82705-2 | 06 Months | 2019-11-26 | 2020-05-26 | Internal | Joel Mamalateo Deduyo | Y | TE82705-2 | Waters | Waters - Production | Y | N/A
13 | 0222 | 515 Final Test Station | Waters | Pump | TE82534 | 06 Months | 2019-11-26 | 2020-05-26 | Internal | Joel Mamalateo Deduyo | Y | TE82534 | Waters | 5th Floor | Y | N/A
14 | 0223 | HPMV Tester | Waters | NA | TE82625-1 | 06 Months | 2020-01-16 | 2020-07-16 | Internal | Joel Mamalateo Deduyo | Y | TE82625-1 | Waters | Waters - Production | Y | N/A
15 | 0224 | 2690 Pump Head Tester | Waters | NA | TE82519 | 06 Months | 2019-11-27 | 2020-05-27 | Internal | Joel Mamalateo Deduyo | Y | TE82519 | Waters | 5th Floor | Y | N/A
16 | 0225 | Injector Seal Pack Tester | Waters | TE82503-4 | TE82503-4 | 06 Months | 2020-01-16 | 2020-07-16 | Internal | Joel Mamalateo Deduyo | Y | TE82503-4 | Waters | Waters - Production | Y | N/A
17 | 0226 | Electronic Weigh Balance | AND | HR200 | 12315717 | 06 Months | 2020-01-13 | 2020-07-13 | Micro Precision | Joel Mamalateo Deduyo | N | 90913 | Waters | Waters - Production | Y | N/A
18 | 0233 | Column Heater | Waters | WAT038040 | H01CHM352M | 12 Months | 2020-01-29 | 2021-01-29 | Internal | Joel Mamalateo Deduyo | Y | 81020 | Waters | Waters - Production | Y | N/A
19 | 0234 | Column Heater | Waters | WAT038040 | M00CHM910M | 12 Months | 2020-01-29 | 2021-01-29 | Internal | Joel Mamalateo Deduyo | Y | 81021 | Waters | Waters - Production | Y | N/A
20 | 0235 | Column Heater | Waters | WAT038040 | H01CHM335M | 12 Months | 2020-01-29 | 2021-01-29 | Internal | Joel Mamalateo Deduyo | Y | 81022 | Waters | Waters - Production | Y | N/A
21 | 0236 | Column Heater | Waters | WAT038040 | MX6JM7863M | 12 Months | 2020-01-30 | 2021-01-30 | Internal | Joel Mamalateo Deduyo | Y | 81018 | Waters | Waters - Production | Y | N/A
22 | 0237 | Column Heater | Waters | WAT038040 | L97CHM711M | 12 Months | 2020-01-30 | 2021-01-30 | Internal | Joel Mamalateo Deduyo | Y | 81019 | Waters | Waters 5th Floor | Y | N/A
23 | 0238 | Column Heater | Waters | WAT038040 | MX6KM9444M | 12 Months | 2020-04-02 | 2021-04-02 | Internal | Joel Mamalateo Deduyo | Y | 81017 | Waters | Waters 5th Floor | Y | N/A
24 | 0239 | Column Heater | Waters | WAT038040 | MX5MM5993M | 12 Months | 2020-01-30 | 2021-01-30 | Internal | Joel Mamalateo Deduyo | Y | 81023 | Waters | Waters - Production | Y | N/A
25 | 0244 | Column Heater | Waters | WAT038040 | MX6JM7867M | 12 Months | 2020-01-30 | 2021-01-30 | Internal | Joel Mamalateo Deduyo | Y | 81016 | Waters | Waters - Production | Y | N/A
26 | 0245 | Digital Multimeter | Fluke | 110 | 85730619 | 12 Months | 2019-12-16 | 2020-12-16 | Micro Precision | Joel Mamalateo Deduyo | N | 96263 | Waters | Waters - Production | Y | N/A
27 | 0246 | Digital Multimeter | Fluke | 179 | 90240092 | 12 Months | 2020-02-13 | 2021-02-13 | Micro Precision | Joel Mamalateo Deduyo | N | 91368 | Waters | Waters - Production | Y | N/A
28 | *** 0249 | Electrical Safety Analyzer | Quadtech | Guardian 6000 | 5037655 | 12 Months | 2020-01-08 | 2021-01-08 | Micro Precision | Joel Mamalateo Deduyo | N | 95037 | Waters | Waters 5th Floor | Y | N/A
29 | 0256 | Pressure Calibrator | Ametek | APC10KGINDG | 9093205 | 12 Months | 2019-11-22 | 2020-11-21 | Micro Precision | Joel Mamalateo Deduyo | N | WT-TE-0331 | Waters | 5th Floor | Y | N/A
30 | 0259 | Injector Seal Pack Tester | Waters | TE82503-5 | TE82503-5 | 06 Months | 2020-02-27 | 2020-08-27 | Internal | Joel Mamalateo Deduyo | Y | TE82503-05 | Waters | Waters - Production | Y | N/A
31 | 0278 | Electronic Weigh Balance | AND | HR200 | 12316922 | 06 Months | 2020-01-13 | 2020-07-13 | Micro Precision | Joel Mamalateo Deduyo | N | 95206 | Waters | Waters - Production | Y | N/A
32 | 0279 | High Voltage Probe | Fluke | 80K-6 | 89200011 | 12 Months | 2019-09-24 | 2020-09-24 | SGS | Joel Mamalateo Deduyo | N | NA | Waters | 5th Floor | Y | N/A
33 | 0280 | Clamp Meter | Kyoritsu | 2033 | 66386 | 12 Months | 2019-07-30 | 2020-07-30 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters | Waters - Production | Y | N/A
34 | *** 0360 | Electrical Safety Analyzer | Quadtech | Guardian 6000 | 6069766 | 12 Months | 2019-11-28 | 2020-11-28 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters | Waters 5th Floor | Y | N/A
35 | 0431 | Picoammeter | Keithley | 485 | 435110 | 12 Months | 2020-02-18 | 2021-02-18 | Keysight Technologies | Joel Mamalateo Deduyo | N | NA | Waters | 5th Floor | Y | N/A
36 | *** 0432 | Electrical Safety Analyzer | Quadtech | Guardian 6000 Plus | 08280865 | 12 Months | 2019-10-03 | 2020-10-03 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters | Waters - Production | Y | N/A
37 | 0458 | Auto Pressure (Flowcell Tester) | Waters | NA | TE-82475-2 | 12 Months | 2019-04-11 | 2020-04-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | 5th Floor | Y | Report submitted but not yet update on the system
38 | 0481 | Liquilaz | Particle Measuring Systems | E20P | 51053 | 12 Months | 2019-08-21 | 2020-08-21 | Particle Measuring Systems | Joel Mamalateo Deduyo | N | NA | Waters | 5th Floor | Y | N/A
39 | 0497 | Electronic Weigh Balance | AND | HR200 | 12326186 | 06 Months | 2020-01-31 | 2020-07-31 | Ming Deng | Joel Mamalateo Deduyo | N | 90610 | Waters | Waters - Production | Y | N/A
40 | 0499 | Electronic Weigh Balance | AND | HR200 | 12324733 | 06 Months | 2019-12-16 | 2020-06-16 | Micro Precision | Joel Mamalateo Deduyo | N | 90605 | Waters | Waters - Production | Y | N/A
41 | 0500 | Electronic Weigh Balance | AND | HR200 | 12326184 | 06 Months | 2019-10-23 | 2020-04-23 | Micro Precision | Joel Mamalateo Deduyo | N | 90608 | Waters | Waters - Production | Y | On-schedule
42 | 0501 | Electronic Weigh Balance | AND | HR200 | 12326187 | 06 Months | 2019-12-16 | 2020-06-16 | Micro Precision | Joel Mamalateo Deduyo | N | 90604 | Waters (1500 Pump) | Waters - Production | Y | N/A
43 | 0502 | Electronic Weigh Balance | AND | HR200 | 12326180 | 06 Months | 2019-12-16 | 2020-06-16 | Micro Precision | Joel Mamalateo Deduyo | N | 90603 | Waters | Waters - Production | Y | N/A
44 | 0503 | Electronic Weigh Balance | AND | HR200 | 12326190 | 06 Months | 2019-12-16 | 2020-06-16 | Micro Precision | Joel Mamalateo Deduyo | N | 90600 | Waters | Waters - Production | Y | N/A
45 | 0504 | Electronic Weigh Balance | AND | HR200 | 12326279 | 06 Months | 2019-12-16 | 2020-06-16 | Micro Precision | Joel Mamalateo Deduyo | N | 90607 | Waters | Waters - Production | Y | N/A
46 | 0592 | Torque Wrench | Snap-on | QD2R1000 | 1009032663 | 06 Months | 2020-03-04 | 2020-09-04 | Ming Deng | Phani | N | NA | Waters | 5th Floor | Y | N/A
47 | 0594 | Dial Depth Gauge | Mitutoyo | 2904S / 7217S | LQY537 / 000943 | 06 Months | 2020-03-06 | 2020-09-06 | Ming Deng | Phani | N | NA | Waters | Waters - Production | Y | N/A
48 | 0595 | Dial Indicator | Mitutoyo | 2514S | JDY576 | 06 Months | 2020-03-06 | 2020-09-06 | Ming Deng | Phani | N | NA | Waters | Waters - Production | Y | N/A
49 | 0597 | Mass Flow Meter | AALBORG | GFM17 | 170609-1 | 12 Months | 2019-07-15 | 2020-07-15 | Caltek | Joel Mamalateo Deduyo | N | NA | Waters | Waters - Production | Y | N/A
50 | 0877 | Torque Wrench | Snap-on | QD1R200 | 0408302147 | 06 Months | 2020-03-04 | 2020-09-04 | Ming Deng | Gary Lai Yam Bee | N | NA | Waters | Waters - Production | Y | N/A
51 | 1109 | Dual Wavelength Detector | Waters | TUV Detector | K05UPT479M | 12 Months | 2020-01-31 | 2021-01-31 | Internal | Joel Mamalateo Deduyo | Y | 91897 | Waters | Waters - Production | Y | N/A
52 | 1394 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E726A | 12 Months | 2019-06-12 | 2020-06-12 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
53 | 1395 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E727A | 12 Months | 2019-07-10 | 2020-07-10 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
54 | 1396 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E728A | 12 Months | 2019-06-28 | 2020-06-28 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
55 | 1397 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E729A | 12 Months | 2019-07-19 | 2020-07-19 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
56 | 1398 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E730A | 12 Months | 2019-07-11 | 2020-07-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
57 | 1399 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E731A | 12 Months | 2019-06-03 | 2020-06-03 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
58 | 1400 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E732A | 12 Months | 2019-06-11 | 2020-06-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
59 | 1401 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E733A | 12 Months | 2019-06-13 | 2020-06-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
60 | 1402 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E734A | 12 Months | 2019-07-19 | 2020-07-19 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
61 | 1403 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E735A | 12 Months | 2019-07-10 | 2020-07-10 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
62 | 1404 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E736A | 12 Months | 2019-06-21 | 2020-06-21 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
63 | 1405 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E737A | 12 Months | 2019-07-11 | 2020-07-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
64 | 1428 | 2489 UV / Visible Detector | Waters | 186248900 | H1387E710A | 12 Months | 2019-07-11 | 2020-07-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
65 | 1429 | 2489 UV / Visible Detector | Waters | 186248900 | D1387E759A | 12 Months | 2019-04-10 | 2020-04-10 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | Report submitted but not yet update on the system
66 | 1430 | 2489 UV / Visible Detector | Waters | 186248900 | D1387E763A | 12 Months | 2019-04-11 | 2020-04-11 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | Report submitted but not yet update on the system
67 | 1431 | 2489 UV / Visible Detector | Waters | 186248900 | D1387E765A | 12 Months | 2019-04-12 | 2020-04-12 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | Report submitted but not yet update on the system
68 | 1432 | 2489 UV / Visible Detector | Waters | 186248900 | D1387E766A | 12 Months | 2020-04-03 | 2021-04-03 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
69 | 1433 | 2489 UV / Visible Detector | Waters | 186248900 | C1387E135M | 12 Months | 2019-04-12 | 2020-04-12 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | Report submitted but not yet update on the system
70 | *** 1492 | Electrical Safety Analyzer | Chroma | 19032 | 19032G000182 | 12 Months | 2019-10-26 | 2020-10-26 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters | Waters - Production | Y | N/A
71 | 1515 | 2489 UV / Visible Detector | WATERS | 186248900 | B1587E607A | 12 Months | 2019-12-26 | 2020-12-26 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
72 | 1516 | 2489 UV / Visible Detector | WATERS | 186248900 | B1587E608A | 12 Months | 2019-12-27 | 2020-12-27 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
73 | 1519 | 2489 UV / Visible Detector | WATERS | 186248900 | B1587E609A | 12 Months | 2019-12-27 | 2020-12-27 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
74 | 1520 | 2489 UV / Visible Detector | WATERS | 186248900 | B1587E610A | 12 Months | 2019-12-26 | 2020-12-26 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
75 | 1577 | Macbeth Lighting | X-Rite | Judge II | C61093261217 | 12 Months | 2019-09-04 | 2020-09-04 | Xrite | Hayati Binte Isshak | N | NA | Waters | Changi South - IQA | Y | N/A
76 | 1781 | Column Heater | WATERS | WAT091727 | C13CHM795R | 12 Months | 2019-07-17 | 2020-07-17 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
77 | 1782 | Column Heater | WATERS | WAT091727 | D12CHM299R | 12 Months | 2019-07-17 | 2020-07-17 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
78 | 1783 | Column Heater | WATERS | WAT091727 | K14CHM398R | 12 Months | 2019-07-18 | 2020-07-18 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
79 | 1784 | Column Heater | WATERS | WAT091727 | L14CHM438R | 12 Months | 2019-07-19 | 2020-07-19 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
80 | 1787 | Column Heater | WATERS | WAT091727 | D15CHM641R | 12 Months | 2019-07-19 | 2020-07-19 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters | Waters - Production | Y | N/A
81 | 2035 | Electronic Weigh Balance | AND | HR200 | 12336741 | 06 Months | 2019-10-23 | 2020-04-23 | Micro Precision | Joel Mamalateo Deduyo | N | NA | Waters (BSM/QSM ) | Waters - Production | Y | On-schedule
82 | 2092 | Electronic Weigh Balance | AND | HR200 | 12337042 | 06 Months | 2019-10-23 | 2020-04-23 | Micro Precision | Joel Mamalateo Deduyo | N | 92207 | Waters (Pump Series) | Waters - Production | Y | On-schedule
83 | *** 2094 | Load Box | Chroma | S14-03 | 14354695 | 12 Months | 2019-12-12 | 2020-12-12 | Micro Precision | Joel Mamalateo Deduyo | N | 86080 | Waters (Generic) | Waters - Production | Y | N/A
84 | *** 2096 | Load Box | Chroma | S14-03 | 14354691 | 12 Months | 2019-05-08 | 2020-05-08 | Micro Precision | Joel Mamalateo Deduyo | N | 86079 | Waters (Generic) | Waters - Production | Y | N/A
85 | *** 2097 | Load Box | Chroma | S14-03 | 7322788 | 12 Months | 2019-05-08 | 2020-05-08 | Micro Precision | Joel Mamalateo Deduyo | N | 86085 | Waters (Generic) | Waters - Production | Y | N/A
86 | 2217 | Benchtop Measuring Machine | Quality Vision International | StarLite 150 | RSL1002609 | 12 Months | 2019-04-23 | 2020-04-22 | Micro Precision | Joel Mamalateo Deduyo | N | 95515 | Waters | Waters - Production | Y | On-schedule
87 | 2348 | 2489 UV / Visible Detector | Waters | 186248900 | C1887E101A | 12 Months | 2020-03-20 | 2021-03-20 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
88 | 2349 | 2489 UV / Visible Detector | Waters | 186248900 | C1887E103A | 12 Months | 2020-03-20 | 2021-03-20 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
89 | 2350 | 2489 UV / Visible Detector | Waters | 186248900 | C1887E106A | 12 Months | 2020-03-13 | 2021-03-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
90 | 2351 | 2489 UV / Visible Detector | Waters | 186248900 | D1887E275A | 12 Months | 2019-04-08 | 2020-04-08 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | Report submitted but not yet update on the system
91 | 2352 | Column Heater Module | Waters | WAT038040 | M17CHM834M | 12 Months | 2019-06-13 | 2020-06-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
92 | 2353 | Column Heater Module | Waters | WAT038040 | M17CHM835M | 12 Months | 2019-06-13 | 2020-06-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
93 | 2354 | Column Heater Module | Waters | WAT038040 | M17CHM839M | 12 Months | 2019-06-13 | 2020-06-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
94 | 2355 | Column Heater Module | Waters | WAT038040 | M17CHM840M | 12 Months | 2019-06-13 | 2020-06-13 | Internal | Joel Mamalateo Deduyo | Y | NA | Waters (eAlliance) | Waters - Production | Y | N/A
95 | 2769 | Electronic Weigh Balance | AND | HR200 | 12303631 | 06 Months | 2020-02-07 | 2020-08-07 | Micro Precision | Joel Mamalateo Deduyo | N | 82184 | Waters | Waters - Production | Y | N/A

## OESR's (hidden)
Log Number | Engineering Group | ASP TSP No | TE Tool No | STATUS | Problem Summary | Justification
171714 | Test Engineering | WAT242102 | N/A | OPEN | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | Part Number not mention in the TSP by default is considered untested part and subject to non compliance during audit
193814 | Test Engineering | 735001613 | TE82882 | OPEN | The final test record for Detector(TE82882-All and TE82969-1) and SDS ( TE82943-All) has incomplete of 'start time' and 'end time' print out. The second of time is disappeared from test record. Please kindly update the final test record template to print out the complete time, including the seconds on test record. Please refer to attachment on page 2 for the details. | Trunacated part of any test station report in any printed report is non compliance to good documentation practices for both Flex
193244 | Test Engineering | WAT271129 | 186000139 | OPEN | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. | Tester shutdown is still within acceptable numbers but post frustration to test techncian performing the task. This is very evident if you are testing 4 untis at the same time and ystem hang. Though data was collected between Aliance Final Test1, 3 and 5, number of hang happens more in Final Test 5, as when ew highlight, more of the hang happened in tester #3.

## 193244 (hidden)
Log Number | Engineering Group | ASP TSP No | TE Tool No | STATUS | Problem Summary | Justification
171714 | Test Engineering | WAT242102 | N/A | OPEN | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | Part Number not mention in the TSP by default is considered untested part and subject to non compliance during audit
193814 | Test Engineering | 735001613 | TE82882 | OPEN | The final test record for Detector(TE82882-All and TE82969-1) and SDS ( TE82943-All) has incomplete of 'start time' and 'end time' print out. The second of time is disappeared from test record. Please kindly update the final test record template to print out the complete time, including the seconds on test record. Please refer to attachment on page 2 for the details. | Trunacated part of any test station report in any printed report is non compliance to good documentation practices for both Flex
193244 | Test Engineering | WAT271129 | 186000139 | OPEN | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. | Tester shutdown is still within acceptable numbers but post frustration to test techncian performing the task. This is very evident if you are testing 4 untis at the same time and ystem hang. Though data was collected between Aliance Final Test1, 3 and 5, number of hang happens more in Final Test 5, as when ew highlight, more of the hang happened in tester #3.

## Pump Head (HotPart) (hidden)
Update as if 04 Dec 2019
Below data per latest sound measurement conducted today.
Present : Aldwin. Ivy, Desmond, Guan Yee, Eugene
As per discussion today the data gathered today shall be used as reference final reference for the Sound comparison of Hot Part vs Known Good Part..
Observations:
The sound level of the Hot Part Motor (68.2 dB) is still high when installed to production unit (Hot Part Motor + New#6)
There is an average of 5.0 dB difference between the Hot Part compared to 6 samples of production units.
After re-sitting ang tightening the pump coupling shaft (WAT271467) of the Hot Part Motor, the sound level is still higher.
SN | Date | Comparison | General Sound Level ( A ) dB (min) | General Sound Level ( A ) dB (max)
1 | 2019-12-04 | Hot Part (1st try) | 66.4 | 68.4
2 | 2019-12-04 | New#1-1 | 62 | 64.6
3 | 2019-12-04 | New#2-1 | 62.6 | 65.2
4 | 2019-12-04 | New#3-1 | 61.4 | 63.1
5 | 2019-12-04 | Hot Part (2nd try) | 68.4 | 70
6 | 2019-12-04 | Hot Part (3rd try) | 68.2 | 69.5
7 | 2019-12-04 | New#4-1 | 65.1 | 67.6
8 | 2019-12-04 | New#5-1 | 61.8 | 62.5
9 | 2019-12-04 | New#5-2 | 61.7 | 62.3
10 | 2019-12-04 | New#6-1 | 64 | 64.7
11 | 2019-12-04 | Hot Part (4th try) | 67.4 | 70.8
12 | 2019-12-04 | New 6 Motor + Hot Part Body (1st try) | 62.8 | 64.5
13 | 2019-12-04 | New#6-2 | 66.8 | 67.7
14 | 2019-12-04 | Hot Part Motor + Hot Part Body (after Tightened) | 69.3 | 70
15 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 1) | 68.1 | 69.4
16 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 2) | 67.9 | 69.5
17 | 2019-12-04 | Hot Part Motor + New#6 | 66.8 | 68.2
18 | 2019-12-04 | New#6 Motor + Hot Part Body | 63.7 | 64.2
2019-12-09 | Final report submitted to WPPL QA Team. Pending final dispotion. Eng'g Team activity Close

## Alliance FT3 Tracker (hidden)
Alliance Final Tester 3 Freeze OESR S190059/OESR 193244 Alliance Final Test 3 Frequently Freeze During Final Test
Item | Serial Number | Freeze (Y/N) | Time | Prepared By | Remarks
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50

## IQR TE82943 UPDATE (hidden)
Used Known Good Unit QSM+ (F19QSP275A) and BSM (F19UPB765A)
Week 0f (estimate) | System Number | 2019-06-03 | 2019-06-03 | 2019-06-10 | 2019-06-17 | 2019-06-24 | 2019-07-01 | 2019-07-08 | 2019-07-15 | 2019-07-22 | 2019-07-29 | 2019-08-05 | 2019-08-12 | 2019-08-19 | 2019-08-26 | 2019-09-02 | 2019-09-09 | 2019-09-16 | 2019-09-23
WPPL Flextronics | TE82943-11 | 2019-06-03: TMD 010161 v01; 2019-06-03: IQ
| ACTUAL | 2019-06-03: TEST SOFTWARE UPDATE 07 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-10: 10 Jun 2019 186018080 F19QSP275A 14 Jun 2019 186015001 F19UPB765A; 2019-06-17: PREPARED BY: CYH 19 Jun 2019; 2019-06-24: REVIEWED BY: EM 19 Jun 2019 EJ 20 Jun 2019 APROVED BY: TKC 28 Jun 2019
| TE82943-12 | 2019-06-03: PLAN; 2019-06-10: IQ
| ACTUAL | 2019-06-17: TEST SOFTWARE UPDATE 25 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-24: 26 Jun 2019 186018080 F19QSP275A 26 Jun 2019 186015001 F19UPB765A; 2019-07-08: PREPARED BY: CYH 09 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-13 | 2019-06-03: PLAN; 2019-06-17: IQ
| ACTUAL | 2019-07-01: TEST SOFTWARE UPDATE 09 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-07-08: 10 Jul 2019 186018080 F19QSP275A 11 Jul 2019 186015001 F19UPB765A PREPARED BY: CYH 12 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-14 | 2019-06-03: PLAN; 2019-06-24: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 18 Jul 2019 186018080 F19QSP275A 19 Jul 2019 186015001 F19UPB765A; 2019-07-22: PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-15 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software 18 Jul 2019 186015001 F19UPB765A; 2019-07-22: 22 Jul 2019 186018080 F19QSP275A PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-16 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 07 Aug 2019 186015080 F19QSP275A 08 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-17 | 2019-06-03: PLAN; 2019-07-08: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 08 Aug 2019 186015080 F19QSP275A 07 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-18 | 2019-06-03: PLAN; 2019-07-15: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-08-12: 14 Aug 2019 - testing completed; 2019-08-19: 21 Aug 2019 Resubmit : EQID 0497 Test report with updated calibration report and Evauation Result - Not allowed to used due to somer error on 90g reading during calibration; 2019-08-26: TMD010661 v 01 To be extended until end Sept 19 29 Aug 2019 Submitted TE89243-18 test report to YHC using AND HR 200 SN 12337042 (EQID 2092) 30 Aug 2019 WPPL ISOISO13485 Audit; 2019-09-02: WPPL ISO13485 Audit 30 Aug 2019-02 Sep 2019 TMD010661 v 02 Start at Flex 02 Sep 2019 extended until end Sept 19 AND HR 200 ITS verification completed - within specs. 02 Sep 2019 AND HR 200 Miing Deng re-calibration within specs 05 Sep 2019; 2019-09-09: KC Tan on business trip until 18 Sep 19; 2019-09-16: KC Tan sign-off 18 Sep 19

## Consined Alliance Training Plan (hidden)
Known Good Alliance Transfer Plan To Flex Senai
Mounting and dismantling of PCBA to UUT - 2695, Alarm Message (Related to Allianc 2695 only) - Instrument Specification (Basic related to PCBA) | WW 40 | WW 41 | WW 42 | WW 43 | WW 44 | WW 45 | REMARKS
Plan
Actual
Generic Preventive Maintenance Procedure For Test Equipment (735004259)
Plan
Actual
Transfer of Alliance Known Good Unit to Senai
Plan
Actual
Flex Senai to use Known Good Unit for Troubleshooting
Plan
Actual
Flex Senai to source Spare Part such as cables to enhance toubleshooting
Plan
Actual

## TUV Shunt Energy Level (hidden)
TUV Flow Cell | 55% to 150 %
Energy Setting as per ASP
Shunt ID

## Pump Spares Part (hidden)
Waters ID | Description | Part Number | Spare Part Description | OESR | SOURCE BY: |  | Remarks
[5 empty] | FLEX (Yes/No) | WPPL (Yes/No)
TE82705-2 | 1500 Series Final Test Station | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
|  |  | Waters Pump Control Module |  |  | Yes
|  | SC2072 ASSY181310-02 | National Instrument Controller Card |  |  | Yes
TE82534 | 515 Final Tester |  |  |  |  | Yes
|  | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes | No
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes | No
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
TE82613-2 | Pump Mechanic Run In Module |  | TE82613 PUMP MECHANIC RUN IN MODULE |  |  | Yes
TE82851 | 515 Bezel Assemby Tester |  | TE82851 , 515 BEZEL ASSEMBLY TESTER Both Waters customized PCBA PN is not available |  |  | Yes

## Alliance Pump HDD Imaging (hidden)
Alliance / pump |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.50 | TE 82534 - 2 | 515 Pump Final Test | 2019-02-12 | Completed
10.249.26.51 | TE 82705 - 2 | 1500 pump Final Test | 2019-03-12 | Completed
10.249.26.20 | Final Test - 1 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.21 | Final Test - 2 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.22 | Final Test - 3 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.23 | Final Test - 4 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.24 | Final Test - 5 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.25 | Final Test - 6 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.26 | TE 82526 - 1 | Waters Alliance Test Station ( Pumptray Test - 1 ) | 2019-03-13 | Completed
10.249.26.27 | TE 82526 - 2 | Waters Alliance Test Station ( Pumptray Test - 2 ) | 2019-03-13 | Completed
10.249.26.28 | TE 82526 - 3 | Waters Alliance Test Station ( Pumptray Test - 3 ) | 2019-03-13 | Completed
10.249.26.29 | TE 82519 | Waters Alliance Test Station ( Pump Head Mechanics ) | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.32 | TE 82573 - 1 | E-Alliance Rear Panel Test | 2019-03-12 | Completed
10.249.26.34 | AutoLoader - 2 | Alliance AutoLoader Station 2 | 2019-03-12 | Completed
10.249.26.33 | AutoLoader - 3 | Alliance AutoLoader Station 3 | 2019-03-12 | Completed
10.249.26.35 | TE 82520 - 1 | Alliance Carousel/Carrier Alignment Station -1 | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.36 | TE 82520 - 2 | Alliance Carousel/Carrier Alignment Station -2 | 2019-03-13 | Backup thru Windows OS, not able to use Acronis Imaging Software
10.249.26.37 | TE 82546 | Front Panel Station | 2019-03-13 | Completed
10.249.26.38 | TE 0084018-1 | Heater Cooler Tester | 2019-03-28
10.249.26.90 | QA Workstation | QA Workstation/SQT.ORT | 2019-02-11 | Completed
NA | PMS | Particle Measurement |  | Plan : 30 Apr 19
PLAN : Completion Date

## BSM.QSM HDD Imaging (hidden)
BSM /QSM |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.2 | Waters SISNL1 | Solectron Instrument Serial Number 1 | 2019-03-12 | Completed
10.249.26.3 | Waters SPNL1 | Solectron Part Number Labeling 1 | 2019-03-12 | Completed
10.249.26.4 | WATERSCOFP1 | WATERS Certificate Of Funtional Performance 1 | 2019-03-12 | Completed
10.249.26.9 | TE 82800 - 2 | Acquity Actuator Tester #2 | 2019-02-13 | Completed
10.249.26.10 | TE 82800 - 3 | Acquity Actuator Tester #3 | 2019-03-12 | Completed
10.249.26.11 | TE 82943 - 11 | Acquity SDS Tester #11 | 2019-02-13 | Completed
10.249.26.12 | TE 82943 - 12 | Acquity SDS Tester #12 | 2019-03-12 | Completed
10.249.26.13 | TE 82943 - 13 | Acquity SDS Tester #13 | 2019-02-13 | Completed
10.249.26.14 | TE 82943 - 14 | Acquity SDS Tester #14 | 2019-02-13 | Completed
10.249.26.15 | TE 82943 - 15 | Acquity SDS Tester #15 | 2019-02-13 | Completed
10.249.26.16 | TE 82943 - 16 | Acquity SDS Tester #16 | 2019-02-13 | Completed
10.249.26.17 | TE 82943 - 17 | Acquity SDS Tester #17 | 2019-03-12 | Completed
10.249.26.18 | TE 82943 - 18 | Acquity SDS Tester #18 | 2019-03-13 | Completed
10.249.26.55 | Genealogy SDS | Acquity SDS Genealogy Stn (BSM/QSM) | 2019-03-12 | Completed
PLAN : Completed

## Detector HDD Imaging (hidden)
Detectors / Cleanroom |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.40 | TE 82903 - 6 | TUV Final Test Station | N/A | White Lace \ CD ROM Not able to activate Acronis Imaging Software
10.249.26.42 | TE 82903 - 8 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.43 | TE 82903 - 9 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.44 | TE 82903 - 10 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.45 | TE 82964 - 1 | TUV Pretest System | 2019-03-13 | Completed
10.249.26.75 | TE 82969 - 1 | PDA Linearity Tester | 2019-03-13 | Completed
10.249.26.76 | TE 83057-4 | PDA Download Station | 2019-03-13 | Completed
10.249.26.78 | TE 82882 - 2 | 2998/PDA Final Test System | 2019-03-13 | Completed
10.249.26.79 | TE 82882 - 8 | 2998/PDA Final Test 8 | 2019-03-13 | Completed
10.249.26.80 | TE 82882 - 9 | 2998/PDA Final Test 9 | 2019-03-13 | Completed
10.249.26.81 | TE 82882 - 10 | 2998/PDA Final Test 10 | 2019-03-13 | Completed
10.249.26.82 | TE 82882 - 11 | 2998/PDA Final Test 11 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.84 | TE 82882 -13 | PDA Final Test System 13 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.85 | TE 82882 -14 | PDA Final Test System 14 | 2019-03-14 | Completed
10.249.26.63 | TE 82918 - 2 | 2998 Fine Alignment Sys | 2019-03-14 | Completed
10.249.26.64 | TE 82868 - 2 | 2998 Aperture Alignment System | 2019-03-14 | Completed
10.249.26.65 | TE 82951 - 2 | 2998 Mask Alignment Sys | 2019-03-14 | Completed
10.249.26.66 | TE 82916 - 2 | Flowcell Alignment Fixture | 2019-03-14 | Completed
10.249.26.68 | TE 82973 - 1 | 2998 Flowcell Noise & Drift System | NA | CD-ROM unable to read
10.249.26.72 | TE 82917 - 1 | 2998 Rough Alignment | 2019-03-14 | Completed
10.249.26.73 | TE 82880 - 2 | PDA Rough Alignment Sys | 2019-03-14 | Completed
10.249.26.74 | TE 82881 - 2 | PDA Fine Alignment Stn | 2019-03-14 | Completed
10.249.26.56 | Genealogy SDS | Acquity SDS Genealogy Stn (Detector) | 2019-03-15 | Completed: Backup thru Windows OS
10.249.26.59 | TE83096-2 | Universal PDA Alignment | 2019-03-16 | Completed: Backup thru Windows OS
PLAN : Completion Date