## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2
161048 | Test Engineering | 279001852 | 2016-09-07 | N/A | Desmond Ann | OPEN |  | In preparation for the localization of the 279001852 Assy Drive Injector Housing with Venture GES to have them self support the assembly drive injector housing requirements from 2017 onwards. Below is the list of testers we require for the localization: 1) TE82848 2) TE83027 3) TE83061 Not Required 4) TE83062 Not Required | Sept 22, 2016: The CEA for the 2 testers TE83027 (Matt) and TE82848 (Paul) is approved. The CEA number is WPPL16015. No need to build the two remaining testers TE83061 and TE83062. 27 FEB 2017 This project is currently on hold. 27 APR 2017, as far as I know this project has been cancelled, based on that this OESR is being closed with no action taken. A new OESR can be created should this project be reinstated. No verification \ validation necessary as no action was taken April 28, 2017: Reopen to build testers from WPPL. When is this needed? Oct 25, 2017: Desmond will ask Edwin and provide CEA to build or to close out this request. 30 JAN 2018, No further information from Edwin or Desmond, should this be closed? 27-FEB-2018 No further information from Edwin or Desmond, should this be closed? 17 DEC, 2018: Edwin advise still discussing on the transfer under 2019 AOP Feb 25 March 2019 Edwin can we close this? | Matt Brule | Paul Deschenes | Waiting for Edwin, Desmond to follow up with Edwin; Close as per Edwin
161145 | Manufacturing Engineering | NA | 2016-10-12 | 999119905 | Desmond Ann | OPEN |  | Flex feedback that the fixture 999119905 is having a different press. It is using 99901154 instead of 999041264. The physical fixture do not have item 2, 5, 8,9,11 as stated in the 999119905 fixture drawing. Drawing for 999041250 is different from the physical part. Please see attached for details. | Nov 30, 2016: Correction is being made to the drawing two versions of the same part. Singapore uses a pneumatic press vs Milford Waters is a manual press. Tool qualification is not required documentation only. June 28, 2017: Both Singapore and Milford have MANUAL press, and Singapore has the latest version of the press due to cost saving changes per Steve. If the marked-up print was already provided to Steve to show the difference of the physical press in Singapore versus SAP drawing, Steve should help looking into updating the drawing to reflect the latest version. | Stephen Wright
161186 | Test Engineering | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN |  | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | May 30, 2017: Model shop will cut the chasis for control box in first week of June. Chau continues putting the fixtures together. Chau to check with Eric about effect of the new PDA chip on these two universal testers. June 27, 2017: Model Shop just complete two sheetmetal boxes today to be picked up for the Universal Fine Alignement Fixture. Eric confirms that no effect of new PDA chip on these two universal testers. July 27, 2017: Works in progress. Sep 26, 2017: Plan to ship at end of month so Eric can set it up at Flex. (A few alignment tools might be shipped later on.) Oct 24, 2017: Shipped and being set up at Flex. Building Target Cell, Aperture. Mar 6, 2018: Flex has all parts short of the 2998 target. Oct 26, 2018: Waiting for WPPL to work on capital to build 1 or 2 control boxes. Dec 18, 2018: Edwin already replied on 2Dec18 that it is ok to utilize the budget for the CEA. Waiting for clarification from WPPL to know which budget to use and to build 1 or 2 control boxes. | Eric Turner | Chau Nguyen | Pending for Eric and Eric is out for sergery ; Need
171657 | Test Engineering | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OPEN |  | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. | Eric Turner |  | OFI
171694 | Test Engineering | NA | 2017-04-27 | N/A | Yeap Siew Yee | OPEN |  | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | June 27, 2017: Both of these to be fixed with test software upgrade for SIgmasure, TE83026 is first followed by TE83013. 27-JUL-2017: this will be corrected during up-commoing software revision for Camstar. Sept 26, 2017: Work has started on an update to TE83013, TE83026 is next. 23-JAN-2018 Hope to get time in February to get this moving 30 JAN 2018 Working on TE83013, complete update (win 10, SIgmasure 7, fix OESRs etc) 27-FEB-2018 Actively working on this, Test running working on EDV and bugs as they are found 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. TE83026 will be started soon 20, AUG 2018, Fixing Ecord test bug, one duplicate tester built. 27-FEB-2019 Hope to get time to get this going in 2019 | Matt Brule |  | Matt will follow up this
171714 | Test Engineering | WAT242102 | 2017-05-05 | N/A | Yeap Siew Yee | OPEN |  | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | June 27, 2017: Have not looked into this one yet, should not be an issue, except specification alignment will take time. 27-JUL-2017: Have not looked into yet, will require specification alignment 20, AUG 2018 No action taken yet 29-Nov-2018 No Change 17-DEC-2018 No Change 27-FEB-2019 No Change | Matt Brule |  | Separated language and spec? check with Matt, Edwin check with Quality, Desmond to compile list with OFI related to language; YH has license of OESR
171726 | Manufacturing Engineering | N/A | 2017-05-09 | 000119850 | Desmond Ann | OPEN |  | Tool 119850 is specified in ASP 289001994 but was not physically available in Flex. Milford to fabricate and consign the tool to WPPL. Charge Account is WP21 615150. | Ordered fixture found blue line copy in tool folder, manual drawing. June 28, 2017: Steve will create drawing to get it fabricated. August 30, 2017: WPPL received the tool, and waiting for dwg to be released by Steve into SAP. | Stephen Wright | Roland Poisson
171914 | Manufacturing Engineering | N/A | 2017-07-31 | 999119592 | Desmond Ann | OPEN |  | While raising an ECR to update the PRO WAT045248, Roland found that tool 999119592 does not appear in SAP and could not locate a hard copy in the file. See attached on the drawing sketch for the fixture to be release into SAP. | Jul 31, 2017: The tooling drawing 999119592 will require at least a month to be released into the SAP. Desmond and Mark is okay with TMD010368 while waiting for tool drawing from Steve unless noticed otherwise. Desmond and Steve will follow up to get the tool drawing done to ECR. Aug 02, 2017: Desmond has volunteered to create a model file in solid works because this will speed up the completion of this request. SW Aug 2. 17Dec18 - Pending Steve Wright to update the drawing. | Stephen Wright | Mike Aubin
171930 | Manufacturing Engineering | N/A | 2017-08-03 | N/A | Desmond Ann | OPEN |  | We have fabricated a jig to help in the assembly of the PCB. This jig is proposed to be used for all the detectors models listed in the matrix. The jig will have 4 bases to be used based on the model, see attached on the base used per the product matrix. This jig will be able to press the CPU board evenly into the standoff to the Personality card and also prevent flexing of the PCBAs during assembly. We will need Milford help to assign a tooling number and upload the drawing into SAP. | August 30, 2017: When Steve releases dwg into SAP, Desmond will complete ASP update. 17Dec18 - Yee Hwa will be performing the evaluation will be performed in Jan 2019. | Stephen Wright
171993 | Test Engineering | 735002064 | 2017-08-25 | TE83013 | Yeap Siew Yee | OPEN |  | TE83013-1 is observed to experience a test sequence error at "Ecord Interface Test" by skipping to instruct to connect "Ecord" to the UUT while testing for 186015042. They have recently changed the tester PCBA P/N:210000367 (ASSY PCB CPSM CONTROLPCB) Rev H because existing P/N:210000367 Rev G was faulty on 14-Aug-2017. Refer to videos for failure occurrence. | 25-SEP-2017: In progress. 30-JAN-2018 Actively working on this 27-FEB-2018 working on bug fixes and EDV 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. 25, June 2018 Emailed Desmond this AM with proposed upgrade plan 20, AUG 2018 working on Ecord bug. 24-Nov-2018 Working on bug as time allows expect resolution soon, working to prove theory about Ecord formatting 17 DEC 2018 Same status (Pending Matt Brule.) 23-FEB-2019, expect to get freed up to complete this soon | Matt Brule |  | wait for matt
172141 | Manufacturing Engineering | N/A | 2017-10-23 | 999029310 | Desmond Ann | OPEN |  | Flex feedback that the tooling 999029310 Fixture wheel to motor 2998 classic. Following sub-drawing are missing and not released:- 1. 999041435 Column, support filter wheel to motor. 2. 999041438 Bridge filter wheel to motor. | These parts exist in another fixture 999026549. They will be released with fixture 999026549. See ECO# 17Dec18 - Pending Steve Wright to provide ECO number to update drawing. | Stephen Wright
182626 | Test Engineering | N/A | 2018-06-27 | TE82943 | Desmond Ann | OPEN |  | Missing serial number at the bottom of the test report (Ser.No) of Acquity BSM 18K PLUS and Acquity Arc QSM test in some Test record and it is encountered across all test stations TE82943 #1,#19,#20,#21,#22 and #24. Please kindly refer to the examples "BSM Test Record.JPG" and "QSM Test Record.JPG". Do help to resolve the issue. | Oct 2018: Made a Test Program code modification on the TE Lab PC and is under evaluation. 17Dec1: Check with Venture to see if this is resolved. 29Jan2019: It has not been resolved and still require support to solve the issue. 29-Jan-2019: A fix for this intermittent condition has been added to the latest test programs. It was going to be released along with Firebrick firmware in January. However, Firebrick firmware is now being rolled up into Gold firmware. So now it will be released in April with the Gold firmware release. Feb: same status. Apr23: same status. expecting release soon. | Edward Jackson |  | Wait for gold firmware release
182922 | Manufacturing Engineering | N/A | 2018-11-16 | N/A | Cheah Su Fern | OPEN |  | There are few errors found on the Part Drawing for Belt Tightening Fixture 999038298. Details will be attached in the email. | It looks like maybe a repair was made in Singapore by the CM that needs addressing. Area of concern is clearance only Desmond requested a print change to open up tolerance for concerns of an audit. | Stephen Wright
182944 | Manufacturing Engineering | N/A | 2018-11-27 | N/A | Desmond Ann | OPEN |  | We are being informed today that the Sharp Autobagger is having printing issue located in Richport. The barcode is unable to be scanned due to the poor printing and also the P/N print out is not clear. This is impacting the shipment for the instruments in the Venture as the packaging printing is being rejected by the Quality team. This parts are pack as accessories for the Column Heater. We do not have support for the Sharp Autobagger locally as there is no service rep here and this machine is more than 8 years old. 1) Please advise if a TMD can be given to UAI this condition of printing. 2) Please advise if we can use other brands of Autobagger which have support locally to replace the Sharp Autobagger. Is there any other brands that Milford is recommending? Details and photos are as follows. Defect: Poor Printing. Part Number: W1405003287-R/CRC Part Description: Fitting 10-32 One Piece Modified. Supplier/Vendor: Richport Technology Pte Ltd Qty Affected: 147/147 ~ 100% Instruments Affected: 186015045 (ACQUITY 30cm CHA). |  | SuFern Cheah
183008 | Test Engineering | 735000888/7 | 2018-12-19 | TE82817 | LUO HUAZHEN | OPEN |  | During business review recently, we found that TE82817 is at high risk as it’s only tester to support both PCM and TCM testing (Sharing the same PC) and tester/PC is old. If this tester down, it can cause production line of TCM/PCM to be stopped. This test station has IBM M50 PC and test box which are very old. Upon checking, IBM M50 pc was manufactured in 2003 and its OS is window XP. Test box for PCM is also an old test box. Please refer to attached photos for reference. Therefore we’re seeking your advice if we can request to upgrade tester/PC and test box for TE82817, or another TE82817. | Currently supporting WPPL efforts for redesign. | Mark Jussaume | Bill Bates | Close. To upgrade PC with software transfer. Low priority
193069 | Manufacturing Engineering | N/A | 2019-01-22 | N/A | Chan Yee Hwa | OPEN |  | ECO 9221720 & 9221721 changed SL Power Supply 279000852 to Delta Power Supply 279003675 for various Detectors. SigmaSure power supply part number need to update accordingly. We will need 7 model to be change today and 4 model to be change tomorrow due to a depletion. Request to update for below part number TODAY: 186002487 186017002 186015034 186017012 186017013 186002998 186015053 186299800 186015031 Request to update for below part number TOMORROW: 186248900 186017003 186015032 186015033 Thank you very much. |  | Matt Franklin
193112 | Manufacturing Engineering | 279002224 | 2019-02-11 | N/A | Chan Yee Hwa | OPEN |  | Sigmasure changes in Detector 186015034 Acquity UPC2 PDA. Noticed additional Part Number 279002225 require to be scan since 02/10/2019 in Sigmasure. Previously were only 5 PN to be scan and now 6 in total. Request to remove additional PN 279002225 from 186015034. Please refer to attachment in email. |  | Matt Franklin | Roland Poisson
193124 | Test Engineering | N/A | 2019-02-14 | TE82526 | Chan Yee Hwa | OPEN |  | Request additional Pump Tray Tester TE82526 to support Alliance Production for high demand. The tester is currently running at 92% capacity; Besides, we also see high chance of tester break down according to past years data. Please provide the cost for WPPL to raise CEA. Thanks. |  | Mark Jussaume | Chau Nguyen | CEA is available. Mark to work on this. Mid priority due to volume is going down. Moderization will be taking longer.
193132 | Test Engineering | N/A | 2019-02-15 | N/A | Chan Yee Hwa | OPEN |  | Request help from Eric Turner to program 2 Shunts with Shunt ID for Flex production use. Thanks. |  | Eric Turner |  | Wait for Eric back
193148 | Test Engineering | 735000615 | 2019-02-25 | TE82760 | Luo HuaZhen | OPEN |  | During meeting with Venture for First Pass Yield (FPY), the FLR detector yield was 3 consecutive months lower than target yield (97% for UPLC detector and 90% for 2475 detector). The venture feedback they're seeing Signal to Noise Ratio failure increased, it's the top failure after Raman/Rayleigh failure. The units are failed the 1st test with readings approximately zero (~0.00) and the 2nd test (retest) passed with over thousands (>1000). 55.56% of Signal/Noise Ratio failure are failed with readings nearly to "zero". And we tried to use Diagnostic on test program to troubleshoot this failure, but it's always blanked out. Could you please kindly help to check this issue and advise the possible root cause? Thanks |  | Matt Franklin | Eric Turner | Wait for Eirc back
193161 | Manufacturing Engineering | 289012412 | 2019-03-01 | N/A | Chan Yee Hwa | OPEN |  | Request Steve help to release a new jig to SAP. This jig is a holding fixture for Alliance new seal pack 289012412, to be replace current jig 999044874. Drawing will be attached in Email. Thanks. | Looked over supplied prints and requested to have missing information added.6-Mar.Yee Hwa is planning to have changes completed by 12-Mar. New prints supplied, documentation is much better and will be release on ECO # | Stephen Wright
193162 | Test Engineering | N/A | 2019-03-01 | TE82988 | Luo HuaZhen | OPEN |  | WPPL received this chassis from CEI Batam after our spare chassis was used for this repair. The failure reported by CEI as shown as below. Could you please help to repair this chassis by replacing Power Supply in Chassis? Should I send this faulty Chassis to Milford for repair and test? Please kindly advise. Thank you 1. One power supply at GEOTEST chassis GX7310 having intermittently on-off problem. The Geotest chassis is in GTS tester TE82988-9. 2. Power Supply part number is PRM560. Requesting a replacement unit. 3. This GTS tester is mainly for testing PCBA for Waters UK (Micromass). | Rick Davis on March 01, 2019: "We do not repair the instruments in a GTS. We send them back to the manufacturer and have them repaired. I suggest you contact the appropriate Micromass person and have them arrange to have the chassis repaired. Chau, you need to assign this to someone from Micromass." | Luo Hua Zhen |  | No Update
193182 | Manufacturing Engineering | WAT207000 | 2019-03-07 | N/A | Chan Yee Hwa | OPEN |  | Flex has reviewed current assembly process and found multiple HPLC Pump PRO files need to be update to capture actual PN, new tooling number, and extra tools needed for assembly. Request help to update below PRO files, redlined files will be attached in email; WAT207000 WAT207033 WAT207035 WAT207066 Thanks. | Note the annual volume has been: 88 in 2018, 130 in 2017, 128 in 2016, 65 in 2015, 84 in 2014. | Stephen Wright
193201 | Manufacturing Engineering | 289001458 | 2019-03-15 | N/A | Liu Jingxuan | OPEN |  | Tooling fixture number is stated wrongly in drawing 289001458 under Note 2. The correct tool number should be 999028067 instead of 999026067. | Request for resources submitted 23 April. | Stephen Wright
193204 | Manufacturing Engineering | N/A | 2019-03-15 | N/A | Chan Yee Hwa | OPEN |  | Request part number update in Sigma Sure per ECO9227819 requirement. For QSM+ 186015080 Change PN from 279001864 to 289003741. Thanks. |  | Matt Franklin | Roland Poisson
193210 | Test Engineering | N/A | 2019-03-15 | TE83000 | Paul Deschenes | OPEN |  | Flextronics has an ICT fixture they are not presently using. TE83000-2 used to test the 210000520 PCBA . I would like to have that fixture shipped to Milford for an update. |  | Luo Hua Zhen | Desmond Ann | 5544
193232 | Test Engineering | N/A | 2019-04-01 | TE83026 | Luo HuaZhen | OPEN |  | Venture Singapore reported their CMA tester TE83026-1 is down due to Test PC window XP could not be loaded on 29 March 2019. Hang at window XP loading step. Venture TE tried to recover the PC by loading PC image files but unsuccessful. So they replaced the test PC with another spare PC (Windows XP). Now need your help on the system configuration. The new PC IP address is the same as original TE83026-1 IP address 10.249.40.27 An UUT will be connected to the tester. Please kindly help on this configuration and advise the validation requirement once the tester is fixed. The production is down now so they need your assist urgently. Thank you |  | Matt Brule |  | Closed
193244 | Test Engineering | WAT271129 | 2019-04-04 | 186000139 | Luo HuaZhen | OPEN |  | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. |  | Matt Brule |  | update to Matt
193266 | Test Engineering | N/A | 2019-04-12 | N/A | Chan Yee Hwa | OPEN |  | Would like to request for new shunts to support detector production test capacity at Flex Sg. We'll need below shunt with programmed chips. Thanks. i. 289002450 -- TUV Shunt ; 3 Ea. ii. 289002227 -- PDA Small Shunt; 1 Ea. iii. 289002226 -- PDA Medium Shunt; 3 Ea. |  | Eric Turner | Chau Nguyen | Edwin is looking into self built for this shunt
193267 | Test Engineering | 735001613 | 2019-04-15 | TE82882 | Luo HuaZhen | OPEN |  | New ECO is required Flex Singapore to supply 2 new spare Part Number 700012245 and 700012243. Please kindly help to review/update the TSP 735001613 to support new part number 700012245 and 700012243 test. Below is request sent by Flex, FYI 1. Problem Statement. - TE82882 PDA Final Test and TSP 735001613 Rev W does not have Spare OB PN 700012245 - Optics Bench, eLambda (On-Semi) and 700012243 - Assy, Optics Bench, 2998 (On-Semi) 2. Acceptance criteria that is not met. (e.g. relevant specification number and revision) - TE82882 PDA Final Test and TSP 735001613 Rev W must have to support Production run of the 2x Spare OB. 3. Relevant records (e.g. Incomin QA, Test Records, Photos, trackise NCRs numbers and related issue) - Relevant PDA Family Test Report for Spare Optic Bench have dedicated PN on the test station and TSP. 4. Info on whether this issue impacts revenue. - This impact the spare bench testing. |  | Matt Brule | Eric Turner | A gap? New ECR should notify to TE. So that ECO can include the new PN; Check with Ying ying how to convey the right to relevant party. Check with Lisa
193269 | Test Engineering | 735000615 | 2019-04-15 | TE82760 | Luo HuaZhen | OPEN |  | Tester show error message “serial number scanned Does not match store value.” . It seems failed to communication between UUT and tester. Compared with another FLR tester TE82760-4, no 1394 net adapter shows under network connection. and tester TE82760-4 is working fine. Please kindly help to check if any software or setting needs to be updated for tester TE82760-5 to work as per normal. Thank you |  | Matt Brule | Eric Turner | Closed