## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2 | WPPL Updates
161145 | Manufacturing Engineering | NA | 2016-10-12 | 999119905 | Desmond Ann | OPEN |  | Flex feedback that the fixture 999119905 is having a different press. It is using 99901154 instead of 999041264. The physical fixture do not have item 2, 5, 8,9,11 as stated in the 999119905 fixture drawing. Drawing for 999041250 is different from the physical part. Please see attached for details. | Nov 30, 2016: Correction is being made to the drawing two versions of the same part. Singapore uses a pneumatic press vs Milford Waters is a manual press. Tool qualification is not required documentation only. June 28, 2017: Both Singapore and Milford have MANUAL press, and Singapore has the latest version of the press due to cost saving changes per Steve. If the marked-up print was already provided to Steve to show the difference of the physical press in Singapore versus SAP drawing, Steve should help looking into updating the drawing to reflect the latest version. June 03, 2019: Pending Steve to update. | Stephen Wright |  | Chau to follow up with Jaffar and or Edwin
161186 | Test Engineering | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN |  | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | May 30, 2017: Model shop will cut the chasis for control box in first week of June. Chau continues putting the fixtures together. Chau to check with Eric about effect of the new PDA chip on these two universal testers. June 27, 2017: Model Shop just complete two sheetmetal boxes today to be picked up for the Universal Fine Alignement Fixture. Eric confirms that no effect of new PDA chip on these two universal testers. July 27, 2017: Works in progress. Sep 26, 2017: Plan to ship at end of month so Eric can set it up at Flex. (A few alignment tools might be shipped later on.) Oct 24, 2017: Shipped and being set up at Flex. Building Target Cell, Aperture. Mar 6, 2018: Flex has all parts short of the 2998 target. Oct 26, 2018: Waiting for WPPL to work on capital to build 1 or 2 control boxes. Dec 18, 2018: Edwin already replied on 2Dec18 that it is ok to utilize the budget for the CEA. Waiting for clarification from WPPL to know which budget to use and to build 1 or 2 control boxes. June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. | Eric Turner | Chau Nguyen | Milford is working on it
171657 | Test Engineering | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OPEN |  | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. Feb 26, 2020: Low Priority Per Eric | Eric Turner |  | WPPL is ok to put it as OFI but it should have a completion date.
171694 | Test Engineering | NA | 2017-04-27 | N/A | Yeap Siew Yee | OPEN |  | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | June 27, 2017: Both of these to be fixed with test software upgrade for SIgmasure, TE83026 is first followed by TE83013. 27-JUL-2017: this will be corrected during up-commoing software revision for Camstar. Sept 26, 2017: Work has started on an update to TE83013, TE83026 is next. 23-JAN-2018 Hope to get time in February to get this moving 30 JAN 2018 Working on TE83013, complete update (win 10, SIgmasure 7, fix OESRs etc) 27-FEB-2018 Actively working on this, Test running working on EDV and bugs as they are found 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. TE83026 will be started soon 20, AUG 2018, Fixing Ecord test bug, one duplicate tester built. 27-FEB-2019 Hope to get time to get this going in 2019 May 2019: Still at same status, second on project priority list 26 Feb 2020: Matt is working on this and should be completed soon. | Matt Brule |  | Milford to advise the estimated complete date.
171714 | Test Engineering | WAT242102 | 2017-05-05 | N/A | Yeap Siew Yee | OPEN |  | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | June 27, 2017: Have not looked into this one yet, should not be an issue, except specification alignment will take time. 27-JUL-2017: Have not looked into yet, will require specification alignment 20, AUG 2018 No action taken yet 29-Nov-2018 No Change 17-DEC-2018 No Change 27-FEB-2019 No Change July 02, 2019: WPPL : Matt reply in May 2019, Procedure update, low priority, no change, requires specification alignment Feb 26, 2020: Specification alignment will take time and Low Priority Per Matt | Matt Brule |  | WPPL is ok to put it as OFI but it should have a completion date.
171993 | Test Engineering | 735002064 | 2017-08-25 | TE83013 | Yeap Siew Yee | OPEN |  | TE83013-1 is observed to experience a test sequence error at "Ecord Interface Test" by skipping to instruct to connect "Ecord" to the UUT while testing for 186015042. They have recently changed the tester PCBA P/N:210000367 (ASSY PCB CPSM CONTROLPCB) Rev H because existing P/N:210000367 Rev G was faulty on 14-Aug-2017. Refer to videos for failure occurrence. | 25-SEP-2017: In progress. 30-JAN-2018 Actively working on this 27-FEB-2018 working on bug fixes and EDV 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. 25, June 2018 Emailed Desmond this AM with proposed upgrade plan 20, AUG 2018 working on Ecord bug. 24-Nov-2018 Working on bug as time allows expect resolution soon, working to prove theory about Ecord formatting 17 DEC 2018 Same status (Pending Matt Brule.) 23-FEB-2019, expect to get freed up to complete this soon July 02, 2019: Matt reply in May 2019, First on his project priority list 26 Feb 2020: Matt is working on this and should be completed soon. | Matt Brule |  | Matt -- completed soon
172012 | Manufacturing Engineering | N/A | 2017-09-01 | N/A | Chau Nguyen | OPEN |  | Work with Chau on Target Cells and other required tools and associated drawings to complete and transfer the two fixtures to WPPL by Mid-October 2017 . | Need a list of the targets SW......... not enough info provided ......Oct 2017. SW Oct 26, 2018: Chau will sort through components on hand to build spares. June 03, 2019: Universal Coarse Fine Alignment station pending Milford to fabricate. June 17, 2019: Universal Coarse Fine Alignment station already built and shipped. This request is to document the process to make Target Cells and other required tools to build spares when required. If there is enough materials on hands and available to build any target, it would be built if required. Chau Nguyen. June 17, 2019: WPPL: Milford to advise more details on what is required for this OESR. These will be required when 2 control boxes completed. Request for completion date of 2 control boxes from Eric to provide to Finance in WPPL. 26 Feb 2020: Gather info to move forward | Stephen Wright | Chau Nguyen | Chau will work with Stephen on this request
182922 | Manufacturing Engineering | N/A | 2018-11-16 | N/A | Cheah Su Fern | OPEN |  | There are few errors found on the Part Drawing for Belt Tightening Fixture 999038298. Details will be attached in the email. | It looks like maybe a repair was made in Singapore by the CM that needs addressing. Area of concern is clearance only. Desmond requested a print change to open up tolerance for concerns of an audit. June 03, 2019: Pending Steve to update the drawings. Drawing in Concept status. July 02, 2019: WPPL : Pending Steve to update | Stephen Wright |  | Chau to follow up with Jaffar and or Edwin
193124 | Test Engineering | N/A | 2019-02-14 | TE82526 | Chan Yee Hwa | OPEN |  | Request additional Pump Tray Tester TE82526 to support Alliance Production for high demand. The tester is currently running at 92% capacity; Besides, we also see high chance of tester break down according to past years data. Please provide the cost for WPPL to raise CEA. Thanks. | June 03, 2019: Mark has begun looking at this, and is reviewing best path forward regarding component and PC obsolescence. CEA# WPPL19003 Updated CAD file with new bracket and bracket was added manual parts list S.W. Sept. 5 Fluidic diagram needs to be updated per change. 24 JAN 2020, MB: Additional system, and parts upgrade kits, shipped Jan 7. | Mark Jussaume | Mike Aubin | Mark will ship tester WPPL update: New Pump Tray tester is set up and completed the validation. The upgrade of the existing testers is in progress.
193244 | Test Engineering | WAT271129 | 2019-04-04 | 186000139 | Luo HuaZhen | OPEN |  | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. | July 02, 2019: WPPL : Matt reply that he is working on it. Nov 15, 2019: The frequency of PC freeze sent to Matt. 26 Feb 2020: Matt is working on it and coordinating with HuaZhen | Matt Brule |  | HZ resend the freq of PC freeze to Matt WPPL update: Resend the data log to Matt Brule on 20 Dec 2019. Milford to advise status
193266 | Test Engineering | N/A | 2019-04-12 | N/A | Chan Yee Hwa | OPEN |  | Would like to request for new shunts to support detector production test capacity at Flex Sg. We'll need below shunt with programmed chips. Thanks. i. 289002450 -- TUV Shunt ; 3 Ea. ii. 289002227 -- PDA Small Shunt; 1 Ea. iii. 289002226 -- PDA Medium Shunt; 3 Ea. | Request for CEA from WPPL to place actual order in SAP to determine which components available and which to be sourced to complete building these shunts. June 03, 2019: YH to raise PO for Shunts. July 02, 2019: WPPL : Pending for WPPL to raise CEA (and or PO?) to enter order. 26 Feb 2020: Shunts built and waiting for shipping information from WPPL to ship | Eric Turner | Chau Nguyen | Chau is working on this (Emailed Purchasing and Finance on Dec 23, 2019 to start building shunts required)
193382 | Test Engineering | N/A | 2019-06-06 | TE82743 | Luo HuaZhen | OPEN |  | Venture is running QDA production too. Some of testers are using dynamic IP address and those testers are sharing the same network with Milford testers. Therefore it can cause IP address conflict, while tester from Milford is power off and the tester from QDA production is assigned a same IP address by server. Can Milford assign a range of IP address for QDA testers? I will work with local team/Waters UK to check if they can assign a fixed IP address within this range for each of testers for QDA production. So we can avoid the same issue happening again. | July 02, 2019: WPPL: Milford to advise the range of IP address to use for UK equipments. This is urgent to get IP conflicts resolved for CM Production. 31-Oct-2109: Per WPPL, info was provided to have 50 static IP to be used for UK testers. Email sent on 30 Sep 2019. Pending for Milford/UK to advise. 26 Feb 2020: Chau has emailed to contact Raymond but no responses yet | Matt Brule |  | WPPL has sent the request info to Raymond and waiting for feedback from Raymond. Chau will contact with Raymond on this case
193411 | Test Engineering | N/A | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN |  | New FLR Detector tester (Win7) don’t get the Factory Order and TMD # on new test report template and the new template format is slightly different from xp template although the same parameters were tested. To standardize the report, we need Milford help to update the report templates to the same format as the one on the old testers(WinXP). Please refer to page 2 for details of differences on 2 report templates. | July 02, 2019: WPPL : Pending Eric to update. | Eric Turner |  | Eric is working on this standardization and WPPL to extend TMD based on completeion timeline from Eric. WPPL update: TMD is extended. Team is reviewing the udpated test record.
193413 | Test Engineering | 735000615 | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN |  | TSP735000615 Rev P, page 8 of 9, the TSP is missing a step during keypad test after passing the functional test . For 2475 mode, we need to set the date for the lamp installed. Press “DIAG”, “7” ,” 1” as per red line TSP attached at page 2 of this OESR Please review this step and update the TSP if necessary. Thanks | July 02, 2019: WPPL : Pending for Milford to update TSP. Feb 26, 2020: Low Priority, and also reassigned to Matt per Eric. If it's OFI, provide estimated completion timeline. | Matt Brule | Eric Turner | Chau follow up with Eric to determine if it's OFI
193459 | Test Engineering | 735002010 | 2019-07-29 | TE82943 | Luo HuaZhen | OPEN |  | Since last year, we were facing multiple(4 times) BSM/QSM tester breakdown due the weighing balance faulty in Flextronics. Each time it will take around 3 months to get the weighing balance repaired in Japan due to lack of repair parts in ITS Singapore (AnD local representative). The repair turnaround time is very long and costly (around USD1500 for each repair). So we checked with ITS again recently regarding supportability of this HR200 model. They sent us a product discontinuation letter and its successor model GR-200 info as attached Can we replace the current old HR-200 weighing balance with its successor mode GR-200 since these HR200 are frequently breaking down? If can, how can we proceed with this replacement? Or do you have any other better solution for HR200 support? Please refer to GR200 user manual and discontinuation letter in Pg2 & 3. Thank you | 21-Aug-2019: A discussion took place between Jaaffar Ismail, Matt Brule, and Ed Jackson. We decided to first pursue a direction of maintaining the HR-200 balances. We will ask Paul Deschenes next week if he can identify any spare balances throughout the building that we may be able to use to support the TE82943 testers. The long term solution, at least a year away, is to replace the TE82943 testers with a new design that will eliminate the balances by replacing them with flow meters. Ed J.. 23-Sep-2019 BB - Shipping a spare HR-200 balance to HuaZhen's attention. This will need Calibration once received in Singapore/when installed on tester setup. 31-Oct-2019: Per WPPL, the spare balance was received and tester is working now. Ed suggests to leave this OESR open until he completes the viablity study of new model because the old model is no longer available to be purchased from supplier. 26 Feb 2020: New scale being evaluated and Chau is working on a custom Top Plate to provide to Ed for further testing. | Edward Jackson | Paul Deschenes | Chau to follow up with Ed
193460 | Test Engineering | N/A | 2019-07-29 | TE82475 | Luo HuaZhen | OPEN |  | Based on tester utilization calculation, this pressure decay tester in Flex has been used up 91.56% of total tester capacity. So we plan to buy another TE82475 tester by raising CEA. Please kindly advise the total price and estimated complete date to build this tester. Thank you Tester utilization data, FYI Tester Utilization based on Q1 & Q2’20 numbers: Prime is 79.3% Prime + Spares is 91.56% Q3 & Q4’19 numbers: Prime is 78.2% Prime + Spares is 91.5% | [20Dec2019] Redeisgn under way 26 Feb 2020: Redeisgn under way | Mark Jussaume |  | Milford to provide comletion timeline. Potential production impact due to no backup tester. WPPL update: WPPL is working on CEA
193551 | Test Engineering | 735002011 | 2019-09-02 | TE83032 | Luo HuaZhen | OPEN |  | The existing optics bench alignment stations are using a Golden flow cell to align all the optics benches. The flow cell used to align 5029 FLR optics benches had been using since year 2013 product transferred period and it had never been replaced. Venture observed the white colour residue on the glass window of the flow cell and it can not be removed although it was purging/priming through with MeOH for long hours. Refer to attached slide. Venture should have one back up golden flow cell in any case of flow cell degrading over a few years of usage and also for not interrupting production line while it waiting Milford to send over new Golden flow cell. | 26 Feb 2020: Reassigned to Matt Frankline per Eric and updates seen below: We are working on the improvement and we are going to order some optics benched from Singapore to do some testing but currently they are dealing with a quality issue that we are waiting to get resolved before we can order the benches. Singapore should be up to date on this project. | Eric Turner | Matt Franklin | Chau to follow up on this case
193552 | Test Engineering | 73500615 | 2019-09-02 | TE82760 | Luo HuaZhen | OPEN |  | On the month of Jul, Venture production found some Ram-Ray Valley Energy failures during FLR detectors final test for 5029 models. The UUTs were passed when they swapped with new flow cells. The failed flow cells were used in another UUTs and passed the test. Refer to attached slides for failure details. Raised OESR to follow up and get some advices from WPPL/Milford Engineering to look into the possible root causes of this failure. | 26 Feb 2020: Reassigned to Matt Frankline per Eric and updates seen below: This should no longer be an issue. I remember checking the test data and everything seemed to be back to normal. If they are no longer having this issue this OESR should be closed. | Eric Turner | Matt Franklin | Chau to follow up on this case
193563 | Test Engineering | 735001613 | 2019-09-06 | TE82969 | Luo HuaZhen | OPEN |  | Flex feedback the latest PDA Final test TSP 735001613 Rev Y has part numbers which is listed in the other TSP 735002531. This will create confusion. Please remove the duplicate numbers per attached redline TSP and also some other updates shown in the table. | 26 Feb 2020: Will get it done soon | Eric Turner | Matt Brule | Milford will get it done soon
193757 | Test Engineering | N/A | 2019-11-08 | TE83021 | Luo HuaZhen | OPEN |  | Venture reported this fixture shows error message ' Serial CMD XCVR. The operation time out'. Suspect the CPU board of this fixture is faulty. Please kindly send us a right firmware for this CPU board for this repair or send us a CPU with firmware. Right now 2 pcs of PCB are pending for this test. | Firmware sent 13-NOV-2019, Matt B | Matt Brule |  | Repair in progress, the fixture has been shipped to Matt. WPPL will close TMD once fixture is returned from Milford.
193758 | Manufacturing Engineering | 186015032 | 2019-11-08 | N/A | Michael Cotter | CLOSED |  | On a returned ACQUITY PDA detector, the slit appeared to be damaged, physically, gouged, scraped, as if someone (production, test, customer, field service, etc.) had to struggle with this component. When the repair/refurb department removed the slit, they found it to be quite tight (an interference fit) in the optics bench hole, when this component typically drops in and comes out without much trouble. It raises the question if this was simply a one-off, an isolated incident (perhaps the slit was misshapen, due to rough handling), or is this something that we see in production? Generally, how does this component typically fit in production? Do we ever have difficulty mounting this component to the optics bench? And if so, how often, and have we inspected the parts (slit optics bench)? | Desmond has updated to mike on 22 Nov 2019. Going to close it after chau discussed with Mike. 26 Feb 2020: Closed | Desmond Ann | Luo Hua Zhen | Desmond has updated to mike on 22 Nov 2019. Going to close it after Chau discussed with Mike
193786 | Test Engineering | 735002510 | 2019-11-21 | TE82865 | Luo HuaZhen | OPEN |  | CEI request for test program update to allow enough time for plate actuation adjustment. There for reducing the false failure due to test program run out time while adjusting the flow control knob. Please kindly advise a way to extend the time for this adjusting or help to remotely update the test software. Below is the detail of reported issue, CEI is facing numbers of failure on the adjust plate latch actuation speed during functional setup for CPSO model. Production adjusted according to the adjust Latch Actuation Speed step shown on screen by preset plate latch flow knob controls by loosening the stops nuts, then rotate the flow control knob clockwise all the way, then turn counterclockwise 1.5 turns on the upper flow control and 0.75 turns on the lower flow control. Production feedback when click 'OK' they don’t have enough time to adjust the knob before test program pop out to "fail". This will result in final test fail and affecting FPY. | 19-BEC-2019 Matt B. CM looking at adjusting "preset" process to avoid failure. If successful this will be moved to OFI and test program will be updated at next change. | Richard Davis | Matt Brule | WPPL is waiting for prodution order to evaluation if preset can be adjusted to advoid error
193814 | Test Engineering | 735001613 | 2019-12-17 | TE82882 | Luo Huazhen | OPEN |  | The final test record for Detector(TE82882-All and TE82969-1) and SDS ( TE82943-All) has incomplete of 'start time' and 'end time' print out. The second of time is disappeared from test record. Please kindly update the final test record template to print out the complete time, including the seconds on test record. Please refer to attachment on page 2 for the details. | This is a known issue with the test record formatting file and will be updated as time is available. 26 Feb 2020: Changed to OFI with no clear date on completion |  |  | Please provide estimated date to complete
203852 | Test Engineering | N/A | 2020-01-10 | TE83044 | Luo HuaZhen | OPEN |  | Flex Malaysia request to have a spare CPU with firmware for the following tester fixtures. Most of these CPU are obsoleted. Please kindly advise a spare solution. Can you please send us a spare if you have same part number? Or please advise if can use 520 CPU2K board with correct firmware version from you? Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 pls refer to page 2 attachment 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 | 26 Feb 2020: FYI that Tony is no longer with Waters | Paul Deschenes | Tony Fontes | Milford to advise the status
203933 | Test Engineering | N/A | 2020-02-07 | TE82984 | Luo HuaZhen | OPEN |  | Venture Singapore feedback the Column Manager Testers have different cut off point for temperature. Some units required to have retested, then can Pass. It's affecting the FPY. If UUT’s column temperature is lower than 20degC, test software(Ver1.1) cut off the heater/cooler test and displays as fail. This failure mode is observed on CM-Aux (5049. 5039) and Classic CM (5007, 5008 and 5009) but do not have similar issue on CM (5043 & 5038) Test software (Ver1.3) Please kindly review CM-Aux and classic CM test software (Ver 1.1) to enhance the test methodology for heater/cooler test if it's required. | 26 Feb 2020: Matt is reviewing and will assign |  |  | Milford to advise the status
203936 | Test Engineering | N/A | 2020-02-10 | TE83044 | Luo HuaZhen | OPEN |  | Flex Malaysia GTS has 2 fixtures down, TE83044 and TE83002. Both fixtures were found its CPU2k board faulty. Flex needs a replacement of this CPU board. Please kindly send a replacement CPU2k board with firmware loaded for this 2 fixtures repair. Or please send the firmware and schematic of fixtures to allow team to work on troubleshooting. The original CPU card is 210000273 Rev G. If fixtures cannot be fixed before 19 Feb 2020, the shipment of following part numbers will be affected. 186015031; 186015032; 186015033 and 186015080 | 26 Feb 2020: Matt is reviewing and will assign
203955 | Test Engineering | N/A | 2020-02-17 | OESR | BILL IRONS | OPEN |  | TEST OF LIMITED ACCESS ACCOUNT | 26 Feb 2020: Bill is testing OESR system for updates. No actions taken. | Bill Irons |  | Milford to provide more info
203966 | Test Engineering | WAT271129 | 2020-02-19 | TE82573 | Mike Cotter | OPEN |  | Please provide the injector adjustment test values (top and width) for 50 recent instruments. Also, please provide the injector adjustment test values (top and width) for one particular instrument (D19SM4575A). Thanks, Mike |  | Desmond Ann | Luo Hua Zhen | WPPL udpate: Data has been sent to Mike. WPPL to clsoe this OESR
203983 | Test Engineering | 735000615 | 2020-02-26 | TE82760 | Chan Yee Hwa | OPEN |  | OFI: Current test report generated from all TE82760 do not show the firmware version of 186015029, 186017006. Request to include the firmware version in test report for next software update | 26 Feb 2020: Matt is reviewing and will assign |  |  | Milford to advise the status