## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2
161048 | Test Engineering | 279001852 | 2016-09-07 | N/A | Desmond Ann | OPEN |  | In preparation for the localization of the 279001852 Assy Drive Injector Housing with Venture GES to have them self support the assembly drive injector housing requirements from 2017 onwards. Below is the list of testers we require for the localization: 1) TE82848 2) TE83027 3) TE83061 Not Required 4) TE83062 Not Required | Sept 22, 2016: The CEA for the 2 testers TE83027 (Matt) and TE82848 (Paul) is approved. The CEA number is WPPL16015. No need to build the two remaining testers TE83061 and TE83062. 27 FEB 2017 This project is currently on hold. 27 APR 2017, as far as I know this project has been cancelled, based on that this OESR is being closed with no action taken. A new OESR can be created should this project be reinstated. No verification \ validation necessary as no action was taken April 28, 2017: Reopen to build testers from WPPL. When is this needed? Oct 25, 2017: Desmond will ask Edwin and provide CEA to build or to close out this request. 30 JAN 2018, No further information from Edwin or Desmond, should this be closed? 27-FEB-2018 No further information from Edwin or Desmond, should this be closed? 17 DEC, 2018: Edwin advise still discussing on the transfer under 2019 AOP Feb 25 March 2019 Edwin can we close this? | Matt Brule | Paul Deschenes | WPPL: Closed
161145 | Manufacturing Engineering | NA | 2016-10-12 | 999119905 | Desmond Ann | OPEN |  | Flex feedback that the fixture 999119905 is having a different press. It is using 99901154 instead of 999041264. The physical fixture do not have item 2, 5, 8,9,11 as stated in the 999119905 fixture drawing. Drawing for 999041250 is different from the physical part. Please see attached for details. | Nov 30, 2016: Correction is being made to the drawing two versions of the same part. Singapore uses a pneumatic press vs Milford Waters is a manual press. Tool qualification is not required documentation only. June 28, 2017: Both Singapore and Milford have MANUAL press, and Singapore has the latest version of the press due to cost saving changes per Steve. If the marked-up print was already provided to Steve to show the difference of the physical press in Singapore versus SAP drawing, Steve should help looking into updating the drawing to reflect the latest version. June 03, 2019: Pending Steve to update. | Stephen Wright |  | WPPL : Pending Steve to update
161186 | Test Engineering | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN |  | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | May 30, 2017: Model shop will cut the chasis for control box in first week of June. Chau continues putting the fixtures together. Chau to check with Eric about effect of the new PDA chip on these two universal testers. June 27, 2017: Model Shop just complete two sheetmetal boxes today to be picked up for the Universal Fine Alignement Fixture. Eric confirms that no effect of new PDA chip on these two universal testers. July 27, 2017: Works in progress. Sep 26, 2017: Plan to ship at end of month so Eric can set it up at Flex. (A few alignment tools might be shipped later on.) Oct 24, 2017: Shipped and being set up at Flex. Building Target Cell, Aperture. Mar 6, 2018: Flex has all parts short of the 2998 target. Oct 26, 2018: Waiting for WPPL to work on capital to build 1 or 2 control boxes. Dec 18, 2018: Edwin already replied on 2Dec18 that it is ok to utilize the budget for the CEA. Waiting for clarification from WPPL to know which budget to use and to build 1 or 2 control boxes. June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box | Eric Turner | Chau Nguyen | WPPL : CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box
171657 | Test Engineering | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OPEN |  | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. | Eric Turner |  | WPPL: Low priority per Eric.
171694 | Test Engineering | NA | 2017-04-27 | N/A | Yeap Siew Yee | OPEN |  | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | June 27, 2017: Both of these to be fixed with test software upgrade for SIgmasure, TE83026 is first followed by TE83013. 27-JUL-2017: this will be corrected during up-commoing software revision for Camstar. Sept 26, 2017: Work has started on an update to TE83013, TE83026 is next. 23-JAN-2018 Hope to get time in February to get this moving 30 JAN 2018 Working on TE83013, complete update (win 10, SIgmasure 7, fix OESRs etc) 27-FEB-2018 Actively working on this, Test running working on EDV and bugs as they are found 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. TE83026 will be started soon 20, AUG 2018, Fixing Ecord test bug, one duplicate tester built. 27-FEB-2019 Hope to get time to get this going in 2019 | Matt Brule |  | WPPL: Matt reply in May: Still at same status, second on my project priority list
171714 | Test Engineering | WAT242102 | 2017-05-05 | N/A | Yeap Siew Yee | OPEN |  | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | June 27, 2017: Have not looked into this one yet, should not be an issue, except specification alignment will take time. 27-JUL-2017: Have not looked into yet, will require specification alignment 20, AUG 2018 No action taken yet 29-Nov-2018 No Change 17-DEC-2018 No Change 27-FEB-2019 No Change | Matt Brule |  | WPPL : Matt reply in May 2019, Procedure update, low priority, no change, requires specification alignment
171914 | Manufacturing Engineering | N/A | 2017-07-31 | 999119592 | Desmond Ann | OPEN |  | While raising an ECR to update the PRO WAT045248, Roland found that tool 999119592 does not appear in SAP and could not locate a hard copy in the file. See attached on the drawing sketch for the fixture to be release into SAP. | Jul 31, 2017: The tooling drawing 999119592 will require at least a month to be released into the SAP. Desmond and Mark is okay with TMD010368 while waiting for tool drawing from Steve unless noticed otherwise. Desmond and Steve will follow up to get the tool drawing done to ECR. Aug 02, 2017: Desmond has volunteered to create a model file in solid works because this will speed up the completion of this request. SW Aug 2. June 03, 2019: WPPL will submit to Steve by 30th June. 17Dec18 - Pending Steve Wright to update the drawing. Waiting for completed prints from Desmond Ann. Mar 2019 | Stephen Wright | Mike Aubin | WPPL : WPPL JingXuan submitted to Steve 21Jun19, Pending Steve to update
171930 | Manufacturing Engineering | N/A | 2017-08-03 | N/A | Desmond Ann | OPEN |  | We have fabricated a jig to help in the assembly of the PCB. This jig is proposed to be used for all the detectors models listed in the matrix. The jig will have 4 bases to be used based on the model, see attached on the base used per the product matrix. This jig will be able to press the CPU board evenly into the standoff to the Personality card and also prevent flexing of the PCBAs during assembly. We will need Milford help to assign a tooling number and upload the drawing into SAP. | August 30, 2017: When Steve releases dwg into SAP, Desmond will complete ASP update. 17Dec18 - Yee Hwa will be performing the evaluation will be performed in Jan 2019. Duplicate jig is required for Waters Milford. June 03, 2019: WPPL will submit to Steve by 15th July. | Stephen Wright |  | WPPL : WPPL JingXuan submitted to Steve 21Jun19, Pending Steve to update
171993 | Test Engineering | 735002064 | 2017-08-25 | TE83013 | Yeap Siew Yee | OPEN |  | TE83013-1 is observed to experience a test sequence error at "Ecord Interface Test" by skipping to instruct to connect "Ecord" to the UUT while testing for 186015042. They have recently changed the tester PCBA P/N:210000367 (ASSY PCB CPSM CONTROLPCB) Rev H because existing P/N:210000367 Rev G was faulty on 14-Aug-2017. Refer to videos for failure occurrence. | 25-SEP-2017: In progress. 30-JAN-2018 Actively working on this 27-FEB-2018 working on bug fixes and EDV 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. 25, June 2018 Emailed Desmond this AM with proposed upgrade plan 20, AUG 2018 working on Ecord bug. 24-Nov-2018 Working on bug as time allows expect resolution soon, working to prove theory about Ecord formatting 17 DEC 2018 Same status (Pending Matt Brule.) 23-FEB-2019, expect to get freed up to complete this soon | Matt Brule |  | WPPL: Matt reply in May 2019, First on my project priority list
172012 | Manufacturing Engineering | N/A | 2017-09-01 | N/A | Chau Nguyen | OPEN |  | Work with Chau on Target Cells and other required tools and associated drawings to complete and transfer the two fixtures to WPPL by Mid-October 2017 . | Need a list of the targets SW......... not enough info provided ......Oct 2017. SW Oct 26, 2018: Chau will sort through components on hand to build spares. June 03, 2019: Universal Coarse Fine Alignment station pending Milford to fabricate. June 17, 2019: Universal Coarse Fine Alignment station already built and shipped. This request is to document the process to make Target Cells and other required tools to build spares when required. If there is enough materials on hands and available to build any target, it would be built if required. Chau Nguyen. | Stephen Wright | Chau Nguyen | WPPL: Milford to advise more details on what is required for this OESR.
172141 | Manufacturing Engineering | N/A | 2017-10-23 | 999029310 | Desmond Ann | OPEN |  | Flex feedback that the tooling 999029310 Fixture wheel to motor 2998 classic. Following sub-drawing are missing and not released:- 1. 999041435 Column, support filter wheel to motor. 2. 999041438 Bridge filter wheel to motor. | These parts exist in another fixture 999026549. They will be released with fixture 999026549. See ECO# 17Dec18 - Pending Steve Wright to provide ECO number to update drawing. | Stephen Wright |  | WPPL: Closed
182626 | Test Engineering | N/A | 2018-06-27 | TE82943 | Desmond Ann | OPEN |  | Missing serial number at the bottom of the test report (Ser.No) of Acquity BSM 18K PLUS and Acquity Arc QSM test in some Test record and it is encountered across all test stations TE82943 #1,#19,#20,#21,#22 and #24. Please kindly refer to the examples "BSM Test Record.JPG" and "QSM Test Record.JPG". Do help to resolve the issue. | Oct 2018: Made a Test Program code modification on the TE Lab PC and is under evaluation. 17Dec1: Check with Venture to see if this is resolved. 29Jan2019: It has not been resolved and still require support to solve the issue. 29-Jan-2019: A fix for this intermittent condition has been added to the latest test programs. It was going to be released along with Firebrick firmware in January. However, Firebrick firmware is now being rolled up into Gold firmware. So now it will be released in April with the Gold firmware release. Feb: same status. Apr23: same status. expecting release soon. | Edward Jackson |  | WPPL: Validation in progress.
182922 | Manufacturing Engineering | N/A | 2018-11-16 | N/A | Cheah Su Fern | OPEN |  | There are few errors found on the Part Drawing for Belt Tightening Fixture 999038298. Details will be attached in the email. | It looks like maybe a repair was made in Singapore by the CM that needs addressing. Area of concern is clearance only Desmond requested a print change to open up tolerance for concerns of an audit. June 03, 2019: Pending Steve to update the drawings. Drawing in Concept status. | Stephen Wright |  | WPPL : Pending Steve to update
193069 | Manufacturing Engineering | N/A | 2019-01-22 | N/A | Chan Yee Hwa | OPEN |  | ECO 9221720 & 9221721 changed SL Power Supply 279000852 to Delta Power Supply 279003675 for various Detectors. SigmaSure power supply part number need to update accordingly. We will need 7 model to be change today and 4 model to be change tomorrow due to a depletion. Request to update for below part number TODAY: 186002487 186017002 186015034 186017012 186017013 186002998 186015053 186299800 186015031 Request to update for below part number TOMORROW: 186248900 186017003 186015032 186015033 Thank you very much. |  | Matt Franklin |  | WPPL: Closed
193112 | Manufacturing Engineering | 279002224 | 2019-02-11 | N/A | Chan Yee Hwa | OPEN |  | Sigmasure changes in Detector 186015034 Acquity UPC2 PDA. Noticed additional Part Number 279002225 require to be scan since 02/10/2019 in Sigmasure. Previously were only 5 PN to be scan and now 6 in total. Request to remove additional PN 279002225 from 186015034. Please refer to attachment in email. |  | Matt Franklin | Roland Poisson | WPPL: Closed
193124 | Test Engineering | N/A | 2019-02-14 | TE82526 | Chan Yee Hwa | OPEN |  | Request additional Pump Tray Tester TE82526 to support Alliance Production for high demand. The tester is currently running at 92% capacity; Besides, we also see high chance of tester break down according to past years data. Please provide the cost for WPPL to raise CEA. Thanks. | June 03, 2019: Mark has begun looking at this, and is reviewing best path forward regarding component and PC obsolescence. CEA# WPPL19003 | Mark Jussaume | Chau Nguyen | WPPL : Mark reply in May2019, Mark has begun looking at this, and is reviewing best path forward regarding component and PC obsolescence. CEA# WPPL19003
193132 | Test Engineering | N/A | 2019-02-15 | N/A | Chan Yee Hwa | OPEN |  | Request help from Eric Turner to program 2 Shunts with Shunt ID for Flex production use. Thanks. |  | Eric Turner |  | Closed
193148 | Test Engineering | 735000615 | 2019-02-25 | TE82760 | Luo HuaZhen | OPEN |  | During meeting with Venture for First Pass Yield (FPY), the FLR detector yield was 3 consecutive months lower than target yield (97% for UPLC detector and 90% for 2475 detector). The venture feedback they're seeing Signal to Noise Ratio failure increased, it's the top failure after Raman/Rayleigh failure. The units are failed the 1st test with readings approximately zero (~0.00) and the 2nd test (retest) passed with over thousands (>1000). 55.56% of Signal/Noise Ratio failure are failed with readings nearly to "zero". And we tried to use Diagnostic on test program to troubleshoot this failure, but it's always blanked out. Could you please kindly help to check this issue and advise the possible root cause? Thanks |  | Matt Franklin | Eric Turner | WPPL: It is under monitoring while waiting for Eric to enable to diagnostic feature.
193162 | Test Engineering | N/A | 2019-03-01 | TE82988 | Luo HuaZhen | OPEN |  | WPPL received this chassis from CEI Batam after our spare chassis was used for this repair. The failure reported by CEI as shown as below. Could you please help to repair this chassis by replacing Power Supply in Chassis? Should I send this faulty Chassis to Milford for repair and test? Please kindly advise. Thank you 1. One power supply at GEOTEST chassis GX7310 having intermittently on-off problem. The Geotest chassis is in GTS tester TE82988-9. 2. Power Supply part number is PRM560. Requesting a replacement unit. 3. This GTS tester is mainly for testing PCBA for Waters UK (Micromass). | Rick Davis on March 01, 2019: "We do not repair the instruments in a GTS. We send them back to the manufacturer and have them repaired. I suggest you contact the appropriate Micromass person and have them arrange to have the chassis repaired. Chau, you need to assign this to someone from Micromass." | Luo Hua Zhen |  | Closed
193201 | Manufacturing Engineering | 289001458 | 2019-03-15 | N/A | Liu Jingxuan | OPEN |  | Tooling fixture number is stated wrongly in drawing 289001458 under Note 2. The correct tool number should be 999028067 instead of 999026067. | Request for resources submitted 23 April. Work was completed May 22 and will be released on ECO June 03, 2019: Pending Steve to update. | Stephen Wright |  | WPPL : Pending Steve to update
193244 | Test Engineering | WAT271129 | 2019-04-04 | 186000139 | Luo HuaZhen | OPEN |  | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. |  | Matt Brule |  | WPPL : Matt reply that he is working on it.
193266 | Test Engineering | N/A | 2019-04-12 | N/A | Chan Yee Hwa | OPEN |  | Would like to request for new shunts to support detector production test capacity at Flex Sg. We'll need below shunt with programmed chips. Thanks. i. 289002450 -- TUV Shunt ; 3 Ea. ii. 289002227 -- PDA Small Shunt; 1 Ea. iii. 289002226 -- PDA Medium Shunt; 3 Ea. | Request for CEA from WPPL to place actual order in SAP to determine which components available and which to be sourced to complete building these shunts. June 03, 2019: YH to raise PO for Shunts. | Eric Turner | Chau Nguyen | WPPL : Pending for WPPL to raise PO.
193267 | Test Engineering | 735001613 | 2019-04-15 | TE82882 | Luo HuaZhen | OPEN |  | New ECO is required Flex Singapore to supply 2 new spare Part Number 700012245 and 700012243. Please kindly help to review/update the TSP 735001613 to support new part number 700012245 and 700012243 test. Below is request sent by Flex, FYI 1. Problem Statement. - TE82882 PDA Final Test and TSP 735001613 Rev W does not have Spare OB PN 700012245 - Optics Bench, eLambda (On-Semi) and 700012243 - Assy, Optics Bench, 2998 (On-Semi) 2. Acceptance criteria that is not met. (e.g. relevant specification number and revision) - TE82882 PDA Final Test and TSP 735001613 Rev W must have to support Production run of the 2x Spare OB. 3. Relevant records (e.g. Incomin QA, Test Records, Photos, trackise NCRs numbers and related issue) - Relevant PDA Family Test Report for Spare Optic Bench have dedicated PN on the test station and TSP. 4. Info on whether this issue impacts revenue. - This impact the spare bench testing. |  | Matt Brule | Eric Turner | WPPL : Validation in progress.
193345 | Manufacturing Engineering | N/A | 2019-05-14 | 999045819 | Cheah Su Fern | OPEN |  | Drive Wheel Assembly of Spiral Wrap Fixture (999045819) need to be replaced due to wear and tear. Thus, would like to request Stephen Wright to purchase Drive Wheel Assembly from Automated Machine Design and ship to WPPL as replacement. | Steve on May 14, 2019: I would suggest to buy two sets for these are custom molded unless you are confident that the second machine the wheels will not fail in the near future. Parts ordered May 16 2019. SW. PO was released on June 4 there was an issue with Account number. June 03, 2019: Pending part fabricate by AMD. | Stephen Wright |  | WPPL : Pending part fabricate by AMD
193347 | Test Engineering | 735000629 | 2019-05-16 | TE83031 | Cheah Su Fern | OPEN |  | Bill of Material (BOM) for Spare kits (700010517) consist of “PN : 279003225 - ASSY OPTICS BENCH 2475 W/O FLOW CELL” and some packaging. 279003225 is use in both instrument 186247500 and 186017006 in Venture. Both 186247500 and 186017006 have different Flow Cell. Flow cell used in 186247500 is 289007905. Flow cell used in 186017006 is 289010016. Venture need to perform Functional test for the Optics Beach before ship it out. The Functional test require Flow cell to test together with Optics Bench. Please advise which Flow Cell should Venture use to test the 279003225? |  | Eric Turner |  | WPPL : Closed
193349 | Test Engineering | 735001645 | 2019-05-17 | TE82956 | Luo HuaZhen | OPEN |  | TE82956-2 failed test with high ambient temp (130 deg.C). Could it be this Thermistor faulty? What is PN of this Thermistor? Please kindly send us a repair part if you think it's a thermitor faulty too and you have a spare of thermistor. Thank you. Below is message from Venture, FYI TE82956-2 cannot perform "Column Heater Test" due to Thermistor Error. During initializing of the test, Ambient Temp is shown as 130 (Deg.C) instead of ambient temperature. Please refers to attached screenshorts for reference. We have performed as below. 1) Checked all cable connections between tester & UUT and PC & Tester. No issue found. 2) Checked cables connection inside of tester.No issue found. 3) Shut Down and restart PC.Then,retested. Still having same failure. 4) Tried to test using (x3) differnet UUTs and still having same failure. Request the help on this tester issue. | 2 | Tony Fontes |  | WPPL : Verification in progress in Venture.
193363 | Manufacturing Engineering | WAT271087 | 2019-05-23 | N/A | Liu Jingxuan | OPEN |  | Tube Press Tool 999027699 is called out in Tooling Required on Sheet 6 of 8, which is no longer in use. Please remove Tube Press Tool 999027699 in Grind A4 of Sheet 6 of 8 in WAT271087 PRO file. | Requested information is the tool is required. 5/23 SW Request replied on 24 May. Pending Steve | Stephen Wright |  | WPPL: Closed
193373 | Test Engineering | 735000698 | 2019-05-29 | TE82774 | Edward Jackson | OPEN |  | IQ has not been preformed on assemblies to be tested on this tester. Please perform all validations required. See Emma Pierce in Milford if there are any validation questions or concerns. Thank you. |  | Luo Hua Zhen | Desmond Ann | WPPL: Pending execution of IQ.
193374 | Manufacturing Engineering | N/A | 2019-05-30 | N/A | Liu Jingxuan | OPEN |  | Based on the DHR, sigmasure entry have input the details for SN:D18QSM164A, PN:186015018. However, no record can be found when search the Genealogy information for D18QSM164A. |  | Eric Turner |  | WPPL: Closed
193379 | Test Engineering | WAT271129 | 2019-06-05 | N/A | Mark Jussaume | OPEN |  | This request is for "width" test data from Alliance instrument Final test. Supply 100 data entries from previously tested instruments. While reviewing the test specs for the new Alliance test system EDV, we discovered a discrepancy between the instrument performance and TSP specification for the "width". This needs to be addressed before moving forward with the new test system as well as updating the current test. The Performance Specification document lists the "width" spec at 70 - 90 steps as limits and the TSP has 60 - 90 steps. In an effort to align the test to the instrument performance document, we are looking at changing the spec to 70 - 90 steps. Before we do so I need to understand how this may affect test yields. |  | Yeehwa Chan | Luo Hua Zhen | WPPL: Closed
193382 | Test Engineering | N/A | 2019-06-06 | TE82743 | Luo HuaZhen | OPEN |  | Venture is running QDA production too. Some of testers are using dynamic IP address and those testers are sharing the same network with Milford testers. Therefore it can cause IP address conflict, while tester from Milford is power off and the tester from QDA production is assigned a same IP address by server. Can Milford assign a range of IP address for QDA testers? I will work with local team/Waters UK to check if they can assign a fixed IP address within this range for each of testers for QDA production. So we can avoid the same issue happening again. |  | Bill Bates |  | WPPL: Milford to advise the range of IP address to use for UK equipments.
193393 | Manufacturing Engineering | N/A | 2019-06-11 | 999041977 | Chan Yee Hwa | OPEN |  | WPPL has shipped over 2 backpressure assembly to Milford attn to Anthony for Calibration Purpose. Please help to calibrate and return to WPPL by end Jun as Keypoint (Supplier) will require the tool for calibration in early July. Thanks. |  | Paul Deschenes |  | WPPL : Milford to advise status on the calibration.
193410 | Test Engineering | N/A | 2019-06-24 | TE82895 | Luo HuaZhen | OPEN |  | Tester TE 82895 software unable to run due to disk space full ('C' Driver). Temporary solution was transferred some data files from Big size folder location : C:\Empower\Oracle\admin\WAT6\bdump to D driver: D:\Empower\Oracle\admin\WAT6\bdump. The tester can be running now after data transferred (clean up C driver). Please kindly review these *.trc files and advise actions to these autogenerated files. Flex can do manual delete or delete by scripts regularly if the files are not important. |  | Bill Bates |  | WPPL: Pending Milford to advise.
193411 | Test Engineering | N/A | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN |  | New FLR Detector tester (Win7) don’t get the Factory Order and TMD # on new test report template and the new template format is slightly different from xp template although the same parameters were tested. To standardize the report, we need Milford help to update the report templates to the same format as the one on the old testers(WinXP). Please refer to page 2 for details of differences on 2 report templates. |  | Eric Turner |  | WPPL : Pending Eric to update.
193413 | Test Engineering | 735000615 | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN |  | TSP735000615 Rev P, page 8 of 9, the TSP is missing a step during keypad test after passing the functional test . For 2475 mode, we need to set the date for the lamp installed. Press “DIAG”, “7” ,” 1” as per red line TSP attached at page 2 of this OESR Please review this step and update the TSP if necessary. Thanks |  | Eric Turner |  | WPPL : Pending for Milford to update TSP.