## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2 | Minutes on Feb 27, 2019 | Minutes on March 30, 2020 and Apr 30, 2020 | WPPL update on 04 May 2020 | WPPL update on 29 May 2020
161145 | Manufacturing Engineering | NA | 2016-10-12 | 999119905 | Desmond Ann | OFI |  | Flex feedback that the fixture 999119905 is having a different press. It is using 99901154 instead of 999041264. The physical fixture do not have item 2, 5, 8,9,11 as stated in the 999119905 fixture drawing. Drawing for 999041250 is different from the physical part. Please see attached for details. | Nov 30, 2016: Correction is being made to the drawing two versions of the same part. Singapore uses a pneumatic press vs Milford Waters is a manual press. Tool qualification is not required documentation only. June 28, 2017: Both Singapore and Milford have MANUAL press, and Singapore has the latest version of the press due to cost saving changes per Steve. If the marked-up print was already provided to Steve to show the difference of the physical press in Singapore versus SAP drawing, Steve should help looking into updating the drawing to reflect the latest version. June 03, 2019: Pending Steve to update. | Stephen Wright |  | Desmond and HuaZhen will meet with Edwin to devise a plan to take care of it if possible. | OFI is okay with WPPL
161186 | Test Engineering | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN |  | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | Open since 2017, June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. March 30, 2020: Matt will dig in tomorrow, and understand what is needed, and we will see if we can get someone on it. APR 30, 2020: To advise a date - Huazhen | Eric Turner | Chau Nguyen | Chau to confirm with Eric if Matt and his group are working on it. | View updates from Matt | Checked with Flex, can put this acitivity KIV till 30 Jun | KIV
171657 | Test Engineering | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OFI |  | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. March 30, 2020: Switched from Eric to Matt, hoping to get some traction on this this month, while folks are working @ home | Matt Brule |  | HuaZhen will confirm with Venture and inform everyone involved. | OFI after WPPL verified with CM | OFI, Verified with CM
171694 | Test Engineering | NA | 2017-04-27 | N/A | Yeap Siew Yee | OPEN |  | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | Open Since 2017, 27-FEB-2019 Hope to get time to get this going in 2019 May 2019: Still at same status, second on project priority list March 30, 2020: HTCH updates looking like 2nd quarter completion APRIL 30, 2020: End Jun to complete the TE83026 upgrade. | Matt Brule |  | On-going | View updates from Matt |  | Stephen Wright to check with Matt B on the progress
171714 | Test Engineering | WAT242102 | 2017-05-05 | N/A | Yeap Siew Yee | CLOSED |  | Proposed to update TSP Doc No: WAT242102 Rev 4 1) To include Waters PN 279003385, 279002592, 279003250 on the cover sheet | June 27, 2017: Have not looked into this one yet, should not be an issue, except specification alignment will take time. 27-JUL-2017: Have not looked into yet, will require specification alignment 20, AUG 2018 No action taken yet 29-Nov-2018 No Change 17-DEC-2018 No Change 27-FEB-2019 No Change July 02, 2019: WPPL : Matt reply in May 2019, Procedure update, low priority, no change, requires specification alignment March 30, 2020: Hoping to get some traction on this this month, while folks are working @ home. | Matt Brule |  | WPPL will investigate if it's low priority to close or change status | Switch to OFI
171993 | Test Engineering | 735002064 | 2017-08-25 | TE83013 | Yeap Siew Yee | CLOSED |  | TE83013-1 is observed to experience a test sequence error at "Ecord Interface Test" by skipping to instruct to connect "Ecord" to the UUT while testing for 186015042. They have recently changed the tester PCBA P/N:210000367 (ASSY PCB CPSM CONTROLPCB) Rev H because existing P/N:210000367 Rev G was faulty on 14-Aug-2017. Refer to videos for failure occurrence. | Open since 2017, 23-FEB-2019, expect to get freed up to complete this soon July 02, 2019: Matt reply in May 2019, First on his project priority list March 30, 2020: HTCH updates looking like 2nd quarter completion APRIL 30, 2020: End Jun to complete the TE83026 upgrade. | Matt Brule |  | On-going | View updates from Matt | This OESR related to OESR 171694, so can close this OESR and keep track under OESR 171694
172012 | Manufacturing Engineering | N/A | 2017-09-01 | N/A | Chau Nguyen | OFI |  | Work with Chau on Target Cells and other required tools and associated drawings to complete and transfer the two fixtures to WPPL by Mid-October 2017 . | Need a list of the targets SW......... not enough info provided ......Oct 2017. SW Oct 26, 2018: Chau will sort through components on hand to build spares. June 03, 2019: Universal Coarse Fine Alignment station pending Milford to fabricate. June 17, 2019: Universal Coarse Fine Alignment station already built and shipped. This request is to document the process to make Target Cells and other required tools to build spares when required. If there is enough materials on hands and available to build any target, it would be built if required. Chau Nguyen. June 17, 2019: WPPL: Milford to advise more details on what is required for this OESR. These will be required when 2 control boxes completed. Request for completion date of 2 control boxes from Eric to provide to Finance in WPPL. | Stephen Wright | Chau Nguyen | Chau will continue working with Stephen on this request | Taking care of higher priority projects and will get to it. Chau will email Steve the information collected so far and work with Steve to write up an instructional document to make these target cells and tools. KIV - 1st Aug 2020 30th Apr 2020 - OFI Part no# of Targets cells or take photos. Guan Yee to assist | Attach with the target cells PN list.
182922 | Manufacturing Engineering | N/A | 2018-11-16 | N/A | Cheah Su Fern | CLOSED |  | There are few errors found on the Part Drawing for Belt Tightening Fixture 999038298. Details will be attached in the email. | It looks like maybe a repair was made in Singapore by the CM that needs addressing. Area of concern is clearance only. Desmond requested a print change to open up tolerance for concerns of an audit. June 03, 2019: Pending Steve to update the drawings. Drawing in Concept status. July 02, 2019: WPPL : Pending Steve to update | Stephen Wright |  | Desmond and HuaZhen will meet with Edwin to devise a plan to take care of it if possible. | Steve will provide ECO 30th Apr 2020: To close the OESR
193124 | Test Engineering | N/A | 2019-02-14 | TE82526 | Chan Yee Hwa | CLOSED |  | Request additional Pump Tray Tester TE82526 to support Alliance Production for high demand. The tester is currently running at 92% capacity; Besides, we also see high chance of tester break down according to past years data. Please provide the cost for WPPL to raise CEA. Thanks. | June 03, 2019: Mark has begun looking at this, and is reviewing best path forward regarding component and PC obsolescence. CEA# WPPL19003 Updated CAD file with new bracket and bracket was added manual parts list S.W. Sept. 5 Fluidic diagram needs to be updated per change. 24 JAN 2020, MB: Additional system, and parts upgrade kits, shipped Jan 7. March 30, 2020: WPPL provided everything necessary (we believe) and are actively updating and validating | Mark Jussaume | Mike Aubin | Mark can close if no further actions is required. | Closed
193244 | Test Engineering | WAT271129 | 2019-04-04 | 186000139 | Luo HuaZhen | OPEN |  | Flex Singapore feedback this Alliance Final Test workstation 3 has frequently hanging issue. They request to replace this workstation with a windows 7 workstation. Can you please help to access if this PC hang issue can be fixed? or have to replace the workstation? If replacing workstation, can you please advise the spec of workstation for Alliance Final Tester? Thanks. Below is message from Flex, 1. Problem Statement : Alliance Final Test Station 3 having frequent hang during Final Test execution . Refer to attached file. 2. Acceptance Criteria Not Met: The test station must be able to execute Final Tester of 4 units simultaneously without system hang during test . Not able to complete test an need to restart the CPU leaving down station 3. Info on whether this issue impacts revenue : No impact yet to revenue but sure one day will hit when the numbers are up. This is to prepare for upsize of Alliance. | July 02, 2019: WPPL : Matt reply that he is working on it. March 30, 2020: Have not had time to look at this again. APRIL 30, 2020: On going troubleshooting | Matt Brule |  | On-going | View updates from Matt |  | Stephen Wright to check with Matt B on the progress
193266 | Test Engineering | N/A | 2019-04-12 | N/A | Chan Yee Hwa | CLOSED |  | Would like to request for new shunts to support detector production test capacity at Flex Sg. We'll need below shunt with programmed chips. Thanks. i. 289002450 -- TUV Shunt ; 3 Ea. ii. 289002227 -- PDA Small Shunt; 1 Ea. iii. 289002226 -- PDA Medium Shunt; 3 Ea. | Request for CEA from WPPL to place actual order in SAP to determine which components available and which to be sourced to complete building these shunts. June 03, 2019: YH to raise PO for Shunts. July 02, 2019: WPPL : Pending for WPPL to raise CEA (and or PO?) to enter order. Duplicate part. no qualification required. | Chau Nguyen | Eric Turner | Chau to ask Matt and his group to ship it to WPPL to bring to CM. | Bill Bates has shipped 27th Apr - Airwaybill? Huazhen - to close the OESR if received. | Waiting for the shunts from Milford
193382 | Test Engineering | N/A | 2019-06-06 | TE82743 | Luo HuaZhen | CLOSED |  | Venture is running QDA production too. Some of testers are using dynamic IP address and those testers are sharing the same network with Milford testers. Therefore it can cause IP address conflict, while tester from Milford is power off and the tester from QDA production is assigned a same IP address by server. Can Milford assign a range of IP address for QDA testers? I will work with local team/Waters UK to check if they can assign a fixed IP address within this range for each of testers for QDA production. So we can avoid the same issue happening again. | July 02, 2019: WPPL: Milford to advise the range of IP address to use for UK equipments. This is urgent to get IP conflicts resolved for CM Production. 31-Oct-2109: Per WPPL, info was provided to have 50 static IP to be used for UK testers. Email sent on 30 Sep 2019. Pending for Milford/UK to advise. 26 Feb 2020: Chau has emailed to contact Raymond but no responses yet | Chau Nguyen |  | Chau to ask management to work with Raymond to help resolve the issue. | Chau or WPPL can create BT request to have WPPL contact and work directly with BT to resolve the issue as Raymond in BT probably left Waters. APRIL 30th: Switch added - Huazhen to verify Equipment IP address to be fixed* - Huazhen to ensure. To close this OESR. Huazhen to review with team on Monday. | Switch added with 48 ports, confirmed with team. Use IP address list from server to check with UK see if can fix the IP address -YH; After getting fixed IP, to update Bill Iron on latest list.
193411 | Test Engineering | N/A | 2019-06-24 | TE82760 | Luo HuaZhen | CLOSED |  | New FLR Detector tester (Win7) don’t get the Factory Order and TMD # on new test report template and the new template format is slightly different from xp template although the same parameters were tested. To standardize the report, we need Milford help to update the report templates to the same format as the one on the old testers(WinXP). Please refer to page 2 for details of differences on 2 report templates. | July 02, 2019: WPPL : Pending Eric to update. March 30, 2020: Without rewriting the test code, there is nothing we can do at this point. I think they have done some validation and may be allowing for its use. Please update after your meeting. | Eric Turner |  | Eric updated the tempalte | Closed
193413 | Test Engineering | 735000615 | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN |  | TSP735000615 Rev P, page 8 of 9, the TSP is missing a step during keypad test after passing the functional test . For 2475 mode, we need to set the date for the lamp installed. Press “DIAG”, “7” ,” 1” as per red line TSP attached at page 2 of this OESR Please review this step and update the TSP if necessary. Thanks | July 02, 2019: WPPL : Pending for Milford to update TSP. 26 Feb 2020: Reassigned to Matt per Eric. March 30, 2020: Hoping to get some traction on this this month, while folks are working @ home. | Matt Brule | Eric Turner | Matt Brule has accepted the reassignment on Feb 26, 2020. | View updates from Matt |  | Matt B to update; redline was provided
193459 | Test Engineering | 735002010 | 2019-07-29 | TE82943 | Luo HuaZhen | OFI |  | Since last year, we were facing multiple(4 times) BSM/QSM tester breakdown due the weighing balance faulty in Flextronics. Each time it will take around 3 months to get the weighing balance repaired in Japan due to lack of repair parts in ITS Singapore (AnD local representative). The repair turnaround time is very long and costly (around USD1500 for each repair). So we checked with ITS again recently regarding supportability of this HR200 model. They sent us a product discontinuation letter and its successor model GR-200 info as attached Can we replace the current old HR-200 weighing balance with its successor mode GR-200 since these HR200 are frequently breaking down? If can, how can we proceed with this replacement? Or do you have any other better solution for HR200 support? Please refer to GR200 user manual and discontinuation letter in Pg2 & 3. Thank you | 21-Aug-2019: A discussion took place between Jaaffar Ismail, Matt Brule, and Ed Jackson. We decided to first pursue a direction of maintaining the HR-200 balances. We will ask Paul Deschenes next week if he can identify any spare balances throughout the building that we may be able to use to support the TE82943 testers. The long term solution, at least a year away, is to replace the TE82943 testers with a new design that will eliminate the balances by replacing them with flow meters. Ed J.. 23-Sep-2019 BB - Shipping a spare HR-200 balance to HuaZhen's attention. This will need Calibration once received in Singapore/when installed on tester setup. 31-Oct-2019: Per WPPL, the spare balance was received and tester is working now. Ed suggests to leave this OESR open until he completes the viability study of new model because the old model is no longer available to be purchased from supplier. 10-Jan-2020: The GR200 internal function table parameters were updated to match the HR200 parameters. A temporary rig was made to fix the TE82943 balance tubes into a beaker placed inside the GR200 balance. The BSM test program was run several times to collect data and verify the operation. The GR200 balance values matched the HR200 balance values for the UUT, as expected. A procedure page detailing how to setup the GR200 balance was created and is located in the design file in the Test Equipment folder under TE82943 Project. EdJ. 10-Jan-2020: I am requesting some Mechanical Engineering help for the next step in this project. The GR200 balance will require a top plate to hold the fluidic interface. Same design as the HR200 but modified to fit the GR200. A quick discussion and viewing will show the idea clearly. In addition, the GR200 will require the addition of the existing style tubes that reach into the beaker. Thank you. EdJ. 10 JAN 2020: MB reassigned to Chau to get a mechanical resource assigned. | Chau Nguyen | Edward Jackson | Chau to prioritize to provide mechancial design for Ed. | Taking care of higher priority projects and will get to it. OFI till new design.
193460 | Test Engineering | N/A | 2019-07-29 | TE82475 | Luo HuaZhen | OPEN |  | Based on tester utilization calculation, this pressure decay tester in Flex has been used up 91.56% of total tester capacity. So we plan to buy another TE82475 tester by raising CEA. Please kindly advise the total price and estimated complete date to build this tester. Thank you Tester utilization data, FYI Tester Utilization based on Q1 & Q2’20 numbers: Prime is 79.3% Prime + Spares is 91.56% Q3 & Q4’19 numbers: Prime is 78.2% Prime + Spares is 91.5% | 26 Feb 2020: Redesign under way March 30, 2020: New tester in design proof stage now APRIL 30, 2020: In progress | Mark Jussaume |  | HuaZhen will meet Edwin to review and provide CEA to Milford when approved. | View updates from Matt |  | Stephen Wright to check with Mark J on the progress
193551 | Test Engineering | 735002011 | 2019-09-02 | TE83032 | Luo HuaZhen | OPEN |  | The existing optics bench alignment stations are using a Golden flow cell to align all the optics benches. The flow cell used to align 5029 FLR optics benches had been using since year 2013 product transferred period and it had never been replaced. Venture observed the white colour residue on the glass window of the flow cell and it can not be removed although it was purging/priming through with MeOH for long hours. Refer to attached slide. Venture should have one back up golden flow cell in any case of flow cell degrading over a few years of usage and also for not interrupting production line while it waiting Milford to send over new Golden flow cell. | 30 APRIL 2020: 5 x optic bench sent. Pending Matt Franklin to use on testing. | Eric Turner | Matt Franklin | Edwin is aware of the concern and issue from Quality. HuaZhen will follow up with Edwin to support Matt to get it done. | Matt Franklin is working on it and will provide to keep provide progress. |  | Matt F is working on this project, almost complete with the testing; Now it's working on how many pilot run units required in Q3;
193552 | Test Engineering | 73500615 | 2019-09-02 | TE82760 | Luo HuaZhen | OPEN |  | On the month of Jul, Venture production found some Ram-Ray Valley Energy failures during FLR detectors final test for 5029 models. The UUTs were passed when they swapped with new flow cells. The failed flow cells were used in another UUTs and passed the test. Refer to attached slides for failure details. Raised OESR to follow up and get some advices from WPPL/Milford Engineering to look into the possible root causes of this failure. |  | Eric Turner | Matt Franklin | HuaZhen will follow up with CM and confirm to close. | HuaZhen will provide information from CM to close it out. | Info was provided. Closed | Stephen Wright to check and close this OESR
193757 | Test Engineering | N/A | 2019-11-08 | TE83021 | Luo HuaZhen | OPEN |  | Venture reported this fixture shows error message ' Serial CMD XCVR. The operation time out'. Suspect the CPU board of this fixture is faulty. Please kindly send us a right firmware for this CPU board for this repair or send us a CPU with firmware. Right now 2 pcs of PCB are pending for this test. | Firmware sent 13-NOV-2019, Matt B March 30, 2020: Performing FCT in instrument via TMD while fixture is repaired, refurbed and prepped for conversion to GTS II | Matt Brule |  | On-going Modify TSP to reflect instrument testing. | View updates from Matt | Matt: On-going Modify TSP to reflect instrument testing. | Matt B to revise TSP
193786 | Test Engineering | 735002510 | 2019-11-21 | TE82865 | Luo HuaZhen | OPEN |  | CEI request for test program update to allow enough time for plate actuation adjustment. There for reducing the false failure due to test program run out time while adjusting the flow control knob. Please kindly advise a way to extend the time for this adjusting or help to remotely update the test software. Below is the detail of reported issue, CEI is facing numbers of failure on the adjust plate latch actuation speed during functional setup for CPSO model. Production adjusted according to the adjust Latch Actuation Speed step shown on screen by preset plate latch flow knob controls by loosening the stops nuts, then rotate the flow control knob clockwise all the way, then turn counterclockwise 1.5 turns on the upper flow control and 0.75 turns on the lower flow control. Production feedback when click 'OK' they don’t have enough time to adjust the knob before test program pop out to "fail". This will result in final test fail and affecting FPY. | 19-BEC-2019 Matt B. CM looking at adjusting "preset" process to avoid failure. If successful this will be moved to OFI and test program will be updated at next change. March 30, 2020: CM looking at, using a preset, as a stop gap to program upgrade. Update from WPPL? | Matt Brule | Richard Davis | WPPL to follow up with CM and provide updates | View updates from Matt. HuaZhen will provide information to close it out. Closed | Info was provided. Closed | Stephen Wright o check and close
203852 | Test Engineering | N/A | 2020-01-10 | TE83044 | Luo HuaZhen | OPEN |  | Flex Malaysia request to have a spare CPU with firmware for the following tester fixtures. Most of these CPU are obsoleted. Please kindly advise a spare solution. Can you please send us a spare if you have same part number? Or please advise if can use 520 CPU2K board with correct firmware version from you? Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 pls refer to page 2 attachment 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 | March 30, 2020: Collecting all necessary PCBs and firmware APRIL 30, 2020: Get information on GTS with individual CPU under teststand. Install the fixture with CPU - Huazhen | Paul Deschenes | Tony Fontes | Waiting for updates from Paul | View updates from Matt. (Tony was no longer with Waters) | Install actual fixture and update Matt the date of each fixture. - HZ | Firemware version was checked by Matt B; Matt B may send the spare with firmware loaded
203936 | Test Engineering | N/A | 2020-02-10 | TE83044 | Luo HuaZhen | OPEN |  | Flex Malaysia GTS has 2 fixtures down, TE83044 and TE83002. Both fixtures were found its CPU2k board faulty. Flex needs a replacement of this CPU board. Please kindly send a replacement CPU2k board with firmware loaded for this 2 fixtures repair. Or please send the firmware and schematic of fixtures to allow team to work on troubleshooting. The original CPU card is 210000273 Rev G. If fixtures cannot be fixed before 19 Feb 2020, the shipment of following part numbers will be affected. 186015031; 186015032; 186015033 and 186015080 | March 30, 2020: Collecting all necessary PCBs and firmware | Paul Deschenes | Bill Bates | On-going | To review if we can close? -Huazhen | Can close this OESR as a samilar OESR 203852 is under tracking | Stephen Wright o check and close
203992 | Test Engineering | n/a | 2020-02-28 | TE83051 | Jian Wang | OPEN |  | Please provide a 2475-Acquity FLR optics bench test system and a 2998/Acquity PDA optics bench test system to the repair center, Beijing. | March 30, 2020: Matt is working on it, and this is not WPPL (?!?!) | Matt Brule |  |  | View updates from Matt. Close this OESR |  | Stephen Wright to check and close
204004 | Test Engineering | N/A | 2020-03-06 | TE82743 | Luo HuaZhen | OPEN |  | CEI Batam report their RACAL system is down. The test fixture TE82839 failed power up initialization. Vref measured 0V spec. 4.09V to 4.103V. CEI team suspect the following 2 modules could be the cause for failure. Please kindly advise if you have below requested parts for repair or troubleshooting tips to fix the initialization failing issue. 1. 1260-100, PN: 407655-01T (Adapt A-Switch Carrier module with 4 cards inside); 2. 6065P-8, PN:70-0036-011 (Message base servant interrupter) The following order shipment may be affected. WA3362-G070 42 PO# 4502444505 USA WA6471-G043 10 PO# 4502447279 USA WA6472-G061 15 PO# 4502452989 USA WA7044-G029 20 PO# 90490 EQM | March 30, 2020: Working on it. | Paul Deschenes |  |  | Closed |  | Stephen Wright to check and close
204089 | Test Engineering | N/A | 2020-04-23 | TE82720-1 | Luo HuaZhen | OPEN |  | A spare SMU Keithley 236 failed calibration on 19 Mar 2020. Waters Quality checked with Flex Malaysia and found that this spare SMU was used for testing in last Aug 2019 until Sept. 2019. So, Teams needs to know this Keithley 236 Operation range in this tester to do risk assessment for the PCBA were shipped during last Aug 2019 to Sep 2019. The failing calibration report and a test record for a unit tested on Aug 22 2019 are attached at Page 2, for your reference. Please kindly advise what is Keithley 236 operation range in terms Current & Voltage on this Front End Tester. Because test record shows in ADC counts which we cannot do comparison with Failing calibration report values (Current & Voltage) from Keysight. Could you please advise on whether the out of spec range from failing report is affecting the tester operating range? Please kindly advise the disposition about the PCBA tested in last Aug 2019 to Sep 2019, especially 15 pcs PCBA shipped out as spares. |  | Matt Brule |  |  | Matt Brule working on this- will send email to WPPL |  | Matt B to close this OESR
204131 | Other | 289001974 | 2020-05-18 | Aperture adjustement station | Luo HuaZhen | OPEN |  | Flex reports 2998 Aperture Adjust Station unable to proceed oer 4.5 onwards as ASP 289001974.Reading on position (tenths from centre) specs: +/-2 tenths Alisngment is not caputred. Value remains at '0' even after adjusting the micrometer. Only 100µm aperture is affected. 50µm aperture can still proceed all operation as per normal and station is passed daily calibration. Please kindly advise repair steps for this failure or do a remote troubleshooting on IP address 10.249. 26.64. Flex PDA production is affected by this issue. The picture on page 2 is the reading we cannot see it change on 100m aperture, for your reference. | [6 empty] | Stephen Wright will share troubleshooting tips with local team and work with CM to improve the troubleshooting in future; OESR is under monitoring now