## Sheet1
Log Number | Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Problem Summary | Summary | Assignee 1 | Assignee 2 | Minutes on Feb 27, 2019 | Minutes on March 30, 2020 and Apr 30, 2020 | WPPL update on 04 May 2020 | 25 Jun 2020 review udpate | October 1 2020 update | October 28 2020 Update | November 23, 2020 Update
161145 | ME | NA | 2016-10-12 | 999119905 | Desmond Ann | OFI | Flex feedback that the fixture 999119905 is having a different press. It is using 99901154 instead of 999041264. The physical fixture do not have item 2, 5, 8,9,11 as stated in the 999119905 fixture drawing. Drawing for 999041250 is different from the physical part. Please see attached for details. | Nov 30, 2016: Correction is being made to the drawing two versions of the same part. Singapore uses a pneumatic press vs Milford Waters is a manual press. Tool qualification is not required documentation only. June 28, 2017: Both Singapore and Milford have MANUAL press, and Singapore has the latest version of the press due to cost saving changes per Steve. If the marked-up print was already provided to Steve to show the difference of the physical press in Singapore versus SAP drawing, Steve should help looking into updating the drawing to reflect the latest version. June 03, 2019: Pending Steve to update. | Stephen Wright |  | Desmond and HuaZhen will meet with Edwin to devise a plan to take care of it if possible. | OFI is okay with WPPL
161186 | TE | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | Open since 2017, June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. March 30, 2020: Matt will dig in tomorrow, and understand what is needed, and we will see if we can get someone on it. APR 30, 2020: To advise a date - Huazhen | Steve Wright | Matt Brule | Chau to confirm with Eric if Matt and his group are working on it. | View updates from Matt | Checked with Flex, can put this acitivity KIV till 30 Jun | Matt/Stephen will continue after shutdown; add TE No/Tool. HZ check with Flex and send it | Matt: Q4 2020 estimate completion date | Working on ordering parts with China Testers. Reasign….
171657 | TE | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OFI | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. March 30, 2020: Switched from Eric to Matt, hoping to get some traction on this this month, while folks are working @ home | Matt Brule |  | HuaZhen will confirm with Venture and inform everyone involved. | OFI after WPPL verified with CM | OFI, Verified with CM |  |  | TSP update next week | Still pending
171694 | TE | NA | 2017-04-27 | N/A | Yeap Siew Yee | Open | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | Open Since 2017, 27-FEB-2019 Hope to get time to get this going in 2019 May 2019: Still at same status, second on project priority list March 30, 2020: HTCH updates looking like 2nd quarter completion APRIL 30, 2020: End Jun to complete the TE83026 upgrade. | Matt Brule |  | On-going | View updates from Matt |  | Target to complet it by Q3 2020 | Primier products added; TE0083013 will be shipping to Venture; target to complete upgrade by Q4 2020; | Will Clean out today or tomorrow | Pending, keeping open to track updates to TE83013 and TE83026
172012 | TE | N/A | 2017-09-01 | N/A | Chau Nguyen | OFI | Work with Chau on Target Cells and other required tools and associated drawings to complete and transfer the two fixtures to WPPL by Mid-October 2017 . | Need a list of the targets SW......... not enough info provided ......Oct 2017. SW Oct 26, 2018: Chau will sort through components on hand to build spares. June 03, 2019: Universal Coarse Fine Alignment station pending Milford to fabricate. June 17, 2019: Universal Coarse Fine Alignment station already built and shipped. This request is to document the process to make Target Cells and other required tools to build spares when required. If there is enough materials on hands and available to build any target, it would be built if required. Chau Nguyen. June 17, 2019: WPPL: Milford to advise more details on what is required for this OESR. These will be required when 2 control boxes completed. Request for completion date of 2 control boxes from Eric to provide to Finance in WPPL. | Stephen Wright | Chau Nguyen | Chau will continue working with Stephen on this request | Taking care of higher priority projects and will get to it. Chau will email Steve the information collected so far and work with Steve to write up an instructional document to make these target cells and tools. KIV - 1st Aug 2020 30th Apr 2020 - OFI Part no# of Targets cells or take photos. Guan Yee to assist | Attach with the target cells PN list. |  |  | Will look at this while duplicating optics testers (161186)
193413 | TE | 735000615 | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN | TSP735000615 Rev P, page 8 of 9, the TSP is missing a step during keypad test after passing the functional test . For 2475 mode, we need to set the date for the lamp installed. Press “DIAG”, “7” ,” 1” as per red line TSP attached at page 2 of this OESR Please review this step and update the TSP if necessary. Thanks | July 02, 2019: WPPL : Pending for Milford to update TSP. 26 Feb 2020: Reassigned to Matt per Eric. March 30, 2020: Hoping to get some traction on this this month, while folks are working @ home. | Matt Brule | Eric Turner | Matt Brule has accepted the reassignment on Feb 26, 2020. | View updates from Matt |  | In progess | Thought this had been reassigned, it had not will complete next week | Thought this had been reassigned, it had not will complete next week | Still Pending
193459 | TE | 735002010 | 2019-07-29 | TE82943 | Luo HuaZhen | OFI | Since last year, we were facing multiple(4 times) BSM/QSM tester breakdown due the weighing balance faulty in Flextronics. Each time it will take around 3 months to get the weighing balance repaired in Japan due to lack of repair parts in ITS Singapore (AnD local representative). The repair turnaround time is very long and costly (around USD1500 for each repair). So we checked with ITS again recently regarding supportability of this HR200 model. They sent us a product discontinuation letter and its successor model GR-200 info as attached Can we replace the current old HR-200 weighing balance with its successor mode GR-200 since these HR200 are frequently breaking down? If can, how can we proceed with this replacement? Or do you have any other better solution for HR200 support? Please refer to GR200 user manual and discontinuation letter in Pg2 & 3. Thank you | 21-Aug-2019: A discussion took place between Jaaffar Ismail, Matt Brule, and Ed Jackson. We decided to first pursue a direction of maintaining the HR-200 balances. We will ask Paul Deschenes next week if he can identify any spare balances throughout the building that we may be able to use to support the TE82943 testers. The long term solution, at least a year away, is to replace the TE82943 testers with a new design that will eliminate the balances by replacing them with flow meters. Ed J.. 23-Sep-2019 BB - Shipping a spare HR-200 balance to HuaZhen's attention. This will need Calibration once received in Singapore/when installed on tester setup. 31-Oct-2019: Per WPPL, the spare balance was received and tester is working now. Ed suggests to leave this OESR open until he completes the viability study of new model because the old model is no longer available to be purchased from supplier. 10-Jan-2020: The GR200 internal function table parameters were updated to match the HR200 parameters. A temporary rig was made to fix the TE82943 balance tubes into a beaker placed inside the GR200 balance. The BSM test program was run several times to collect data and verify the operation. The GR200 balance values matched the HR200 balance values for the UUT, as expected. A procedure page detailing how to setup the GR200 balance was created and is located in the design file in the Test Equipment folder under TE82943 Project. EdJ. 10-Jan-2020: I am requesting some Mechanical Engineering help for the next step in this project. The GR200 balance will require a top plate to hold the fluidic interface. Same design as the HR200 but modified to fit the GR200. A quick discussion and viewing will show the idea clearly. In addition, the GR200 will require the addition of the existing style tubes that reach into the beaker. Thank you. EdJ. 10 JAN 2020: MB reassigned to Chau to get a mechanical resource assigned. | Chau Nguyen | Edward Jackson | Chau to prioritize to provide mechancial design for Ed. | Taking care of higher priority projects and will get to it. OFI till new design. |  |  | This is at OFI status and pending new BSM test. | This is at OFI status and pending new BSM test. | Going to remove from list no need to track?
193460 | TE | N/A | 2019-07-29 | TE82475 | Luo HuaZhen | OPEN | Based on tester utilization calculation, this pressure decay tester in Flex has been used up 91.56% of total tester capacity. So we plan to buy another TE82475 tester by raising CEA. Please kindly advise the total price and estimated complete date to build this tester. Thank you Tester utilization data, FYI Tester Utilization based on Q1 & Q2’20 numbers: Prime is 79.3% Prime + Spares is 91.56% Q3 & Q4’19 numbers: Prime is 78.2% Prime + Spares is 91.5% | 26 Feb 2020: Redesign under way March 30, 2020: New tester in design proof stage now APRIL 30, 2020: In progress | Mark Jussaume |  | HuaZhen will meet Edwin to review and provide CEA to Milford when approved. | View updates from Matt |  | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req.
193551 | TE | 735002011 | 2019-09-02 | TE83032 | Luo HuaZhen | OPEN | The existing optics bench alignment stations are using a Golden flow cell to align all the optics benches. The flow cell used to align 5029 FLR optics benches had been using since year 2013 product transferred period and it had never been replaced. Venture observed the white colour residue on the glass window of the flow cell and it can not be removed although it was purging/priming through with MeOH for long hours. Refer to attached slide. Venture should have one back up golden flow cell in any case of flow cell degrading over a few years of usage and also for not interrupting production line while it waiting Milford to send over new Golden flow cell. | 30 APRIL 2020: 5 x optic bench sent. Pending Matt Franklin to use on testing. | Matt Franklin |  | Edwin is aware of the concern and issue from Quality. HuaZhen will follow up with Edwin to support Matt to get it done. | Matt Franklin is working on it and will provide to keep provide progress. |  | In progress of tester set up. | In progress of tester set up. | In progress of tester set up. | In progress of tester set up.
193552 | TE | 73500615 | 2019-09-02 | TE82760 | Luo HuaZhen | OPEN | On the month of Jul, Venture production found some Ram-Ray Valley Energy failures during FLR detectors final test for 5029 models. The UUTs were passed when they swapped with new flow cells. The failed flow cells were used in another UUTs and passed the test. Refer to attached slides for failure details. Raised OESR to follow up and get some advices from WPPL/Milford Engineering to look into the possible root causes of this failure. |  | Matt Franklin |  | HuaZhen will follow up with CM and confirm to close. | HuaZhen will provide information from CM to close it out. | Info was provided. Closed | Monitor and Mark F/WPPL working on this OESR |  | Same as 193551 | Same as 193551
193757 | TE | N/A | 2019-11-08 | TE83021 | Luo HuaZhen | OFI | Venture reported this fixture shows error message ' Serial CMD XCVR. The operation time out'. Suspect the CPU board of this fixture is faulty. Please kindly send us a right firmware for this CPU board for this repair or send us a CPU with firmware. Right now 2 pcs of PCB are pending for this test. | Firmware sent 13-NOV-2019, Matt B March 30, 2020: Performing FCT in instrument via TMD while fixture is repaired, refurbed and prepped for conversion to GTS II | Matt Brule |  | On-going Modify TSP to reflect instrument testing. | View updates from Matt | Matt: On-going Modify TSP to reflect instrument testing. | Revise TSP by Matt | In progress Going to ECR to get out from under TMD while repairs are completed, done this month | In progress Going to ECR to get out from under TMD while repairs are completed, done this month | 9238589 completed setting to OFI to track repair of fixture, not high priority as intrument test is sufficient.
203852 | TE | N/A | 2020-01-10 | TE83044 | Luo HuaZhen | OPEN | Flex Malaysia request to have a spare CPU with firmware for the following tester fixtures. Most of these CPU are obsoleted. Please kindly advise a spare solution. Can you please send us a spare if you have same part number? Or please advise if can use 520 CPU2K board with correct firmware version from you? Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 pls refer to page 2 attachment 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 | March 30, 2020: Collecting all necessary PCBs and firmware APRIL 30, 2020: Get information on GTS with individual CPU under teststand. Install the fixture with CPU - Huazhen | Paul Deschenes | Tony Fontes | Waiting for updates from Paul | View updates from Matt. (Tony was no longer with Waters) | Install actual fixture and update Matt the date of each fixture. - HZ | Matt will load firmware and ship it out after shutdown | In pregross Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SGFirmware Part numbers in OESR for 3, need to look at remaining two
204142 | TE | WAT025373 | 2020-05-27 | TE82475 | Luo HuaZhen | OPEN | Flex reports TE82475-2 Flow Cell Decay Tester is not able to test 2x Flow Cell simultaneously. If run both UUT at the same time, test will stop and both UUT will FAIL, If run 1x UUT at the time, both ports are OK. Currently Flex is using 1x UUT per run to support production needs. It may impact revenue if there's surge of order. So please kindly review and advise a way to make testers to run 2x UUT simultaneously or expedite the new Flow Cell Decay Tester build under CEA WPPL2002. |  | Paul Deschenes | [5 empty] | Same as 193460, flow cell decay tester | Same as 193460, flow cell decay tester | Same as 193460, flow cell decay tester Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req.
204144 | TE | N/A | 27—May—20 | TE84004 | Luo HuaZhen | Closed | Flex Malaysia feedback their ICT fixture TE084004- 2 for 2690 Fluidics Driver board 210000416 cannot capture some failures, eg. Needle malfunction failure due to U13 & U7 cannot be detected. Please refer to document attached. They request to add-in detectability/coverage on ICT test of locations on U1,U2,U3,U4,U5,U6,U7, U12,U13,14 and U16 to improve ICT testing. Please kindly advise if this improvement can be done and let team here know if any action is required from Flex Malaysia. |  | Matt Brule | [5 empty] | Need update from Paul? Blieve Paul OK' PetraCarbon? | Need update from Paul? Blieve Paul OK' PetraCarbon? | Paul OK'd working with PetraCarbon to improve if possible, closed.
204156 | TE | 289001405 | 2020-06-17 | TE82973 | Luo HuaZhen | Closed | Flex 2998 Flow Cell not able to pass Noise Test on 2998 Detector (SN:G08998943N) under TE82973-1. Team tried to repair but it’s required to have longer time to get it recovered. Due to urgent shipment, team assessed the risk and propose to use a known good 2998 detector (SN:F20998147A) as a replacement unit to continue production and shipment. This tester is used at subassembly for flow cell test, there’s PDA final test station to perform the noise test as 2nd gate too. Please kindly review and approve this proposed alternate solution (Using a same model passed 186002998 detector as replacement unit to continue testing). |  | Eric Turner | [5 empty] | Suspect this has been done? | Suspect this has been done? Close this one. | Closed, all set.
204249 | TE | 735000589 | 27—Aug—20 | N/A | Luo HuaZhen | Open | TSP 735000589 Rev F, page 7 of 7, steps 4 to 6 are repetitive in 735000615. Venture's practice is following TSP 735000615 for these steps as per training in Wexford during year 2013. As such we would like to remove these steps from this TSP 735000589. Please refer to redline TSP on page 2 |  | Adam Townsend | Eric Turner |  |  |  |  | Change cannot be made right now as an ECR and validation are in process | Change cannot be made right now as an ECR and validation are in process. Can we do this now? | Change cannot be made right now as an ECR and validation are in process. Will be done with ARC Premier update
204253 | TE | 735005108 | 27—Aug—20 | N/A | Luo HuaZhen | Open | We have complete 1st Premier TUV test in Flex as per TSP 735005108. During the run, the following items were redlined by production. Please kindly review and update if necessary. Thanks 1. On page 3, specified file name should be stated in TSP to give a clear instruction for operator to select. 2. On page 3, shunt with part number should be stated in TSP for operator to select. 3. On page 5. Leak Detect Test result should be added into Test Record. The redlined TSP is attached on page 2 |  | Mark Jussaume | [5 empty] | May be done, need to check in SAP | May be done, need to check in SAP. Now ? | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates.
204270 | TE | WAT210700 | 2020-09-07 | TE82526 | Luo HuaZhen | Open | TE82526-1 Pump Tray Tester keeps on restarting when running test. The test box may power on/off intermittently. Prior to reset\restart, the tester is showing “LOW SHOP AIR” although the shop air pressure is with 100 PSI. So, the suspected component as this are related to incoming pressure (Regulator Model: R884-01C and reducing valve as attached.) Please kindly review the attached video in the link below and advise next action or send us a suspected part for this repair. Thanks https://files.waters.com/?ShareToken=03BF55399289EC172AD2134EFB663E2CB434D927 Team here has done the following troubleshooting so far, but still have restarting intermittently 1. Replaced power supply p/n: WAT270471 2. Replaced battery CR2032 on CPU board 3. Unplug and plug-in all cables and connectors 4. Release all pressures from tubing and air switch and retightening tubing 5. Replaced/restored CPU PCB p/n:056370 Impact: may impact revenue if not able to recover in time |  | Bill Bates | [5 empty] | In Process | In Process ark and HuaZhen working on. Need Diade number Mark.... | In Process Mark and HuaZhen working on. Need Diode number Mark. Asked Mark for update
204280 | TE | N/A | 2020-09-14 | TE82988 | Luo HuaZhen | CLOSED | Venture Singapore GTX #3 Tester Failed Multiple Tests on 210000478 PCBA Functional Tests Troubleshooting performed, 1) Try with another 2 boards, all the three 210000478 PCBA exhibited the same failure. 2) Change to Self-Test/Diagnostic Fixture, it shows Pickering Relay Mux (40-612-002) failure (see attached). The resistance reading on over 1.5 ohm spec. Please kindly help on remote diagnostic and advise next action to fix this issue if it’s required from Venture/WPPL. Thanks The tester is down now and affected Part No: 210000478 / 210000397 / 210000400 / 210000969 Immediate Shipment Impact: 52 units Estimated Date: 14 September 2020 1) 700005431, Qty : 20, Shipment Date : 14-Sep-20 2) 186015043, Qty : 32, Shipment Date : 16-Sep-20 |  | Matt Brule | [5 empty] | Pretty sure this can be closed and a different one oponed to pursue self-test fixture issue?? | Pretty sure this can be closed and a different one oponed to pursue self-test fixture issue?? | Closed problem fixed
204320 | TE | 735003678 | 2020-09-30 | 735003678 | Luo HuaZhen | Closed | During performing PC maintenance found at "Step 1.7 Check for unsent test files in all the 3 testers below. Just to inform that there were number of files in the folder Directory C: \program files\sigmaQuest\sigmaProbe\sigmaQueue\Sendbuffer. Testers name as shown below: TE82865-5 CPSO Final Test Station TE82865-6 CPSO Final Test Station TE82865-2 ASO Final Test Station |  | Matt Brule | [5 empty] | More data to move, Next week | More data to move, Next week | Duplicate file moved to backup
204395 | ME |  |  |  |  | Closed | ECO9235494 calls for change to replace 210000520 with 210001232 for Assy, Detector 2998 Classic (Top Level 186002998) However, the Genealogy record in Sigmasure for 186002998 is not updated. It is still showing 210000520 instead of 210001232. Please make the changes in the Genealogy record for 186002998 and replace 210000520 with 210001232. | I check Sigma Sure and the assembly process steps appear to good with the correct PCB number for the build option but the top level product view has the old and new PCB number shown. The assembly process step has the correct part number in it. The top level product will have both part numbers in it, but as long as the old part number is not checked (and it is not checked) then we are good. | Roland Poisson | [7 empty] | Closed no action, looks OK
204443 | TE |  |  |  |  | Open | ICT Fixture TE83053 in CEI Batam failed test at U1 programming Since switching to new approved IC IS25LP032D from ISSI. Fixture is working fine with old IC FL032PIF from Micro Spansion, but this old IC FL032PIF is EOL and CEI has zero stock of this old IC. 54 pcs of this PCB were affected by this failure. CEI cannot pass test at U1 with new IC IS25LP032D from ISSI. Local ICT fixture vendor Patracarbon has accessed the ICT tester remotely and recommend working with ISSI vendor to update ICT programming code. Please kindly provide support CEI Batam to fix this issue. | ICT programming issue | Paul Deschenes | Luo Hua Zhen | [6 empty] | Active
204449 | ME |  |  |  |  | Open | A 2695 Separations Module was returned from a customer site (day 1 failure) for a vial presence sensor failure. The root cause was identified as black RTV on the end of an optical tube, the upper (receiver) tube. Black RTV is used in the assembly process, to hold the optical tube in the sample compartment. The fact that RTV found its way to the end of the tube represents a workmanship or general process issue. Bill Wildman will notify the CM through our WPPL office, with the goal of auditing the line, seeking opportunities for improvement regarding RTV handling. There is a chance that an assembly process night require update, perhaps an additional inspection step (to check for stray RTV). This request should serve as: 1. A notification of the finding (PIM or WPPL may reach out to Mfg Eng, so this is an advance notice). 2. A request to screen the assembly procedure(s) to see if we handle the black RTV with sufficient care to prevent similar issues in the future. Ref: CIT USAABUMKLR22 System Readiness Report 721021649 Image on Pg. 02 of this request | This is likely more of a Milford issue with some assistance from WPPL | Alex Beaudoin
204457 | TE |  |  |  |  | Open | TE82951 Mask Alignment System not able to perform oper 3.2 onwards per ASP due to the Stepping Motor Control II for the camera is not working as power supply for this Stepping Motor Control II is faulty. Below is power supply specification of this faulty power supply adaptor: Vendor: GlobTek, Inc. Part No: TR9CH100N05-Y Model: GT-21089-1818-T3 Input: 100-240 V~ 0.5A MAX, 50-60 Hz Output: +18Vdc 1.0A Output Power: 18W Max Local team searched for a replacement. But it’s difficult to find an exact same part number: TR9CH100N05-Y online. So please kindly advise if Milford has the same power supply adaptor for replacement, or team should search for another equivalent power adaptor for Stepping Motor Control II for the camera and advise if there’ s any concern if using equivalent power adaptor. Please refer to page 2 & 3 Stepping Motor Control II and its power supply adaptor (faulty one). | Need Power supply for camera | Bill Bates | [7 empty] | Active