## Sheet1
Log Number | Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Problem Summary | Summary | Assignee 1 | Assignee 2 | Minutes on Feb 27, 2019 | Minutes on March 30, 2020 and Apr 30, 2020 | WPPL update on 04 May 2020 | 25 Jun 2020 review udpate | October 1 2020 update | October 28 2020 Update | November 23, 2020 Update | December Update, 2020 Update | May Update | Column1 | Column2
161186 | TE | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. | Open since 2017, June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. March 30, 2020: Matt will dig in tomorrow, and understand what is needed, and we will see if we can get someone on it. APR 30, 2020: To advise a date - Huazhen | Bill Bates | Matt Brule | Chau to confirm with Eric if Matt and his group are working on it. | View updates from Matt | Checked with Flex, can put this acitivity KIV till 30 Jun | Matt/Stephen will continue after shutdown; add TE No/Tool. HZ check with Flex and send it | Matt: Q4 2020 estimate completion date | Working on ordering parts with China Testers. Reasign…. |  |  | Need to understand exact needs, just control boxes?
171694 | TE | NA | 2017-04-27 | N/A | Yeap Siew Yee | Open | Files pending in the 'sendbuffer' folders yet to be sent out. 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | Open Since 2017, 27-FEB-2019 Hope to get time to get this going in 2019 May 2019: Still at same status, second on project priority list March 30, 2020: HTCH updates looking like 2nd quarter completion APRIL 30, 2020: End Jun to complete the TE83026 upgrade. | Matt Brule |  | On-going | View updates from Matt |  | Target to complet it by Q3 2020 | Primier products added; TE0083013 will be shipping to Venture; target to complete upgrade by Q4 2020; | Will Clean out today or tomorrow | Pending, keeping open to track updates to TE83013 and TE83026 | Pending, keeping open to track updates to TE83013 and TE83026 | Pending, keeping open to track updates to TE83013 and TE83026
193460 | TE | N/A | 2019-07-29 | TE82475 | Luo HuaZhen | OPEN | Based on tester utilization calculation, this pressure decay tester in Flex has been used up 91.56% of total tester capacity. So we plan to buy another TE82475 tester by raising CEA. Please kindly advise the total price and estimated complete date to build this tester. Thank you Tester utilization data, FYI Tester Utilization based on Q1 & Q2’20 numbers: Prime is 79.3% Prime + Spares is 91.56% Q3 & Q4’19 numbers: Prime is 78.2% Prime + Spares is 91.5% | 26 Feb 2020: Redesign under way March 30, 2020: New tester in design proof stage now APRIL 30, 2020: In progress | TBD |  | HuaZhen will meet Edwin to review and provide CEA to Milford when approved. | View updates from Matt |  | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | Testers in-house in Milford, config to begin as soon as time allows.
193552 | TE |  |  |  | Lou HuaZhen | Open | On the month of Jul, Venture production found some Ram-Ray Valley Energy failures during FLR detectors final test for 5029 models. The UUTs were passed when they swapped with new flow cells. The failed flow cells were used in another UUTs and passed the test. Refer to attached slides for failure details. Raised OESR to follow up and get some advices from WPPL/Milford Engineering to look into the possible root causes of this failure. On the month of Jul, Venture production found some Ram-Ray Valley Energy failures during FLR detectors final test for 5029 models. |  | Matt Franklin | [9 empty] | Added this to May list, had been missing… -A new alignment tool was unsuccessful -Firmware update to correct, but is not coming soon. -Eric / Matt F are looking into adding new firmware to alignment station.
203852 | TE | N/A | 2020-01-10 | TE83044 | Luo HuaZhen | OPEN | Flex Malaysia request to have a spare CPU with firmware for the following tester fixtures. Most of these CPU are obsoleted. Please kindly advise a spare solution. Can you please send us a spare if you have same part number? Or please advise if can use 520 CPU2K board with correct firmware version from you? Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 pls refer to page 2 attachment 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 | March 30, 2020: Collecting all necessary PCBs and firmware APRIL 30, 2020: Get information on GTS with individual CPU under teststand. Install the fixture with CPU - Huazhen | Paul Deschenes |  | Waiting for updates from Paul | View updates from Matt. (Tony was no longer with Waters) | Install actual fixture and update Matt the date of each fixture. - HZ | Matt will load firmware and ship it out after shutdown | In pregross Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SGFirmware Part numbers in OESR for 3, need to look at remaining two Matt to email programming instruction and look up remaining PCBs | In progress Matt to collect and email firmware, such that boards can be programmed in SGFirmware Part numbers in OESR for 3, need to look at remaining two Matt to email programming instruction and look up remaining PCBs | Matt Owes instruction to HuaZhen, and Weimen
204253 | TE | 735005108 | 27—Aug—20 | N/A | Luo HuaZhen | Open | We have complete 1st Premier TUV test in Flex as per TSP 735005108. During the run, the following items were redlined by production. Please kindly review and update if necessary. Thanks 1. On page 3, specified file name should be stated in TSP to give a clear instruction for operator to select. 2. On page 3, shunt with part number should be stated in TSP for operator to select. 3. On page 5. Leak Detect Test result should be added into Test Record. The redlined TSP is attached on page 2 |  | Matt Brule | [5 empty] | May be done, need to check in SAP | May be done, need to check in SAP. Now ? | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. Mark did not completre Matt to pick up
204371 | ME |  |  |  | Lou HuaZhen | Open | Flex report their shunt 289002227 for ASP 279003847 step 8.16 is worn out. Request to have a replacement for this shunt. Also request additional shunt 289002226 to meet production needs. Please kindly advise the build cost for shunt 289002227 and 289002226. So, we can raise a capex request by this week. Thanks |  | Alex Beaudoin | Matt Franklin | [8 empty] | Matt Franklin actively trying to aquire shunts, steve to follow up, looking to close CAPEX Parts are available, Alex to cut orders
204443 | TE |  |  |  | Lou HuaZhen | Open | ICT Fixture TE83053 in CEI Batam failed test at U1 programming Since switching to new approved IC IS25LP032D from ISSI. Fixture is working fine with old IC FL032PIF from Micro Spansion, but this old IC FL032PIF is EOL and CEI has zero stock of this old IC. 54 pcs of this PCB were affected by this failure. CEI cannot pass test at U1 with new IC IS25LP032D from ISSI. Local ICT fixture vendor Patracarbon has accessed the ICT tester remotely and recommend working with ISSI vendor to update ICT programming code. Please kindly provide support CEI Batam to fix this issue. | ICT programming issue | Paul Deschenes | Luo Hua Zhen | [6 empty] | Active Working programming issue with fixture vendor | Active Working programming issue with fixture vendor | Active Working programming issue with fixture vendor, In process with US ICT supplier.
204486 | TE |  |  |  | Desmond Ann | Open | Sigmasure Data issue |  | Matt Brule | [9 empty] | Sigmasure issue,Nneed TE number
214621 | TE |  |  |  | Yeehwa Chan | Open | Requesting a 4th duplicate tester for TE0084054-x (to be named -7) to increase test capacity in Singapore. |  | Bill Bates | [9 empty] | Get update from Bill
214708 | TE |  |  |  |  | Closed | Please be informed that tested data at below two testers in Feb cannot be pushed out to SigmaSure during PC PM. Data can be seen in SigmaQueue folder of their respective testers. 1)TE83026-1 (IP:10.249.40.27) 2)TE83064-11 (IP:10.249.40.31) Please kindly help to check and clear these files under SigmaQueue buffer. |  | Matt Brule
214746 | TE |  |  |  |  | Open | Sigmasure errors |  | Matt Brule | [9 empty] | Will get to soon
214760 | TE |  |  |  | Luo HuaZhen | Closed | TE82800 PC replacement |  | Luo HuaZhen | [9 empty] | Done
214881 | TE |  |  |  | Desmond Ann | Closed | 1. Problem Statement. - PDA Linearity Test Station not able to perform Step 4 Wavelength Accuracy Test due cannot bring into Review. 2. Acceptance criteria that is not met. (e.g. relevant specification number and revision) - UUT shall pass complete test as per TSP 735002531. 3. Relevant records (e.g. Incomin QA, Test Records, Photos, trackise NCRs numbers and related issue) - Refer to attached pictures. 4. Info on whether this issue impacts revenue. - may impact 33 units of 10X of 186015034 and 11x of 186015053 on 14 May 2021 (rework ) if not able to recover tester in time. |  | Eric Turner
214894 | TE |  |  |  | Andy Sawyer | Open | locking up in diagonics has happen three times |  | Adam Townsend | [9 empty] | Trying to include 5 second delay