## Sheet1
Log Number | Group | ASP TSP No | Date | TE Tool No | Originator | STATUS | Problem Summary | Summary | Assignee 1 | Assignee 2 | Minutes on Feb 27, 2019 | Minutes on March 30, 2020 and Apr 30, 2020 | WPPL update on 04 May 2020 | 25 Jun 2020 review udpate | October 1 2020 update | October 28 2020 Update | November 23, 2020 Update | December Update, 2020 Update | FEB Update
161186 | TE | N/A | 2016-10-26 | N/A | Desmond Ann | OPEN | "CEA 16009" is to build the Universal Coarse and Fine Align Station for PDA optics benches. | Open since 2017, June 03, 2019: CEA# WPPL19002 for 2 control box. Pending Milford to Fabricate box. ASK ERIC FOR COMPLETION DATE. March 30, 2020: Matt will dig in tomorrow, and understand what is needed, and we will see if we can get someone on it. APR 30, 2020: To advise a date - Huazhen | Bill Bates | Matt Brule | Chau to confirm with Eric if Matt and his group are working on it. | View updates from Matt | Checked with Flex, can put this acitivity KIV till 30 Jun | Matt/Stephen will continue after shutdown; add TE No/Tool. HZ check with Flex and send it | Matt: Q4 2020 estimate completion date | Working on ordering parts with China Testers. Reasign…. |  |  | To be built at same time as China systems (1st Quarter)
171657 | TE | 735000445 | 2017-04-20 | TE82723 | Edwin Koh | OFI | During product audit, it found that the operator cannot execute to TSP procedure 735000445 on per section 3.0 without completing additional step of Water DHCP download. Understand this additional step is necessary previously due to PC failure and it a temporary fix to access the program for testing. Action options 1. If arrangement is permanent, can TSP procedure 735000445 to be revised 2. Can Test Engineer update the TE82723 to link this additional step back into the standard program. Preferred option 2 if possible. Attached is the supporting information on the additional step and document mark-up. | June 28, 2017: This is a TSP change anybody can do, but will need spec alignment - low volume, low priority. If categorize it as an OFI, an timeline for completion should be provided. March 30, 2020: Switched from Eric to Matt, hoping to get some traction on this this month, while folks are working @ home | Matt Brule |  | HuaZhen will confirm with Venture and inform everyone involved. | OFI after WPPL verified with CM | OFI, Verified with CM |  |  | TSP update next week | Still pending | Still pending | Edwin to look into from WPPL, Matt to look into from Milford. Tried to connec to the tester but couldn't. Need to connect to see if I can correc the issue. OR need markup of requested change.
171694 | TE | NA | 2017-04-27 | N/A | Yeap Siew Yee | Open | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | Open Since 2017, 27-FEB-2019 Hope to get time to get this going in 2019 May 2019: Still at same status, second on project priority list March 30, 2020: HTCH updates looking like 2nd quarter completion APRIL 30, 2020: End Jun to complete the TE83026 upgrade. | Matt Brule |  | On-going | View updates from Matt |  | Target to complet it by Q3 2020 | Primier products added; TE0083013 will be shipping to Venture; target to complete upgrade by Q4 2020; | Will Clean out today or tomorrow | Pending, keeping open to track updates to TE83013 and TE83026 | Pending, keeping open to track updates to TE83013 and TE83026 | Pending, keeping open to track updates to TE83013 and TE83026
193413 | TE | 735000615 | 2019-06-24 | TE82760 | Luo HuaZhen | OPEN | TSP735000615 Rev P, page 8 of 9, the TSP is missing a step during keypad test after passing the functional test . For 2475 mode, we need to set the date for the lamp installed. Press “DIAG”, “7” ,” 1” as per red line TSP attached at page 2 of this OESR Please review this step and update the TSP if necessary. Thanks | July 02, 2019: WPPL : Pending for Milford to update TSP. 26 Feb 2020: Reassigned to Matt per Eric. March 30, 2020: Hoping to get some traction on this this month, while folks are working @ home. | Matt Brule | Eric Turner | Matt Brule has accepted the reassignment on Feb 26, 2020. | View updates from Matt |  | In progess | Thought this had been reassigned, it had not will complete next week | Thought this had been reassigned, it had not will complete next week | Still Pending | Still Pending | Eric to complete with ARC Premier updates this quarter, waiting to close after validation of Premier updates.
193460 | TE | N/A | 2019-07-29 | TE82475 | Luo HuaZhen | OPEN | Based on tester utilization calculation, this pressure decay tester in Flex has been used up 91.56% of total tester capacity. So we plan to buy another TE82475 tester by raising CEA. Please kindly advise the total price and estimated complete date to build this tester. Thank you Tester utilization data, FYI Tester Utilization based on Q1 & Q2’20 numbers: Prime is 79.3% Prime + Spares is 91.56% Q3 & Q4’19 numbers: Prime is 78.2% Prime + Spares is 91.5% | 26 Feb 2020: Redesign under way March 30, 2020: New tester in design proof stage now APRIL 30, 2020: In progress | Mark Jussaume |  | HuaZhen will meet Edwin to review and provide CEA to Milford when approved. | View updates from Matt |  | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | In progess, target to finish by end of this year. | Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | New testers ordered Jan delivery, then config verification etc.
193551 | TE | 735002011 | 2019-09-02 | TE83032 | Luo HuaZhen | OPEN | The existing optics bench alignment stations are using a Golden flow cell to align all the optics benches. The flow cell used to align 5029 FLR optics benches had been using since year 2013 product transferred period and it had never been replaced. Venture observed the white colour residue on the glass window of the flow cell and it can not be removed although it was purging/priming through with MeOH for long hours. Refer to attached slide. Venture should have one back up golden flow cell in any case of flow cell degrading over a few years of usage and also for not interrupting production line while it waiting Milford to send over new Golden flow cell. | 30 APRIL 2020: 5 x optic bench sent. Pending Matt Franklin to use on testing. | Matt Franklin |  | Edwin is aware of the concern and issue from Quality. HuaZhen will follow up with Edwin to support Matt to get it done. | Matt Franklin is working on it and will provide to keep provide progress. |  | In progress of tester set up. | In progress of tester set up. | In progress of tester set up. | In progress of tester set up. | In progress of tester set up. | Pilot testing complete. Working with R&D to understand next steps.
203852 | TE | N/A | 2020-01-10 | TE83044 | Luo HuaZhen | OPEN | Flex Malaysia request to have a spare CPU with firmware for the following tester fixtures. Most of these CPU are obsoleted. Please kindly advise a spare solution. Can you please send us a spare if you have same part number? Or please advise if can use 520 CPU2K board with correct firmware version from you? Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 pls refer to page 2 attachment 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 | March 30, 2020: Collecting all necessary PCBs and firmware APRIL 30, 2020: Get information on GTS with individual CPU under teststand. Install the fixture with CPU - Huazhen | Paul Deschenes | Tony Fontes | Waiting for updates from Paul | View updates from Matt. (Tony was no longer with Waters) | Install actual fixture and update Matt the date of each fixture. - HZ | Matt will load firmware and ship it out after shutdown | In pregross Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SG | In progress Matt to collect and email firmware, such that boards can be programmed in SGFirmware Part numbers in OESR for 3, need to look at remaining two Matt to email programming instruction and look up remaining PCBs | In progress Matt to collect and email firmware, such that boards can be programmed in SGFirmware Part numbers in OESR for 3, need to look at remaining two Matt to email programming instruction and look up remaining PCBs | Matt Owes instruction to HuaZhen, involve Weimen
204142 | TE | WAT025373 | 2020-05-27 | TE82475 | Luo HuaZhen | Closed | Flex reports TE82475-2 Flow Cell Decay Tester is not able to test 2x Flow Cell simultaneously. If run both UUT at the same time, test will stop and both UUT will FAIL, If run 1x UUT at the time, both ports are OK. Currently Flex is using 1x UUT per run to support production needs. It may impact revenue if there's surge of order. So please kindly review and advise a way to make testers to run 2x UUT simultaneously or expedite the new Flow Cell Decay Tester build under CEA WPPL2002. |  | Paul Deschenes | [5 empty] | Same as 193460, flow cell decay tester | Same as 193460, flow cell decay tester | Same as 193460, flow cell decay tester Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | Same as 193460, flow cell decay tester Purchasing off-the-shelf solution, looking into cal / PM and CEA before writing req. | Combined with 193460
204249 | TE | 735000589 | 27—Aug—20 | N/A | Luo HuaZhen | Open | TSP 735000589 Rev F, page 7 of 7, steps 4 to 6 are repetitive in 735000615. Venture's practice is following TSP 735000615 for these steps as per training in Wexford during year 2013. As such we would like to remove these steps from this TSP 735000589. Please refer to redline TSP on page 2 |  | Adam Townsend | Eric Turner |  |  |  |  | Change cannot be made right now as an ECR and validation are in process | Change cannot be made right now as an ECR and validation are in process. Can we do this now? | Change cannot be made right now as an ECR and validation are in process. Will be done with ARC Premier update | Change cannot be made right now as an ECR and validation are in process. Will be done with ARC Premier update | Eric to complete with ARC Premier updates this quarter, waiting to close after validation of Premier updates.
204253 | TE | 735005108 | 27—Aug—20 | N/A | Luo HuaZhen | Open | We have complete 1st Premier TUV test in Flex as per TSP 735005108. During the run, the following items were redlined by production. Please kindly review and update if necessary. Thanks 1. On page 3, specified file name should be stated in TSP to give a clear instruction for operator to select. 2. On page 3, shunt with part number should be stated in TSP for operator to select. 3. On page 5. Leak Detect Test result should be added into Test Record. The redlined TSP is attached on page 2 |  | Mark Jussaume | [5 empty] | May be done, need to check in SAP | May be done, need to check in SAP. Now ? | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. | Reassigned to Mark Jussaume, will be completed as part of the impending ARC Premier updates. Waiting for update from Mark
204258 | ME |  |  |  |  | Closed | The optics bench alignment stations are using a Golden flow cell to align all the optics benches. The current golden flow cell is more than 5 years of age and recently, Venture observed that the backpressure of golden flow cell is >100 psi (spec is max 145psi) There is a constant water flowing through inlet to outlet, no physical leak can be seen, only pressure is high. The 515 pump is ok, verified with other production flow cells, less than 50psi and stable. Tried out the following mehtods but it deoes not solved the issue: - Flushing with MeOH, Acetone and N2 for a few hours. - Flushing in the opposite direction Request to send over a new Golden Flow cell for Optics Alignment Test. Refer to attachement for the flow cell photos. The optics bench alignment stations are using a Golden flow cell to align all the optics benches. The current golden flow cell is more than 5 years of age and recently, Venture observed that the backpressure of golden flow cell is >100 psi (spec is max 145psi) There is a constant water flowing through inlet to outlet, no physical leak can be seen, only pressure is high. The 515 pump is ok, verified with other production flow cells, less than 50psi and stable. Tried out the following mehtods but it deoes not solved the issue: - Flushing with MeOH, Acetone and N2 for a few hours. - Flushing in the opposite direction Request to send over a new Golden Flow cell for Optics Alignment Test. Refer to attachement for the flow cell photos. |  | Matt Franklin | [9 empty] | Closed
204337 | TE |  |  |  |  | Open | 210000526 SFC Personality FCT Fixture [TE83045-2] starts to have vacuum leakage which result in contact issue between pogo pin and test point. This result in UUT always fail at initialization or 5V power up fail. Temporary fixed by adding Kapton tape on the hole location K2, K1, J53, j60,J61 and Fixture area covered by yellow insulation tape, pls refer to attached picture. Please kindly advise a way to fix this vacuum leak issue or should send the fixture back Milford for repair. Thanks |  | Lou Hua Zhen | [9 empty] | HuaZhen working with CEI, no UUTs to test. Temporary solution in place. HuaZhen to review with CEI, and see if it can be fixed.
204371 | ME |  |  |  |  | Open | Flex report their shunt 289002227 for ASP 279003847 step 8.16 is worn out. Request to have a replacement for this shunt. Also request additional shunt 289002226 to meet production needs. Please kindly advise the build cost for shunt 289002227 and 289002226. So, we can raise a capex request by this week. Thanks |  | Alex Beaudoin | Matt Franklin | [8 empty] | Working to find documentation to build. Matt Franklin working on suppling obsolete parts. Take an action to discuss Matt Franklin long term solution.
204443 | TE |  |  |  |  | Open | ICT Fixture TE83053 in CEI Batam failed test at U1 programming Since switching to new approved IC IS25LP032D from ISSI. Fixture is working fine with old IC FL032PIF from Micro Spansion, but this old IC FL032PIF is EOL and CEI has zero stock of this old IC. 54 pcs of this PCB were affected by this failure. CEI cannot pass test at U1 with new IC IS25LP032D from ISSI. Local ICT fixture vendor Patracarbon has accessed the ICT tester remotely and recommend working with ISSI vendor to update ICT programming code. Please kindly provide support CEI Batam to fix this issue. | ICT programming issue | Paul Deschenes | Luo Hua Zhen | [6 empty] | Active Working programming issue with fixture vendor | Active Working programming issue with fixture vendor | Active Working programming issue with fixture vendor, In process with US ICT supplier
204449 | ME |  |  |  |  | Open | A 2695 Separations Module was returned from a customer site (day 1 failure) for a vial presence sensor failure. The root cause was identified as black RTV on the end of an optical tube, the upper (receiver) tube. Black RTV is used in the assembly process, to hold the optical tube in the sample compartment. The fact that RTV found its way to the end of the tube represents a workmanship or general process issue. Bill Wildman will notify the CM through our WPPL office, with the goal of auditing the line, seeking opportunities for improvement regarding RTV handling. There is a chance that an assembly process night require update, perhaps an additional inspection step (to check for stray RTV). This request should serve as: 1. A notification of the finding (PIM or WPPL may reach out to Mfg Eng, so this is an advance notice). 2. A request to screen the assembly procedure(s) to see if we handle the black RTV with sufficient care to prevent similar issues in the future. Ref: CIT USAABUMKLR22 System Readiness Report 721021649 Image on Pg. 02 of this request | This is likely more of a Milford issue with some assistance from WPPL | Alex Beaudoin | [9 empty] | OK to close being driven by quality? Per Alex ASP will be updated to inspect RTV.
204472 | TE |  |  |  |  | Open | Venture observed the following difference on test report/template recently. Please kindly help to check and update report template to fix the following reported issue. Please let us know if remote access to TE82760-7 with UUT is needed. 1) There is no test report print out on TE82760-7 if UUT failed test. TE82760-7 test software does not generate failed test report in "Output File" of test software directory although can see failed records in SimaSure; No physical report printed for failed units but able to print passed units for all the models. The other two final test stations can print out failed report. 2) Noticed missing information on passed test record. The passed report from TE82760-7 is missing firmware version for 186002475,186017006 and 186247500 and appears only one TMD number (not able to show more than one TMD number) Attachments on page 2 & 3 :1) eventlogforfailure.pdf and 2) Testreport_missinginfo.pdf |  | Eric Turner | [9 empty] | Eric to fix during ARC Premier update. Waiting for Premeir Validation
214548 | TE |  |  |  |  | Open | During PM for testers PC in Venture, they found there're files under sigmaquene and sendbuffer folders for testers PC in the attached list. Please kindly be informed and clear it if necessary. |  | Matt Brule | [9 empty] | Sigmasure records to copy, will get to soon.
204472 | ME |  |  |  |  | Open | ECO (9235495, 9235493, 9234970) request to change 210000520 to 210001232, Genealogy is required to be updated for the change per dates attached. |  | Roland Poisson | [9 empty] | Issues with Sigmasure preventing completion.
204573 | TE |  |  |  |  | Open | The APS power adaptor (attached photo) which is to supply 24 V to vent valve drive (solvent delivery system) is out of functioning (no output) and need to replace it. The valve is used on FLR final tester TE82760-5. Would you please send a replacement power supply for this valve? Or advise a part number or equivalent part number for us to order it locally? Please kindly refer to attached photo for this power supply adaptor. |  | Lou HuaZhen | [9 empty] | HuaZhen provided part number for complete spE1:T20are to order from ADC