
	Operations Engineering Service Request
	Log Number:      
To be filled in by Milford Operations Engineering


Please fill out the top section and email to Ops_Eng_Support@waters.com
Press F1 in any field for help (All fields are mandatory. Put N/A if not applicable)
	Date (DD-MMM-YY)
	13-Apr-18

	Originator
	Name:      
	Phone/Ext:     

	Department
	
	Zone/Section/Pole No.
	     

	TE / Tool No.
	82995 - 2
	Associated Part No.
	     

	TE / Tool Name
	Acquity CPSM RTZ Assembly tester
	Software Rev
	16

	Plant / Location
	Venture GES / Singapore
	ASP/TSP
	No.: 735002271
	Rev: G

	   Summary of problem(s) (as detailed as possible and include attachment if required):
   Known good UUT use to test on TE82995-2 show "Pierce Test Failed" see attached print screen capture. Replace the Control Board (P/N: 210000367) inside the test box, problem never resurface.


Milford Operations Engineering use only

	Category
	 FORMCHECKBOX 
  Repair
 FORMCHECKBOX 
  Upgrade

 FORMCHECKBOX 
  Duplicate

 FORMCHECKBOX 
  New

 FORMCHECKBOX 
  Other


	Assignee(s)
	     
	Date Received (DD-MMM-YY) 
	     

	Qualification Required
	 FORMCHECKBOX 
 Yes      FORMCHECKBOX 
 No (Please explain in Summary of action taken)

	Summary of action taken (as detailed as possible and include attachment if required):
      

	Date Completed (DD-MMM-YY)
	     

	Recommendation:



	


Form Revision: 02

