## Sheet2
Row Labels | Count of Log Number
1- High(< 1 week) | 23
2-Normal (<4 weeks) | 47
3- Low (> 4 weeks & above) | 15
(blank)
Grand Total | 85

## TE and ME - Use filters to view
Log Number | Engineering Group | Date | TE Tool No | Category | Problem Summary | Site | Originator | Assignee 1 (UK/ US) | Assignee 2 (WPPL) | Date Completed | Remark/Closure | Priority | STATUS | Priority | Site | (All)
1- High(< 1 week)
S180001 | Test Engineering | 2018-04-13 | TE82995-2 | Repair | Known good UUT use to test on TE82995-2 show "Pierce Test Failed" see attached print screen capture. | Venture | Tan Hong Cheng |  | Desmond | 2018-04-13 | Replace the Control Board (P/N: 210000367) inside the test box, problem never resurface. | 2-Normal (<4 weeks) | CLOSED | 2-Normal (<4 weeks) | Count of Assignee 2 (WPPL) | Column Labels
S180002 | Manufacturing Engineering | 2018-04-20 | 999119657 | Repair | Shuttle Press with Punch Station cannot power on. Open casing found Y3 alarm on. Perform check on all wire connection no loose connection found. Use multi-meter to measure the 24 VDC Output connector found no output. Suspect power supply is faulty or other issue cause Y3 alarm on. | Venture | Tan Hong Cheng |  | Desmond | 2018-04-20 | Venture found an alternate 24 PWS and fixed. 2 spare oringal part number of power supplies received for Milford for future use | 2-Normal (<4 weeks) | CLOSED | 3- Low (> 4 weeks & above) | Row Labels | Open | Grand Total
S180003 | Manufacturing Engineering | 04.26.2018 | QDa Pumping Block Removal Tool | New | When the UUT test failed, it is nessessary to verify with Gold Brick parts fo find the failure. Because of the vacuum sucking pressure during the UUT is operating, the pumping block is always difficult to remove from the analyzer housing. | Venture | Alan Oh |  | Mandy | 2018-07-09 | [09Jul18] Venture have received this removal tool - updated by Mandy | 2-Normal (<4 weeks) | CLOSED |  | Edwin | 1 | 1
S180004 | Test Engineering | 2018-06-27 | TE82943 | Others | Missing serial number at the bottom of the test report (Ser.No) of Acquity BSM 18K PLUS test in some Test record and it is encountered across all test stations TE82943 #1,#19,#20,#21,#22 and #24. Please kindly refer to the "BSM Test Record.JPG" | Venture | Chan Myae Soe | Milford - Ed Jackson | Hua Zhen |  | OESR 182626 [30Aug18] Ed Reply: No updates to report. Expect to look into this anomaly soon. It's not repeatable and completely random from what I've discovered. [19Oct18] Ed update: Made a Test Program code modification on the TE Lab PC and is under evaluation. | 3- Low (> 4 weeks & above) | Open | Status | Hua Zhen | 5 | 5
S180005 | Test Engineering | 06.25.2018 | TE83064-All | Upgrade | Starting from end of April (mainly after software & firmware updates) till now, we are observing high Needle Seal Force failures for all FTN models at all testers. Please refers to attached ppt files for information that we observed about this failure. The purpose of this OESR is to request help from Waters' Test engineer for Needle Seal Force failure and to prequest direction to prevent/reduce this failure. | Venture | Eilwin | Milford-Rick | Hua Zhen |  | OESR 182633 [06Jul18] GES being tasked to measure spring force. [17Jul18] GES have verified 10 springs and confirmed that the spring force are all within specs. To share results with Rick and ask for further advise. [27Jul18] Venture to provide the test results for the 10pcs spring under evaluation. [10Aug18] Rick questioned Load Cell replacement for repair, Venture was informed to provide explaination. Changing spec from 48 to 40 is in discussion in Milford [14Aug18] No load cell(seal wash block) was replaced but re-assemble of load cell was performed. Venture agreed to focus on those re-assemble required units to seek improvement of yeild. Venture will update team with test data on Thursday. [21Aug18] Venture look into the units failed and required reassemble the seal wash block data and will contact with Milford to pursuit a solution [28Aug18] Test Data and Venture Spring Study sent to Rick, no reply so far [31Aug18] Rick advise to have evaluation on Spring Cup with 90 degree interval, Venture was informed [20sep18] The Spring Cup evaluation data sent to Rick and suggested by Edwin to have Venture to get higher spec of Spring from supplier to do evaluation [12Oct18] Rich advised to study the relation between misalignment of wash cup and needle seal force failure issue | 3- Low (> 4 weeks & above) | Open | Open | Mandy | 1 | 1
S180006 | Manufacturing Engineering | 06.27.2018 | N/A | Others | On 13 Jun 2018, Part Number: 186015085. Instrument serial number F18FTP669G barcode label sheet generate by GES system is incorrect. Production operators failed to detect it and scan into Genealogy System. When this error was detected later, production operators try to correct the error and re-entry the correct data result in multiple entry was enter. | Venture | Tan Hong Cheng | Milford-Roland | Desmond | 2018-06-27 | [27Jun18] Data entered into Sigmasure cannot be adjusted. Amend on hardcopy of Sigmasure data. | 2-Normal (<4 weeks) | CLOSED | Closed | Yee Hwa | 2 | 2
S180007 | Test Engineering | 05.29.2018 | TE82736 | Others | Milford requests to have test data for Alliance Display, Degasser and Heater/Cooler test in the past 2 years. | Milford | Milford-Mike |  | Hua Zhen | 2018-06-28 | OESR: 182580. [26Jun18] Discussed with Flex, they will feedback about the feasibility and time required for this job done. [28Jun18] Sent the data to Mike. | 2-Normal (<4 weeks) | CLOSED | category | Grand Total | 9 | 9
S180008 | Test Engineering | 08.24.2018 | TE83013 | Upgrade | E cord interface test failure | WPPL | WPPL - Siew Yee | Milford-Matt | Hua Zhen |  | OESR 171993 Milford will upgrade the tester, in progress [20Sep18] Matt replied on 12 Sep 18 ,'New tester is still pending. Schedule has slipped a bit, as the new test is still a little unreliable. I am working on it and will be looking to you for help with the upgrade schedule as soon as I get it nailed down.' [24Sep18] Matt Brule update: the same status | 3- Low (> 4 weeks & above) | Open | Repair
S180009 | Test Engineering | 06.25.2018 | Leak Detector | New | Leak detector used in QDa and Beacon about to due for calibration. Please define the PM requirement and calibration requirement. | Venture | Loo Kok Leong |  | Siddharth |  | raised PO. 16-Jul-18: The leak bottle has been replaced with a leak bottle from the demo unit so that the calibration status of the instrument is not affected. The leak bottle that was ordered will be installed in the end of July. Servicing to be done in August, when there is no production. 30-Aug-18: The new calibrated leak bottle has been installed in the HLD. The calibration certificate has been given to Su Fern. Since we have not used the HLD a lot in the last year, after speaking with Paul, we have decided that the PM will be done next year after we buy the new HLD. | 2-Normal (<4 weeks) | Closed | Duplicate
S180010 | Manufacturing Engineering |  |  | Others | Please provide a summary of the following tubing used on the ACQUITY column manager test station: For this request, I am interested in the ID and length of each tube, including: 1. Pump Outlet tubing (BSM-SM) 2. UUT interconnect tube1 (SM-UUT) 3. UUT interconnect tube 2 (UUT-Detector) Also the detector/flowcell outlet Back Pressure Regulator (what is the pressure rating - 250 PSI, 500 PSI, other?) It would be good to have this information relatively quickly, so try to determine the tubing details as quickly as possible. Thanks, Mike | Milford | Milford-Mike |  | Hua Zhen | 2018-06-28 | The request has been fulfilled. | 2-Normal (<4 weeks) | CLOSED | New
S180011 | Test Engineering | 06.29.2018 | TE82525-2 | Repair | TE82525-2 Pump Tray Tester is keep on failing on Ripple Test \ Flow Test. Action taken to verify the failure: 1. Pin Plug FV2 during Ripple Test - Test Result PASS 2. Pin Plug FV2 during Flow Test - Test Resut PASS 3. Removed the Pin Plug during Ripple Test - Test Result FAIL and observed small amount of leak solvent during test. 4. Removed the Pin Plug during Flow Test - Test Result FAIL and observed small amount of leak solvent during test. 5. Therefore I conclude that the rootcause is leaky ASFVO Flow Valve. | Flex | Joel | Milford - Mark Jussaume | Hua Zhen |  | [29Jun18] OESR 182634 raised [03Jul18] Request for spares on vent valve P/N 7716-053 [13JUL18] Reassign to tackle in engineering meeting on long term requirement. [17Jul18] Pending for arrival of spares on vent valve P/N 7716-053 from Milford. [23Jul18] Mark placed order and will ship to WPPL once received it [30Jul18] Received valve 7716-053 [10Aug18] Flex replaced valve , it's running but unstable. It's under monitoring. [21Aug18] Flex reported the issue is still exsited and repair in progress [07Sep18] Oil from tester gas line was observed. A new OSER S180063 was raised for this issue. [21Sep18] Suspect the CPU PCB has issue with signal process. A spare CPU will be shipped from Milford. ETA next week | 2-Normal (<4 weeks) | Closed | Upgrade
S180012 | Test Engineering | 2018-01-06 | N/A | Others | We would like to request download the test record from Sigmasure 7 for the following Serial Numbers: PN : 186015080 SN E18QSP539A E18QSP536A E18QSP540A E18QSP542A E18QSP541A | Flex | Joel |  | Desmond | 2018-06-27 | Flex requests printer replacement, wait for Flex to get back latest status about this printer replacement Information provided to Flex | 2-Normal (<4 weeks) | CLOSED | Others
S180013 | Test Engineering | 2018-06-29 | TE82943-1 | New | Production is seeing a spike of 0.3ml/min flow rate accuracy failures (lower spec limit – 0.297ml/min) on QSM (186017000) in Apr . Please kindly refer to the "Brainstorming for 0 3ml flowrate test failure-29 Jun.pptx" for detail study. | Venture | Chan Myae Soe | Milford-Edward Jackson | Hua Zhen |  | [29Jun18] OESR 182632 raised [10Jul18] Remind sent to Ed Jackson [17Jul18] Infor from Milford shared with Venture [19Jul18] Venture check their tools for tubing connected to UUT and update FPY again [21Aug18] Chan did evaluation with increasing freq of chaning tubings and the torque wrench evaluation is ongoing [31Aug18] Venture Chan to review FPY data till Aug and do R1 & R2 Tubings inspect & evaluation on TE1 [19Oct18] OESR 182632 closed: This was closed previously and I'm closing again. The results in the report were not found in Milford and further results showed no spike in failures after the report date. | 3- Low (> 4 weeks & above) | Closed
S180014 | Test Engineering | 2018-07-02 | TE83064-16 | Repair | On 02-July-2018, Pump B of BSM of TE83064-16 has high leakage and cannot continue testing anymore . Both seals of primary and accumulator of pump B are required to be changed due to wear and tear. Affected part numbers to be changed are WAT076152(CRC O-ring 2.016) W1425000930-R (Seal Wash .0787 ID) W1425001037-R (Seal Wash Press Dual Supply) Since we are using these parts at Venture GES production, we replaced them for TE16. | Venture | EiLwin |  | Hua Zhen | 2018-07-02 | Fixed by replacing parts from Venture production and billed to WPPL | 1- High(< 1 week) | CLOSED
S180015 | Test Engineering | 2018-07-02 | 83086-2 | Repair | 1. Analog board 210000399 in GTS test fixture was burnt during testing of Aquity RI Persoanality board 210000699. 2. Request for replacement for 210000339 Analog board | CEI | CA Tan |  | Desmond | 2018-07-02 | 2 PCB from Milford were collected by CA for this repair | 1- High(< 1 week) | CLOSED
S180016 | Test Engineering | 2018-07-04 | TE82998-1 | Repair | 1.Testing for 210000422 Vista Pump PCBA at GTS was down due to bad Degasser board in test fixture. 2. Request for replacemnt and spare degasser board, p/n.725000459 Qaters QPSI Degasser. | CEI | CA Tan |  | Hua Zhen | 2018-07-10 | [04Jul2018] Sent kit PN to Jie Hui for purchasing of this PCB -; [06Jul2018] Christine received MRO [10Jul2018] Only 1 PCB was received, CA from CEI will collect it on 11 Jul 2018 | 1- High(< 1 week) | CLOSED
S180017 | Test Engineering | 2018-07-04 | TE82943-13 | Repair | TE82943 - 13 - Weighing balance hr-200 (s/n:12326184) showing unstable reading.(unable to get the zero reading) Action Taken : 1) Ensure that the bubble spirit level is center of balance. 2) Re-Zero the weighing balance but to no avail. 3) Ensure that balance stability is not influence by breeze or vibration. 4) Remove weighing pan,re insert the pan and Re-zero but no avail. 5) Press and hold the Re-Zero key until balance display CAL.Release the RE-Zero Key.Display shows "Error 1". | Flex | Joel |  | Yee Hwa |  | [05Jul18] Quotation in progress [10Jul18] ITS-ZiLun collected 2 balance from Flex on 09Jul18. To get status from Zilun on quotation and lead time [10Aug18] PO issued. Lead time 3 week. ETA:early Oct [3.Oct] ETA: WW41 | 3- Low (> 4 weeks & above) | Open
S180018 | Test Engineering | 2018-07-04 | TE82943-17 | Repair | TE82943 - 17 - Weighing balance hr-200 (s/n:12326279) has no display.(Faulty display) | Flex | Joel |  | Yee Hwa | 2018-07-20 | [05Jul18] Quotation in progress Similar as item S180017 [10Jul18] ITS-ZiLun collected 2 balance from Flex on 09Jul18. To get status from Zilun on quotation and lead time [20Jul18] Repaired and received on 18Jul18. | 1- High(< 1 week) | CLOSED
S180019 | Manufacturing Engineering | 2018-07-05 | 999027952-2 | Repair | Tooling 999027652-2 pin broken when use on Actuator Assembly for iBSM&QSM. There is no spare pin at Venture GES.(Need 4 pcs) | Venture | Chua Chong Keong |  | Desmond | 2018-07-26 | [05Jul18] GES Chong Keong to advise the reason for damage as the pins are recently changed in Apr 2018. [10Jul18] Chong Keong to follow up with supplier about material quality as all pins broken in 3 months. [26Jul18] Chong Keong received the part replacement - 10pcs. | 1- High(< 1 week) | CLOSED
S180020 | Test Engineering | 2018-01-06 | TE82475 | Repair | TE82475 - Pressure Decay Tester is not able to hold off pressure for position 1. Action Taken : Verified System Transducer during calibration to 1000 PSI - OK Verified Position 1 : during calibration to 1000 PSI - OK Verified Position 2 : during calibration 10 1000 PSI - OK Perform testing as per WAT025373 - to each position and result failed. Position 1 Pneumatic Valve 1 exhibit abnormal temperature during operation compare to other valve . Suspect to have issue with Valve Number 71216SN1BL00N0L111Q3 | Flex | Joel |  | Hua Zhen | 2018-07-13 | [06Jun18] Schematic from Milford forwarded to Flex [07Jun18] Cleaned check valve and fixed tester, Mike will order check valve for future needs [06Jul18] Checked with Mike regarding spare part of check valve [10Jul18] Mike replied they're waiting for check valve and follow up with Flex regarding BOM and tubing availability [13Jul18] Joel to follow up on tubing availability - to be track in weekly engineering minute. | 1- High(< 1 week) | CLOSED
S180021 | Test Engineering | 2018-07-06 | TE82988-12 | New | Tester or system not able to power on. This is due to the " AC Power distribution box " not working.It has been noticed that when the AC switch in the front panel of the tester is switch on there was an arching like sound generated in the AC Power distribution box. The arching sounds like behaviorally the same with the old GTS previously before we have replaced with new power distribution box. | Venture | Jason GES |  | Hua Zhen | 2018-07-20 | [06Jul18] OSER received [10Jul18] To confirm the failure and meanwhile for arrival of repaired Power box from Milford [13Jul18] Email from Paul asked to inspect the power plug for distribution box, Jason at Venture was informed [16Jul18] GES was informed to collect repaired power distribution box from WPPL and verify on new GTS see if the arching is caused by 3-phase power line capacity/stablity. [20Jul18] The issue is fixed by tightening power connector screws and case closed | 2-Normal (<4 weeks) | CLOSED
S180022 | Test Engineering | 2018-07-09 | TE82882-9 | Others | TE82882-9 and TE82882-14 Footer Mismatch with Serial Number. Action Taken: 1. Stop using TE82882-9 amd 14 2. Retest UUT affected to other station 3. Refer to attached print-sample | Flex | Joel |  | Yee Hwa |  | OESR 182647 [10Jul18] Joel will update OESR to reflect the all testers were affected by this issue [17Jul18] Create Sigmasure 7 access for Phani, Joel, Gary and CK. [20Jul18] Sigmasure 7 access have been created but Milford advise detector test data will not be able to share with Flex due to security reasons. TMD to be raise by Yee Hwa for the interim period till the tester is resolved by Eric. [14Aug18] TMD is not required per discussed with WPPL Quality team. Pending Eric to resolve the issue. [12Sep18] Bug fiexed in ECO9224731. | 3- Low (> 4 weeks & above) | Closed
S180023 | Test Engineering | 2018-07-11 | TE83064-13 | Repair | TE83064-13 is having system over pressure due to Column 4 (P/N:186002350) blockage. The column has been flushed in both direction to purge out the blockage but still causing system over pressure. Therefore, request column P/N:186002350 (Qty 1) for TE83064-13. Attached is a reference screen shot of system over pressure for reference. | Venture | EiLwin |  | Hua Zhen | 2018-07-13 | [12Jul18] Hong Cheng come WPPL to collect Column for repair [13Jul18] The overpressure issue was fixed by replacing column 4 | 1- High(< 1 week) | CLOSED
S180024 | Manufacturing Engineering | 2018-07-12 | Label Printer | Others | Found WAT024406 (Label, Serial Shipping Label) content is illegible during packing. Serial Number: G18SM4001A Request Waters personnel to assist to re-print Shipping label at Flex site. Qty: 1pc Root Cause: Mis-handling of Label | Flex | CK Tan |  | Yee Hwa | 2018-07-12 | [12Jul18] Case closed. This OESR is for record purpose | 3- Low (> 4 weeks & above) | CLOSED
S180025 | Test Engineering | 2018-07-12 | TE83064 | Others | Please send in the Test Equipment Qualification Report S/No Document Number Document Title 1 TE83064-10-TEQR_LTERATION 00 Test Equipment Qualification Report for TE83064-10 2 TE83064-11-TEQR_LTERATION 00 Test Equipment Qualification Report for TE83064-11 3 TE83064-12-TEQR_LTERATION 01 Test Equipment Qualification Report for TE83064-12 4 TE83064-13-TEQR_LTERATION 01 Test Equipment Qualification Report for TE83064-13 5 TE83064-14-TEQR_LTERATION 02 Test Equipment Qualification Report for TE83064-14 6 TE83064-15-TEQR_LTERATION 02 Test Equipment Qualification Report for TE83064-15 7 TE83064-16-TEQR_LTERATION 03 Test Equipment Qualification Report for TE83064-16 8 TE83064-17-TEQR_LTERATION 03 Test Equipment Qualification Report for TE83064-17 | Venture | Tan Hong Cheng |  | Edwin | 2018-07-23 | [13Jul18] Already sent IQ report of 2 Testers, will send the rest to Hong Cheng [16Jul18] Hong Cheng will check with their quality department and advise if can close this request after explaining to him about TEQR. [23Jul18] Qualification report was accepted by Venture quality and case closed | 2-Normal (<4 weeks) | CLOSED
S180026 | Test Engineering | 2018-07-12 | TE82998-1 | Repair | The CPU board at the 210000345 GTS fixture (TE#83001-3) suspected not working. To replace with a new board, we will require the CPU firmware. Attached is a request for CPU firmware for GTS FCT fixtures.. The firmware is required if the CPU board at the fixture need to be replace. | CEI | CA Tan |  | Edwin |  | [13Jul18] Edwin to review the follow up on the OESR 182660 [27Jul18] Edwin to follow up with Milford on failure analysis for the burned boards. [10Aug18] Mark was contacted for this issue and OESR 182660 was forwarded to his attention [3thSep18] Forward OESR, include cost of replacement and close issue - Edwin; The cost of pcba is 121.17 each, last quater we ordered 3pcs with total cost 353.61-HZ. | 2-Normal (<4 weeks) | Closed
S180027 | Test Engineering | 2018-07-12 | TE82988-3 | Repair | Tester was failing test " Serial Communication " and Suspected that the CPU is faulty or not able to booth up. Pass board have been tested too but also failing.Screenshot of the failure are attached in the email for reference. | Venture | Jason GES |  | Hua Zhen | 2018-07-30 | [19Jul18] Venture found a replacement CPU from the other fixture and GTS is working with this replacement now, but still need to get a permanent replacement from Milford [23Jul18] Mark will ship 2x cpu with firmware loaded to WPPL [26Jul18] PCBA pass to Jason. | 1- High(< 1 week) | CLOSED
S180028 | Test Engineering | 2018-07-13 | TE82998-1 | Repair | 1.Testing for 210000422 Vista Pump PCBA at GTS was down due to bad Degasser board in test fixture. 2. Request for one (1) replacement and two (2) spare degasser boards, p/n.725000459 Waters QPSI Degasser PCBA. | CEI | CA Tan |  | Hua Zhen | 2018-07-16 | [13Jul18] This GTS Degasser is frquently breakdown, further investiagetion is required. [16Jul18] The board was passed to CA and monitoring the issue now | 1- High(< 1 week) | CLOSED
S180029 | Test Engineering | 2018-07-16 | TE83031-1 | New | The existing optics bench alignment stations are using Golden flow cell to align all the optics benches.These Golden flow cells do not have the proper PM procedure to maintain their performance initially.Venture recently working with WPPL to release PM procedure for these golden flow cells. PM could help to improve flow cell performance for mechanically but it couldn't help when the flow cells are facing optically degaraded and againg. Milford Engineer agreed Venture should have one back up golden flow cell in any case of flow cell degrading over a few years of usage and for not interrupting production while it waiting Milford to send over new Golden flow cell. Request to send over new Golden Flow cells for Optics Alignment Test Stations. Please share with us how to perform the measurement and any specification threshold for replacement. Refer to attached photos of Ref photo diode flow cell. | Venture | MiMi |  | Mandy |  | Email from 16th July from Matt Franklin 'creating a new Photodiode Flow Cell and a New Golden Flow Cell to be used during the optics bench alignment station' Update on 14Aug from Matt, prototypes made and software due to be complete by the end of Aug. May take a couple of iterations to get the process/software to a good point. Expect these will be ready to ship to Sg by the end of the year. Current flow cells to be monitored until then. [15th October 2018] Matt has updated that the new flow cells will most likely be available at end of 18Q1 due to one staff member leaving and additional testing being carried out. | 3- Low (> 4 weeks & above) | Open
S180030 | Test Engineering | 2018-07-17 | TE82774 | Others | Redline TSP 735000698 add Install Dust Cover, Part Number: xxx into the 26 pin D-Sub connector on page 14 of 36. | Venture | Tan Hong Cheng |  | Edwin |  | [17Jul18] check with Hong Cheng to further understand ther request [27Jul18] Edwin to review redline for change to sent to Milford. [20Sep18] Engineering felt this is an isolated case so they do not wish to proceed with new part number unless there is more cases. Recommend to close | 2-Normal (<4 weeks) | Closed
S180031 | Test Engineering | 2018-07-17 | TE82817 | Others | Please help to down down TSP 735000887 in word format for Venture GES to mark up.14 of 36. | Venture | Tan Hong Cheng |  | Edwin | 2018-07-20 | [17Jul18] Edwin to advise if should send word format TSP to Venture? [20Jul18] Venture will redline on PDF. | 2-Normal (<4 weeks) | CLOSED
S180032 | Test Engineering | 2018-07-18 | TE83026 Column Manager | Others | Sigmasure6 could not be reached. "Problem Loading Page - Unable to connect" message is returned eventhough sigmasure7 is working. | Venture | CHAN MYAE SOE |  | Hua Zhen | 2018-07-19 | [19Jul18] Confirmed with Bill the server is down and no solution for time being | 2-Normal (<4 weeks) | CLOSED
S180033 | Test Engineering | 2018-07-18 | TE82995-All | Others | In Signmasure 7, we cannot access the test data of sub-assembly testers TE82995 -2 and TE82995 -3. Please refers to screen shot attached which shows what we could access at Sigmasure 7. Please help to provide test data for TE82995-2 & TE82995-3 from Jan 2017 to till todate. The purpose of this request is to check Needle Seal Force test data trend by sub-assembly testers. | Venture | EiLwin |  | Hua Zhen | 2018-07-23 | [23Jul18] Data sent to Lwin. | 2-Normal (<4 weeks) | CLOSED
S180034 | Test Engineering | 2018-07-20 | TE82882-14 | Upgrade | 1. TE82882-14 Shunt \ Cell Energy 6ms reading is (0) and Outcome FAIL, the 2998 Refresh Final Test Overall Result - PASS , Eric Turner mentioned that "That particular test is not an official spec" 2. TE82882-14 Noise Run result is FAIL and but the result 2998 Referesh Final Test Overall Result - PASS. Flex Acton Taken : Restest UUT and if pass release for shipment Pls. refer to attached file for test reult. | Flex | Joel |  | Mandy |  | OESR Created 30th January Log no. 182351 waiting for software update 13th Aug updates. Software due to be updated within the next two weeks. Update on 30th August - waiting for ECR approval before releasing V4.1 software | 2-Normal (<4 weeks) | Closed
S180035 | Test Engineering | 2018-07-24 | TE82984-01 | New | "Valve Jumper" test step did not stated in Appendix of TSP 735002056. Jumper setting of 186015008 is different from others. Would like to request to add valve jumper step in TSP. Attached graphical user interface in .bmp format. | Venture | Chan Myae Soe |  | Hua Zhen |  | [27Jul18] Request Redline TSP from Chan [10Aug18] The word format of TSP sent to Chan as per request, waiting her reply with Redline TSP [14Aug18] Chan in Venture update the latest TSP has covered the change and case closed | 2-Normal (<4 weeks) | Closed
S180036 | Test Engineering | 2018-07-24 | NA | Others | Request to download test data from 1st Jul 2018 to 24th Jul 2018 (or until the day of download) of 186015042, 186015043, 18601545, 186015007, 186015008 and 186015009. Time zone is China Time | Venture | Chan Myae Soe |  | Hua Zhen | 2018-07-24 | [24Jul18] Data provided to Chan. | 2-Normal (<4 weeks) | CLOSED
S180037 | Test Engineering | 2018-07-24 | TE82987-01 | Repair | No power on Ethernet adapter. Confirmed that ethernet adapter is faulty. | Venture | Chan Myae Soe |  | Hua Zhen | 2018-08-14 | [27Jul18] Email approval given to Chan to purchase an off the shelf ethernet switch. [14Aug18] The switch has been purchased. Case closed | 2-Normal (<4 weeks) | Closed
S180038 | Test Engineering | 1900-07-24 | Certificate Print | Repair | Waters Corporation Instrument Certificate Printing WATERSCOFP1 HP Laser Jet Enterprise M604 Encounter Run-time error '68' Device unavailable, Cannot link to Server. Flex try out below step :- 1.Re-start the system >>Still unable to print and occur the same error . | Flex | Gary |  | Yee Hwa | 2018-07-26 | Bill has reconfigure the server. The error due to file tranferring by Dave Moss. | 2-Normal (<4 weeks) | Closed
S180039 | Manufacturing Engineering | 2018-07-25 | NA | Others | On 5th od March 2018 Venture sent 4 pcs of Gratings P/N 4020000130 to Milford Enginner (Attn: Nicholas Szwaja) for evaluation purpose. Refer to attached Invoice attached and email chain. We would like to know the disposition of this 4 pcs of gratings. | Venture | MiMi |  | Edwin |  | [27Jul18] Edwin to check with Milford for cross charged [3thSep18] To issue PO to Venture on 4 grating - Edwin; Jie Hui has issued PO 2018-326 on 07 Aug 2018 - HZ. | 3- Low (> 4 weeks & above) | Closed
S180040 | Manufacturing Engineering | 2018-07-26 | Server and 3x printer | New | Insert Tape Media and observed that Waters Server IP 10.249.26.1 showing "There are offline or missing virtual drive with preserved cache. Please check th ecable and ensure that all drives are present. Refer to attached photo (a), the lower drived SAS 600GB 15k, the LED is blinking green orange.(b) 2 After the server down, the 3 printer is also not able to work: 2.1 Ceritificate of Functional Performance - Run-Time error ‘68’: Device unavailable 2.2 WATERSSPNL1-The drive or network connection that the shortcut refer to is unavailable 2.3 WATERSISNL1 - The drive or network connection that the shortcut refer to is unavailable | Flex | Joel |  | Yee Hwa | 2018-07-27 | Bill has reconfigure the server. The error due to file tranferring by Dave Moss. | 2-Normal (<4 weeks) | Closed
S180041 | Test Engineering | 2018-07-27 | TE83064-13 | Repair | BSM of TE13 had high leak rate 6169 nL/min at pump A accumulator testing with 17k psi and leak rate 4613 nL/min at pump B primary when testing with 15k psi. Troubleshooting Summary: - Changed each pump seals set for pump A accumulator and pump B primary. - After seals are changed, pump B has no more leakage. But pump A cannot compress pressure up to 17k psi. To solve this problem, changed new assy vent valve 279002426. - After vent valve is replaced, pump A can compress up to 17k psi, but there is still leakage at pump A accumulator so that check valve 289003360 is changed.Then,there is no more leakage at pumpA accumulator. - Re-verification is done by installing back old vent valve 279002426 and pump A cannot compress to 17k psi. | Venture | EiLwin |  | Hua Zhen | 2018-08-10 | [27Jul18] Venture has been asked to provide the repair parts price [10Aug18] The repair was done by Venture and claimed repair cost quaterly | 2-Normal (<4 weeks) | Closed
S180042 | Manufacturing Engineering | 2018-07-30 | NA | New | was notified that the ASP 186015020 IVD Acquity UPLC Sample Organizer, IVD for the info set was different from what we packed (refer to pg 11 of 5). The ASP was stated to throw the info set in between the Unit and the Startup kit. However during the product tranfer the info set was told to place on top of the instrument. Reason was to prevent from damaging the info set when throwing in. Since then we practice according and did not notice ASP show to throw in and ASP 186015020 Packaging, Sample Organizer ASO has changed to place on top of the instrument. Would like you review and change the ASP to place the info set on top of the instrument.hankyou. | CEI | Jason Wang |  | SuFern |  | Pending Desmond guidance on submitting ECR. [21Sep18] The ECR is submitted. [25Sep18] ECR closed | 3- Low (> 4 weeks & above) | Closed
S180043 | Test Engineering | 2018-07-31 | TE82882 | Duplicate | TE82882-8 and TE82882-10 Footer Mismatch with Serial Number. Action Taken: 1.TE82882- PN :18600XXXX SN : G18998922A , G18998919A , G18998895A | Flex | Joel |  | Hua Zhen | 2018-07-31 | [31Jul18] Data sent to Joel | 2-Normal (<4 weeks) | CLOSED
S180044 | Manufacturing Engineering | 2018-08-01 | 999028067 | Repair | During the assembly of Housing Check Valve (405004106) and Holder Cartridge (405004105) as per ASP 289001458 REV C, the swage fixture bushing break and unable to perform the press operation. This is affecting production line down for 186269505 & 186269506 models. | Flex | Gary |  | Desmond |  | [14Aug18] Tool have been fabricated and used in production. [15 Aug 18] Original part has passed to Gary and to be installed same day. | 1- High(< 1 week) | CLOSED
S180045 | Test Engineering | 2018-08-02 | TE82882-8 | Duplicate | TE82882-8 test result mismatch footer: Request to download from server: PN : 186299800 SN : H18998927A | Flex | Joel |  | Hua Zhen | 2018-08-02 | [31Jul18] Data sent to Joel | 2-Normal (<4 weeks) | Closed
S180046 | Test Engineering | 2018-08-02 | TE82999-1 | Repair | 1. Gasket for the GTS test fixture torn, causieng vacuum leakage . 2. To replace the gasket. | CEI | CA Tan |  | Hua Zhen |  | [15Aug18] Checked with CA regarding this repair request [16Aug18] CA feedback no other vendor to do this replacement and seeking another way (Top Cover) to fixing issue. Still checking for details [21Sep18] Team agreed to go ahead with Gasket replacement and CA is arranging the vendor to start this replacement. | 3- Low (> 4 weeks & above) | Closed
S180047 | Test Engineering | 2018-08-06 | TE83064-14 | Repair | TE83064-14 is having system over pressure due to Column 4 (P/N:186002350) blockage. The column has been flushed in both direction to purge out the blockage but still causing system over pressure. Therefore, request column P/N:186002350 (Qty 1) for TE83064-14. Attached is a reference screen shot of system over pressure for reference. | Venture | EiLwin |  | Hua Zhen | 2018-08-06 | [06Aug18] Passed the column to Venture | 1- High(< 1 week) | Closed
S180048 | Test Engineering | 2018-08-06 | TE82882-8 | Duplicate | TE82882-8 test result mismatch footer: Requst to pull-out from server: PN : 186299800 SN : H18998939A H18998940A TE82882-9 test result mismatch footer PN: 186015033 SN: H18UPL128A H18UPL126A | Flex | Joel |  | Hua Zhen | 2018-08-02 | [31Jul18] Data sent to Joel | 2-Normal (<4 weeks) | Closed
S180049 | Manufacturing Engineering | 2018-08-08 | WatersPNL1 | Repair | ASP 735003655 Instrument Serial Number Printing -STD Encounter Quantity Barcode with white line .(Cause by printer head) Has swap with another printer head still having same issue white line at Description This is affecting to all Waters Instrument model. | Flex | Gary |  | Yee Hwa | 2018-08-10 | [10Aug18] Printhead replaced and printer resumed operation on 10th Aug. | 1- High(< 1 week) | Closed
S180050 | Test Engineering | 2018-08-13 | TE82546 | Repair | TE82546 2690 Display/Keypad Test fail as per below : Diag: Modle:display Test DisplayMemoryCheck…DisplayMemoryCheck: Fail, Address 4606CF Val: ff Its is due to Pin 1 broken at PN 270084 Rev 0, ATX 09/03 | Flex | Joel |  | Hua Zhen | 2018-08-15 | [13Aug2018] Flex directly mount the cable 270084 to test board as temp solution while sourcing a replacement connector [15 Aug 18] Joel replaced cable and verified working. | 2-Normal (<4 weeks) | Closed
S180051 | Manufacturing Engineering | 2018-08-14 | NA | Others | ECO 9222683 Life cycle change on Hamamatsu and Optic Bench with new project number . SigmaSure show old Optic Bench project number.(Please refer to attachment) 1) ASP 279002224 Assy Detector ,2998 SFC Del 279002225 , Add 279003843 2) ASP 279002897 Assy, Detector, PDA Del 279002903 ,Add 279003842 3) ASP 279000805 Assy, Detector 2998 Classic Del 279000647 ,Add 279003841 | Flex | Gary | Matt Franklin | Yee Hwa |  | [15 Aug 18] Raised OESR-182727 to Chau to add in new PN to SS7. [19Sep] Added by Matt on 19Sep; Gary to verify 3+3 | 2-Normal (<4 weeks) | Closed
S180052 | Test Engineering | 2018-08-16 | TE82705-2 |  | Unable to flow out solvent at PUMP A when running 1525 unit due to faulty valve(FV6). Need to replaced BSDA025T-NC PV valve (FV6) which control the flow of solvent to PUMP A. | Flex | Joel |  | Yee Hwa | 2018-08-16 | Flex collected Valve from Yee Hwa and the problem was fixed after replacing valve | 1- High(< 1 week) | Closed
S180053 | Manufacturing Engineering | 2018-08-17 | Wat271987 | Others | ECO 9221408 Update Degasser Vent WAT271087 Assy, Vacuum Degasser Vent 2690 . Rev T Encounter sheet 4 of 8 item 19 Part Number WAT271157 without part drawing show. | Flex | Gary |  | Yee Hwa |  | Feedback to Milford Eng. Team, New drawing revV will be release. ECR9224681 | 2-Normal (<4 weeks) | Closed
S180054 | Test Engineering | 2018-08-17 | TE82760 | New | On the month of Jun & July, Venture production found some failures in Ram-Ray Valley Energy during FLR detectors final test for 5029 models. The UUTs were passed when they replaced with new PMTs. The failed PMTs were used in 2475 models detectors, passed the test. Venture requested Hamamatsu to provide the failed PMTs test data and performed some data analysis and some evaluation to find the correlation between Hamamatsu test data and our test based on our knowledge. Refer to attached slides for details. OESR is requested to be raised to follow up and get some advices from Milford Engineer to look into the possible root causes if this failure occur in future. | Venture | MiMi |  | Edwin |  | No failures of this type since 23rd July. 17th August, Edwin will make contact with Milford engineering to get more details on the type of failure that is happening. [3thSep18]: This is a part related issue on Hamamatsu Lamp. Reassigned to Edwin to follow up. [20thSep18]: No advise yet from Matt/ Nic, will continue to follow up | 3- Low (> 4 weeks & above) | Open
S180055 | Test Engineering | 2018-08-17 | TE82503-4 | Repair | TE82503-4 Injector Seal Pack Not able to hold pressure. 1. Ramp and Decay Failed - able to ramp pressure to 5000 PSI but not able to hold-off pressure for the last 3 mins. 2. Action Taken : Verify Calibration : OK, able to calibrate Verify Fittings / tubing : No visible leak from fittings / tubing. Verify Filters : No visible dirt on filters / replaced filters Repair in-progress | Flex | Joel |  | Hua Zhen |  | Fixed by replacing air switch and valve | 1- High(< 1 week) | Closed
S180056 | Test Engineering | 2018-08-21 | NA | Others | Request to download test report of S/N: H18CMP578G (P/N: 186015043). Reason: One test report could not be printed out due to connector from printer to PC. | Venture | Chan Myae Soe |  | Hua Zhen | 2018-08-23 | [23Aug18] sent test record to Chan | 2-Normal (<4 weeks) | Closed
S180057 | Test Engineering | 2018-08-23 | TE83064-15 | Repair | TE83064-15 was having system pressure unstable. Please refers to attached photo. It was required to change I2V check valve for BSM system of tester. Part number required to change is P/N:205001669 (289003706) ASSY, CARTRIDGE, I2V, HEXANE/THF, 1PC . | Venture | EiLwin |  | Hua Zhen |  | [23Aug18] After changing 205001669 (289003706) CARTRIDGE I2V HEXANE/THF KIT (Qty 1) tester can resume testing. | 2-Normal (<4 weeks) | Closed
S180058 | Test Engineering | 2018-08-23 | TE83064-11 | Repair | E83064-11 was severe leakage at pum B primary. Please refers attached screen shot. It required to be changed one seal set of BSM for Pump B primary. | Venture | EiLwin |  | Hua Zhen |  | [23Aug18] After changing BSM seals WAT076152(CRC O-ring 2.016)- Qty 1 W1425000930-R (Seal Wash .0787 ID)- Qty 1 and W1425001037-R (Seal Wash Press Dual Supply) – Qty 1, tester can resume testing. | 2-Normal (<4 weeks) | Closed
S180059 | Test Engineering | 2018-08-23 | TE83064-12 | Repair | TE83064-12 is having 1) blockage at right inlet valve, P/N:289006760 of column manager. 2) leakage at right valve inlet tubing P/N:430003154. Please refers to attached photos for reference. We'd like to request P/N:289006760 (Qty 1) and P/N:430003154 (Qty 1) for replacement. | Venture | EiLwin |  | Hua Zhen |  | [24Aug18] Repair Parts MRO was raised, waiting for delivery. [29Aug18] Faulity parts are replaced. | 2-Normal (<4 weeks) | Closed
S180060 | Manufacturing Engineering | 2018-08-24 | COFP | Others | UUT Serial Number H18TUV969A Certificate of Functional Performance printed WRONG Model Description 5844441 instead of ACQUITY UPLC TUV DETECTOR. Flex would like to request to amend the COFP from 5844441 to ACQUITY UPLC TUV DETECTOR. (refer to scan sample) | Flex | Joel |  | Desmond |  | OESR 182748 [28Aug18] Bill Iron Inactivated the record and can reprint now | 2-Normal (<4 weeks) | Closed
S180061 | Test Engineering | 2018-09-03 | TE82720-2 | Repair | The tester PC always keep rebooting when PC is power on so not able to use for testing. Currently there are boards W1210000814 for testing. The following are the other PCBA list qualified to test using the tester ( W1210000423-R, W1210000399-R, W1279000592/Z, W1279001098, W1081318 ) | Venture |  |  | Hua Zhen |  | [03Sep18] OESR 182765 [05Sep18] Rick will ship a replacement PC on 05 Sep in Milford [07Sep18] ETA 10Sep18(Monday) [20Sep18] Tester was fixed by replacing PC. The PC is online and info is updated. | 1- High(< 1 week) | Closed
S180062 | Test Engineering | 2018-09-03 | TE82903 | Others | SPARE PN : 700010507 SN 35J7V025 not able to test due to TE82903 Tester GUI 700010507 not able to accept the SPARE serial number 35J7V025. TE82903 Test Software 700010507 is showing INVALID SERIAL NUMBER (refer to screen shot 700010507) If using 700010507 with SN format for SN H18VTU025A, the GUI 700010507 is able to proceed on testing.(refer to screen shot SPARE 700010507). If not able to resolve, this will impact 1st shipment of 700010507 (1 unit) by 10 Sep 2018 | Flex |  |  | Edwin |  | 11th Sep Edwin: Action to be closed. Flex should be using the new Water serial no# format per confirmation. Joel has confirmed he is able to proceed with testing using new serial no# format. No further action required. | 1- High(< 1 week) | Closed
S180063 | Manufacturing Engineering | 2018-09-06 | TE82526-2 | Others | TE82525-2 Pump Tray Tester filter detected the presence of oil and may pose risk to operation of pumptray as per assessment with Mandy Stewart (WPPL). Action Taken : - Flex installed another compress air filter before going to pump-trays. Recommendation from Mandy: - To request and replace 3x filters Model : F504-01AH \ SN 9512 - Vendor: Watts Fluidair, Inc Kitter, Maine 0394 | Flex |  |  | Yee Hwa |  | [13Sep18] Discussed with Phani to deliver expectation (applicale to all input): - Tidy up external filter, with labeling of service," where use" - Find out source of oil mist,any way to prevent or reduce - Imporve PM to make sure no oil mist [26 Sep] Develop a project plan on inprovement and PM. To be track in Eng Meeting weekly. | 2-Normal (<4 weeks) | Closed
S180064 | Test Engineering | 2018-09-07 | TE83064-12 | Repair | TE83064-12 was having severe leakage and cannot pass dynamic leak test 06-Sep-18. Please refers to attached screen shot for reference. ASSY VALVE VENT 18 K, P/N:279002426 was required to be changed at BSM of tester. | Venture |  |  | Hua Zhen |  |  | 2-Normal (<4 weeks) | Closed
S180065 | Test Engineering | 2018-09-07 | N/A | Others | Request to download test data from 1st Aug 2018 to 07 Sep 2018 of 186015043. Time zone is China Time | Venture |  |  | Hua Zhen |  | [07Sep18] The data has been supplied to Chan | 2-Normal (<4 weeks) | Closed
S180066 | Test Engineering | 2018-09-13 | TE83086-2 | Repair | Analog board 210000399 in GTS test fixture was burnt during testing of Aquity RI Persoanality board 210000699.2. Request for replacement and spare 210000339 Analog board | CEI |  |  | Hua Zhen |  | [17Sep18] CEI collected the PCBA from WPPL on 14Sep18 | 1- High(< 1 week) | Closed
S180067 | Manufacturing Engineering | 2018-09-13 | 999028048 | Duplicate | Wear and tear of tool. Require replacement. [19Sep] Flex to source spare; no impact to production | Flex |  |  | Yee Hwa |  | Flex to source replacement quotation. 999042121. Quatation provided, decided Yuan Heng on 28Sep ETA: end Oct | 3- Low (> 4 weeks & above) | Open
S180068 | Test Engineering | 2018-09-14 | TE82999-1 | Repair | 1. The 210000412 pcba in the TE82999 CPSM Control GTS fixture found shorted causing test failurefpr 210000367 and 210000903 boards. 2. Request for repalcement board for repair. | CEI |  |  | Hua Zhen |  | [17Sep18] MRO was raised on 14 Sep 18; [20Sep18] CA collected the PCB from WPPL on 20sep18; | 1- High(< 1 week) | Closed
S180069 | Test Engineering | 2018-09-17 | TE83064-15 | Repair | TE83064-15 was having system over pressure blockage at column 4 of Column Manager. Please refers to attached photo. Tried to purge out blockage in both directions of column 4 but blockage still remained. Column 4 part number 186002350 (qty 1) is required to changed. Please provide it since we don't have spare stock in hand. | Venture |  |  | Hua Zhen |  | [20Sep18] The issue is fixed by replacing column | 1- High(< 1 week) | Closed
S180070 | Test Engineering | 2018-09-20 | TE83064-17 | Repair | TE83064-17 was having high leakage at pump B primary. Please refers to attached photo for reference. Pump B primary of BSM of TE17 was required to change BSM seal so that we have changed seals as per below seal list. WAT076152 (CRC O-ring 2.016) ;W1425000930-R (Seal Wash .0787 ID); W1425001037-R (Seal Wash Press Dual Supply) | Venture |  |  | Hua Zhen |  | [21Sep18] Case closed by replacement seals | 1- High(< 1 week) | Closed
S180071 | Test Engineering | 2018-09-20 | TE82988-9 | Repair | 1. One power supply at GEOTEST chassis GX7310 having intermittently on-off problem. The Geotest chassis is in GTS tester TE82988-9. 2. Power Supply part number is PRM560. Requesting a replacement unit. 3. This GTS tester is mainly for testing PCBA for Waters UK (Micromass). | CEI |  |  | Hua Zhen |  | [21Sep18] CEI should provide the clear spec of the power supply and the repair cost arrangement between CEI and Waters UK. [24Sep18] CEI will check with Waters UK about the repair cost [02Oct18] Written to Raymond Cribbin on 24 Sep 18. No reply yet. Updated by CA | 3- Low (> 4 weeks & above) | Open
S180072 | Test Engineering | 2018-09-21 | TE82882-13 | Others | TE82882-13 Test station test software version 4.1 completed test but not able to print report. IP ADDRESS : 10.249.26. 1. Verify network connectivity and able to ping server 2. Verify test print using the computer and is able to print the report 3. Suspect is due to access by tester account is block 4. Flex to leave UUT for Eric to access tonight 5. Raise OESR for support | Flex |  |  | Hua Zhen |  | [21Sep18] OESR182806 was sent to Eric for this urgent case on 21Sep18; [24Sep18] The workaround is print it out from Sigmasure | 1- High(< 1 week) | Closed
S180073 | Test Engineering | 2018-09-21 | TE82969-1 | Repair | TE82969-1 Not able to do the linearity Test for 186015032 IP Address : 10.249.26.75 Test Result as below: % RSD Measurement 1.#QNAN Outcome FAIL Mean Measurement 1.#INF Outcome PASS Std Dev Measurement 1.#QNAN Outcome PASS We will leave one unit for Eric for Eric configuration. | Flex |  |  | Hua Zhen |  | [21Sep18] OESR182805 was sent to Eric for this urgent case on 21Sep18 [24Sep18] The problem is fixed by upgrading software to 2.1, an IQ is required for this tester | 1- High(< 1 week) | Closed
S180074 | Test Engineering | 2018-10-01 | Hipot Tester/Load | Repair | Load Box S14-03 SN 14354681 able to pass calibration. - Hipot Testreading on the instrument is OK but failed Hipot Test - Suspect to be Hipot cabe S02 , HV Lead Set, high & low, 1m , pls. send replacement cable | Flex |  |  | Hua Zhen |  | [02Oct18] Load Box S14-03 SN 14354681 able to pass from other station, but failed at Alliance Hipot at 400Kohm, Need to replace S02 Cable | 2-Normal (<4 weeks) | Closed
S180075 | Test Engineering | 2018-10-01 | Hipot Tester/Load | Repair | Load Box S14-03 SN 14354695 able to pass calibration. - using Resitance test, no reading and observed loose component when shake. - Reapir in-progreass | Flex |  |  | Hua Zhen |  | [02Oct18]Load Box S14-03 SN 14354695 have loose connection. Resolder back and its able to pass testing. Verify at Detector Hi-pot, passed all 3 tests. | 2-Normal (<4 weeks) | Closed
S180076 | Test Engineering | 2018-10-02 | Column Manager | Others | Request to download test data from 1st Sep 2018 to 30 Sep 2018 of 186015043. Time zone is China Time | Venture |  |  | Hua Zhen |  | [02Oct18] Sent data to Chan | 2-Normal (<4 weeks) | Closed
S180077 | Test Engineering | 2018-10-03 | TE83064-15 | Repair | TE83064-15 has leakage and it is confirmed that cannot pass dynamic leak test. Based on checking, we propose to change pump B primary seal set and vent valve 205001476. Please refers to attached file for BSM leak test result. Please acknowledge for the proposed changes. | Venture |  |  | Hua Zhen |  |  | 2-Normal (<4 weeks) | Closed
S180078 | Test Engineering | 2018-10-03 | Hipot Tester/Load | Repair | Load Box S14-03 SN 14354681 able to pass calibration. - Hipot Testreading on the instrument is OK but failed Hipot Test - S02 , HV Lead Set, high & low, 1m , pls. send replacement cable / or approval 2 buy 2 cable, 1 for replacement and 1 for Spare. | Flex |  |  | Hua Zhen |  |  | 2-Normal (<4 weeks) | Closed
S180079 | Manufacturing Engineering | 2018-10-04 | Label Printer | Others | Production encountered one missing label in the label roll which need to reprint the serial number for 186015201 K18UPP007H. I have checked the printer internally and could not find the missing label. Has confirmed this label has missing sticker on it. We need your support to come to CEI to re-print this label. | CEI |  |  | Yee Hwa |  | Reprinted on 5.Oct. | 1- High(< 1 week) | Closed
S180080 | Test Engineering | 2018-10-09 | Hipot Tester/Load | Repair | Found the sleeve of the Power entry adapt cable for Electrical Safety Analyzer Chroma 19032 is loose and because of high voltage testing. The cable has been using for almost 5 years plus. Would like to request for replacement for this cable. | CEI |  |  | Hua Zhen |  | [12Oct18] PO sent to CEI | 2-Normal (<4 weeks) | Closed
S180081 | Manufacturing Engineering | 2018-10-10 | TE82943 | Others | BSM TSP: 735002010 Rev L, Software version: 2.2 QSM TSP: 735002155 Rev H, Software version: 1.7 User not able to sign in to test program at QSM & BSM station. Error message indicates valid ID must be included in Operators.txt file located on PC's Desktop. Verify Operators .txt file, user ID is in the file. 1. Mohd Redzuan; ID: 1074371. Unable to access Tester 17 2. Syed Shamsudin; ID 847838. Unable to access Tester 11 to Tester 18 | Flex |  |  | Yee Hwa |  | [12Oct18] Raised OESR182851 to Milford. | 2-Normal (<4 weeks) | Closed
S180082 | Test Engineering | 2018-10-11 | TE82519 | Repair | TE82519 2690 Pump Mechanics Tester failing on Ramp and Decay test. Action taken to verify the failure: 1)Replaced external interface tubing p/n:WAT024024-07 2)Retightened all internal interface tubing and ensure no leaking in any of them. 3)Replaced inline filter p/n:WAT035190 4)Ensure that Flow valve p/n:AFSVO for FV1 and FV2 has no leakage 5)Replaced Mirco Solenoid valve p/n:M401-37 for Flow Valve 1 due to air leakage at the end of poly tubing going to flow valve 1. 6)Unit Pass after replacing micro solenoid valve p/n:M402-37 for Flow valve 1. Conslusion: We need to have spares of micro solenoid valve p/n:M402-37 in order to refrain from downtime. | Flex |  |  | Hua Zhen |  | [12Oct18] Tester is fixed by Flex | 2-Normal (<4 weeks) | Closed
S180083 | Test Engineering | 2018-10-16 | TE82998 | Others | 1.Due to frequent failure of 725000459 Degasser board in GTS fixture for 210000422 pcba testing, we will like to request one or two spare boards. 2. This is to expedite repair of the fixture when need to prevent downtime to production testing. | CEI |  |  | Hua Zhen |  | [18 Oct 18] The PCB is ready for collection | 2-Normal (<4 weeks) | Closed
S180084 | Test Engineering | 2018-10-17 | Waters Portal | Others | During the ASP and TSP review for 2nd half yearly, i am not able to open all the TSP for CPSO and ASO model in Waters Portal. It show that 'No authority for this Material' . Please see attached pictures. Please advice. Thanks | CEI |  |  | Hua Zhen |  | [18Oct18] sent email to CEI to have details about this request. [19Oct18] OESR 182865 was raised [23Oct18] Asked Wang to list all affected TSP and persons. OESR was raised to Milford. Monitor the status; Milford TE group advised to have BT to support this issue. BT Request ID: B5S4RG | 2-Normal (<4 weeks) | Closed
S180085 | Test Engineering | 2018-10-18 | TE82998-12 | Repair | When loading the test software for W1210000478 & W1210000400 both are for USA PCBA , there is an error loading pop up when executing or clik the " Test UUT " button after selecting PCBA number to test. Please refer to the error screenshot attached in other page. | Venture |  |  | Hua Zhen |  |  | 3- Low (> 4 weeks & above) | Open
S180086 | Test Engineering | 2018-10-25 | TE82988-2 | Repair | 1. The model 210000345 not be able to pass the testing. Error message prompted "Test +5V VCC Voltage Monitor Readback". Note : Currently we have 86 pcs boards pending for test and shipment will be on risk . | Flex |  |  | Hua Zhen |  | [26Oct18] A CPU board was loaded with firmware 1.46 and it's fixed by 25 Oct 2018 | 1- High(< 1 week) | Closed
S180087 | Test Engineering | 2018-10-29 | TE82720-1 | Repair | Waters Universal Front End Test Platform TE82720-1 down not able to initialize the ATE when continue to the diagnostic it shows the “SYSTEM POWER SUPPLIES” Is failing at +5V . We measure the power module there is no +5V from the output . Please advice is it the +5V power module is bad ? | Flex |  |  | Hua Zhen |  | [30Oct18] Advised Flex Senai to check power supply [02Nov18] Fiexed by checking the test fixture | 1- High(< 1 week) | Closed
S180088 | Test Engineering | 2018-10-29 | TE82895 | Repair | . The model 404 not able to continue testing. Fail at "Voltages test". Note : shipment will be at risk . | Flex |  |  | Hua Zhen |  | [30Oct18] Advised Flex Senai to check power supply [02Nov18] Fiexed by restored USB cable | 1- High(< 1 week) | Closed
S180089
S180090
S180091
S180092
S180093
S180094
S180095
S180096
S180097
S180098
S180099
S180100
S180101
S180102
S180103
S180104
S180105
S180106
S180107
S180108
S180109
S180110
S180111
S180112
S180113
S180114
S180115
S180116
S180117
S180118
S180119
S180120
S180121
S180122
S180123
S180124
S180125
S180126
S180127
S180128
S180129
S180130
S180131
S180132
S180133
S180134