## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | TE Tool No Instance | Request | Originator | STATUS | MPP Date | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2 | TE Tool No Count
160595 | Test Engineering | NA | 2016-01-08 | NA | NA | International | Joel(Flex/Chia(WPPL) | CLOSED |  | 2016-01-08 | For security & Control: the following are request for all workstations by Flex Singapore manufacturing; 1) Enable Screensaver (15mins). Prompt for password to resume. 2) Review the “Power Option Properties”, it is set to “Turn OFF monitor After 5 Hours” and at the “Power Option Properties (Advance)”, it will “Prompt for password when computer resumes from standby”. This might be the issue for instances where Milford cannot access Test Stations if the system goes into standby. 3) To disable INTERNET access for TESTENG access. | The screensaver is enabled on all accounts. Admin-level accounts have a short timeout for security reasons, but user-level accounts have a longer timeout (to avoid reentering the password multiple times per day per station). This is standardized across test stations worldwide and is configured on the system images. I wonder why Flex Singapore would require a different timeout than other plants. Was there a particular reason for this request? The monitor should shut down after five hours, but the computer should not go into standby mode. I will make a note to verify this when the images are revised. Internet access is required for support, maintenance, upgrades, test data (SigmaSure server is off-site). It's not possible to run these systems without internet access. No changes were made, so there is no qualification required. | Michael Cotter |  | 357
160596 | Test Engineering | 279003278 | 2016-01-08 | NA | NA | International | Flextronics/Chia (WPPL) | CLOSED |  | 2016-05-25 | Currently Flex have 3 Tech is allow to access to Loging into Velox SigmaSure. Flex have another 3 trained personnel (below) for Assy, Detector,2489 W/O flow cell / Assy , Detector 2998 W/O flow cell and request to grant access to SigmaSure. 1) Amirah Binte Abdullah 927986 2) Ow Eng Neng 931524 3) Amir Nashrul Bin Abd Jalil 912394 | 25-May-2016 - accounts created The SigmaSure genealogy system will be emailing you to validate your account. It will ask you to visit the web address and change your password after you enter the temporary password that is emailed to you. We recommend you use waters as your password. The address to visit is http://10.2.242.127/sigma Jan 26, 2016: Eric was out. Feb 24, 2016: Eric "we have a request to add 3 operators to SigmaSure genealogy at Flex (2998 line). It was my understanding that we were moving towards having Flex use their own genealogy system. We discussed with Edwin before the end of last year and I'm not sure where we are with that. If they need operators added to the tester, I can do that right away, but I want to be sure we plan to use SigmaSure genealogy at Flex, because every user we add to that system slows it down further and we may want to use CSS 7, which is the new SigmaSure." Mar 29, 2016: No Updates and waiting for response from WPPL/CM. Mar 30, 2016: WPPL requests to add these users account to Milford's System. Close when complete. April 9th - In order to do this I will need email addresses of all users. | Eric Turner |  | 357
160617 | Manufacturing Engineering | 279003400 | 2016-01-18 | 999038882 | NA | International | Chia Wee Chun | CLOSED |  | 2016-02-03 | The new tool supplied for the Arc QSM is found to be not correct. The depth of Hex Nut socket is too much. it should be about the thickness of the nut but actual depth is about 3 times . This makes it not possible to tighten the nut for the push button | TRF: SWRT-A6QH46. Modified drawimg to prevent this issue from happening again. | Stephen Wright |  | 357
160618 | Manufacturing Engineering | 279003400 | 2016-01-18 | 999026375 | NA | International | Chia Wee Chun | CLOSED |  | 2016-02-03 | The torque driver 999026375, jams on the nut when tightening, making removal difficult after tightening. | TRF: SWRT-A6QH46 | Stephen Wright |  | 357
160619 | Test Engineering | TBD | 2016-01-18 | N/A | N/A | International | Burt Bliss | CLOSED |  | 2016-06-29 | As discussed, it would be appreciated if WPPL could research and create an ASP to define the packaging of the 186015042. A desired date of completion is Feb-2016. Reference Materials: N2971032_ASP_Rev B MI-WAT-5042-007_Rev02 MI-WAT-N2971032-006_Rev 00 | Jan 26, 2016: ASP will be released into SAP next week for approval. Feb 24, 2016: ECR in process. Mar 30, 2016: ECR 920441 in process. June 29, 2016: Complete per Desmond. Closed. | Wee Chun Chia |  | 357
160639 | Test Engineering | N/A | 2016-01-28 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2016-03-30 | Redesign Plan of 16 Testers on Alliance Testers's List, and strategize to set priorities of development. | Feb 04, 2016: First cut from Flex is in. Chia will discuss with Edwin & Desmond to finalize. The main item for this is TE82774 Alliance Column Heater/Cooler Tester at Venture. Feb 24, 2016: WPPL will discuss to finalize. Mar 30, 2016: Priority List is set as of Feb 04, 2016 by Chia. Closed. | Wee Chun Chia |  | 357
160640 | Test Engineering | N/A | 2016-01-28 | TE82620 | N/A | International | Chau Nguyen | CLOSED |  | 2016-02-03 | Redesign of TE82620 2690 Heater/Chiller Tester: Any input and concern related to testing Inquiry about Yearly Quantity manufactured for Production and Spares | Information provided on Jan 29, 2016. Matt is okay to close the request on Feb 03, 2016. | Desmond Ann |  | 357
160641 | Test Engineering | N/A | 2016-01-28 | TE82736 | N/A | International | Chau Nguyen | CLOSED |  | 2016-03-30 | Redesign TE82736 Alliance Final Test System: Most critical concern/problem about the current system List of failures caused by the current Test System (not by the UUT) Specific details of any desired feature to be incorporated into the new Test System to make things easier, quicker, more reliable. | Feb 24, 2016: Mark will review response from Desmond, and follow up with further inquiries as needed. Mar 30, 2016: Closed and Mark can issue new request on demand. | Desmond Ann |  | 357
160642 | Test Engineering | N/A | 2016-01-28 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2016-11-01 | Provide PM Investigate and incorporate any design improvement for performance and maintenance | Feb 22, 2016: Rick and Chau is working on the Filter Housing Design. Mar 29, 2016: Finalize on one design to make prototype. April 26, 2016: Working on final design for filter enclosure. June 28, 2016: Design Meeting was held in the TE Lab, and a different new design approach will be considered for easier removal and re-installation during PM process. July 27, 2016: Complete design by mid-August. Sept 05, 2016: 4 housings are being made in Model Shop and should be completed by the end of Sept. Rick is working on the PM. | Richard Davis | Chau Nguyen | 357
160643 | Test Engineering | N/A | 2016-01-28 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2016-01-29 | Finalize Cal Box for GTS and provide cost to build one box for WPPL to raise CEA to pilot by working with Essco calibrations. | Current TE projects do not require an OESR. They are recorded in the assignee's project list. | Richard Davis | Chau Nguyen | 357
160644 | Test Engineering | N/A | 2016-01-28 | 999029388 | N/A | International | Chau Nguyen | CLOSED |  | 2016-02-01 | Release 3 drawings to change status to RG from Concept in SAP. | ECR 9204426 submitted to QC 02-01-2016. This ECR will promote drawings to RG status (previously at CP status). This will give WPPL visibility to the files (they don't have access to DIR's at CP status). No qualification required for this administrative change. | Michael Cotter |  | 357
160646 | Test Engineering | N/A | 2016-01-28 | 999029388 | N/A | International | Chau Nguyen | CLOSED |  | 2016-09-07 | Inspect existing Test Gauge at Flex against SAP drawings Investigate availability and capacity of Smart Scope at Flex to determine PM frequency | Feb 22, 2016: Mike Cotter has completed releasing the drawings into SAP per ECO 9204426. Feb 24, 2016: Desmond will check Test Gauge against released dwgs. Mar 30, 2016: Desmond already requests CM to inspect the gauge versus drawing. CM requests a second gauge made for daily production usage while sending the existing gauge out for inspection and finalizing on the smart scope being purchased at CM in June. Apr 26, 2016: Need a charge account to dicuss with Mike Cotter to make a Second Gauge. Paul Lussier reviewed and agreed that the vision system should be quite capable of checking required features. June 29, 2016: Discuss with Mike C. to make a new gauge by using charge account WP21 615150, so WPPL can send the existing gauge out for inspection. July 15, 2016: Mike placed order with Model Shop, and should be done by the second week of August 2016. July 27, 2016: Gauge completed and to be inspected dimensionally. Sept 06, 2016: Test Gauge was shipped. Closed. | Desmond Ann |  | 357
160655 | Test Engineering | NA | 2016-02-04 | TE82705 | 00 | International | Desmond Ann | CLOSED |  | 2016-03-30 | 1500 Series Pump Final Test Station TE82705 ancillary equipment Electronics Weighing Balance 90273,4th digit have 3 missing segment . 2. UUT Instrument Final Test Record is fully dependent to weighing balance reading by tester software. 3. Need replacement unit to avoid line down in the near future. Request for a replacement AND HR-200 balance. WPPL got a quote last year and it cost SGD $5k in Singapore, I believe it will be cheaper in US. Please use account WP21 615150. | Feb 22, 2016: Paul Dechenes received an US quote to be about 50% less in price. Mark will review to buy and ship from US to WPPL. Feb 24, 2016: Desmond already provided charge number and order of the balance will be placed after Mark's approval. March 29, 2016: Paul Deschenes received the balance and shipped to WPPL. Mar 30, 2016: TRF number TFOS-A8GNY5. WPPL will notify to close when receive the scale. | Paul Deschenes | Chau Nguyen | 357
160665 | Manufacturing Engineering | 279001150 | 2016-02-15 | 9120002 | 00 | International | Desmond Ann | CLOSED |  | 2016-06-29 | Flextronics reported that Detector Clean Room need replacement for broken tooling 912002 Reticule Plate. It is currently affecting below model Optic Bench and Units production of the following PN as well as SPARE shipment. 279001150 279002656 WAT081147 279003288 | Feb 16, 2016 - Steve Wright: "The replacement part for Fixture 9120000-2 has been shipped see TRF SWRT-A76M-M. Please proved a valid account number so I can order a few spares, they cost about $150.00 each. ASP's 279001150, 279002656, WAT081147, 279003288 require the use of this tool and need to be updated. They call out for the wrong tool number in the alignment process, 9120002 is listed instead of 9120000-2. Please go thru the ASP's above to see if any other correction need to be made. I can make the corrections once they have been identified." Feb 24, 2016: By the end of the week, Chia will provide charge account and the number of spares needed so Steve can make. Mar 30, 2016: Steve can use account WP21 615150, and update WPPL on the spare parts and cost being charge to the account. | Stephen Wright |  | 357
160675 | Test Engineering | 735001315 | 2016-02-25 | TE82987 | 1 | International | Hong Cheng/Chia | CLOSED |  | 2016-03-30 | Failure Symptom: Unit under Testing not able to obtain IP address from test station. Suspect caused by a NEW Ethernet Card P/N 210000852 The Tester is left with a UUT. Information below. Model: 186015007 Unit Serial Number: B16UPM369G Test station IP address: 10.249.40.22 | Mar 29, 2016: This is NOT a tester issue. Currently running under TMD to fix design firmware issue. Closed | Matt Brule |  | 357
160683 | Test Engineering | NA | 2016-02-29 | TE82882 | 13/14 | International | Joel/Chia | CLOSED |  | 2016-03-30 | The 2 new testers were setup but does not seem to be working when we use them to test 2998s. (It is noted that they are configured for 2998s) It hanged and does not proceed. The IP address are below. TE82882-13 – 10.249.26.84 TE82882-14 - 10.249.26.85 Note: The current tester uses a different USB to RS232 convertor. (Old= Edgeport but current StartTech). Is the addressing and port connection the same as Edgeport? | Mar 29, 2016: Needed additional configuration and correct connections. Closed. | Eric Turner |  | 357
160687 | Test Engineering | NA | 2016-03-08 | TE82503 | 04 | International | Joel/Chia | CLOSED |  | 2016-03-11 | The Keypad for the Injector Tester is worn out. Need replacement. 1) Please help to supply one for replacement. 2) Provide a P/N or specification for future sourcing | I sent 2 replacement key pads to Desmond they will give to Flextronics for repair. this is just a key pad replacement just like changing the keyboard on a computer it will nor effect the tester. | Paul Deschenes |  | 357
160690 | Test Engineering | NA | 2016-03-14 | NA | NA | International | Chia | CLOSED |  | 2016-04-27 | Kindly Review the GROUND LUG and Hardware Stack Configurations for the AcQuity models: Below are some of the examples of different stack configuration that is confusing (See details below). Please define a standard stack-up configuration. I believe this has a safety/regulatory implications. ---------------------------------------------------------------------------------------------------------------- ECR# 9200361 add qty 1 piece of 441001420 Assy Cable Chassis Ground, 10” to (1) 186015043, ASP 279002113, Assy, Column Manager, Acquity, Rev: G. (2) N2971034, ASP 279003538, Assy, Column Manager, Altus A-30, Rev: D Both ASP was update by Mr Bersani. The Cabling stack up sequence is on (1) Page 24 of 29 of ASP 279002113. (2) Page 24 of 30 of ASP 279003538. The stack up sequence is 1 piece of Lock Washer, 3 pieces of ground lugs, 1 pieces of KEPS Nut. Attach document for your reference. This sequence was totally different from the stack up sequence in Acquity CM-AUX and UPC2 CM-AUX both model. (3) 186015049, CM-AUX, ASP 279002135, Assy Proc, Acquity CM-AUX. (4) 186015039, UPC2 CM-AUX, ASP 279002726, Assy Proc, UPC2 CM-AUX. In both ASP the stack up sequence is 1 piece of Ground Lug, 1 Lock Washer, 1 piece of Ground Lug, 1 piece of Lock Washer, 1 piece of Ground Lug and 1 piece of KEPS Nut. Attach document for your reference. (5) We also did a review on 186015038, UPC2 CM-A and notice that there are errors. On page 7 of 29 in ASP 279002717, it mention to use 2 pieces of Lock Washer (P/N: WAT410450) and 1 piece of KEPS Nut (P/N: WAT251620) to assemble valve actuator ground cable (441001420) to chassis ground stud. On page 21 of 29, it mention to use 1 pieces of Lock Washer (P/N: WAT410450) and 1 piece of KEPS Nut (P/N: WAT251620) to assemble 2 ground lugs to the same chassis ground stud. This is impossible as the height of the chassis ground stud is not able to tighten the 2 pieces of KEPS Nut together. The ASP have to revise to change the stack up sequence to remove the extra 1 piece of Lock Washer and 1 piece of KEPS stud mention. Attach document for your reference. Attached document is the history on the qty of Washer and Nut before and after adding extra ground cable. The difference in stack up sequence will cause confusion during assembly process. We would like to suggest standardizing the Stack up sequence for all the Column Managers models mention above in the following sequence: 1 piece of Ground Lug, 1 Lock Washer, 1 piece of Ground Lug, 1 piece of Lock Washer, 1 piece of Ground Lug and 1 piece of KEPS Nut. | Mar 24, 2016: Randy Weaner: "The first connection of ground conductor coming from the AC mains is called the Protective Conductor Terminal. Below is the excerpt from IEC 61010 of the requirements. In particular, (g) requires that if this terminal, (stud) is used for other functional ground wires, the AC protective conductor is on the bottom and secured independently of the others grounds in a reliable method. The standard does not require multiple nuts to secure each functional ground above the protective conductor. The Protective Conductor Terminal is required to be marked as shown. The reason for all this is to alert any person servicing the equipment this connection should not be disconnected. If disconnection is required, care must be taken not to become the ground wire in a energized state. Please let me know if you have any other questions. Randy." Mar 30, 2016: Chia will review the actions taken and notify Milford to close. | Matt Brule |  | 357
160691 | Test Engineering | 0 | 2016-03-14 | 279002260 | 00 | International | Desmond Ann | CLOSED |  | 2016-03-15 | We have received feedback that the XTS Assy which is used for CPSO testing is down due to Latch Home Sensor Error. We have tested it on machine, and as well as offline tester. The latch home indicator is always “ON” even it moves to away position. We suspect that the PCBA 210000552 is faulty. Please help to get a replacement asap, as we currently only left one XTS for CPSO testing. | Ship them 2 pcba's 210000552. Tony can close this OESR when they ship Chau, Rick, I just shipped 2 boards PN :210000552 Tray Latch Sensor to Desmond Ann, attention under TRF Number TFOS-A83JR9, OESR #160691 is Closed There are no qualification requirements, this is just fo fill a request | Richard Davis | Tony Fontes | 357
160715 | Manufacturing Engineering | WAT118140 | 2016-03-28 | NA | NA | International | Desmond Ann | CLOSED |  | 2016-06-29 | With reference to attached photo, the crimping tool to attached ferrule PN WAT022330 reached wear and tear thus we need new head to replace the existing unit ( 2 pcs left and right ) with below dimension: L 24.58mm X W 4.19 mm H 5.72 mm. This is causing a line down in production, we will need the replacement urgently and will appreciate if you can forward us the drawing for this tool and crimp head. | Mar 29, 2016: I (Steve) am glad you were able to locate the parts, I am making a spreadsheet so we can better track and handle parts replacement. Please provide me with a valid account number so I can process the purchase order for replacement crimp fixture parts. Also the vendor has agreed to provide you the drawing so that you can have your local supplier to make future replacement parts. Best Regards. Mar 30, 20106: Close after Steve provides to WPPL drawings from the US Supplier. | Stephen Wright |  | 357
160724 | Test Engineering | N/A | 2016-03-31 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2016-05-18 | Provide compiled list(s) of all spares for Test Equipment and Tooling currently kept at WPPL and each CM. | List of Spares is available and maintained by WPPL with minimum quantity of 1 each. Here are the updates on 4 parts with 0 spares : 1) DVI to VGA Converter: This item can be purchased from an electronic store locally. 2) WAT046-01: These yellow plastic ferrules are available to purchase at anytime. I (Chau) have a bag of 16 pieces, which I can ship upon request. 3) 99-0055-02 Pen-Ray PS-1 Power Supply: It's not available in stock, but it can be purchased directly from supplier. 4) 999028048: WPPL will raise another OESR when require Milford's help to fabricate this tool. | Desmond Ann | Wee Chun Chia | 357
160725 | Test Engineering | NIL | 2016-04-01 | NIL | NIL | International | Desmond Ann | CLOSED |  | 2016-04-01 | Bill Irons have commented that the CEI is offline and Chau have requested to create an OESR on this issue. Attached is the 20 instances that have occurred since 18Nov2013 where we have to reset the sonicwall. We will like to request Milford to investigate what need to be done to resolve the issue. | CEI-Singapore (10.249.55.x) is the only site experiencing this intermittent service outage problem. We do not have this problem with any of the other four (4) sites in Singapore & Indonesia. The network equipment ( Sonic Wall & Switch ) was selected and configured by Waters BT department. 10.249.55.203 Switch 10.249.55.254 Sonic Firewall When Waters BT ( Raymond Danish ) last reviewed the equipment status, he explained that the problem was most likely caused by the local ISP using DHCP. ( see attached email on Pg.02 ) Email chain attached on Pg.03 | Bill Irons |  | 357
150001 | Manufacturing Engineering | - | 2015-01-05 | 119652 | - | International | Chia Wee | CLOSED |  | 2015-09-29 | Flex have received ECO9191850 which updated and revised the tool 119652 on page 2 of the ASP. Flex have the verison that is in Rev K ASP BUT they need the rev V version to be in compliance. Please help to supply the current approved version. | Joe has moved to NPI group and already discussed with Ken Voisine about the request. WPPL needs an updated tool. Burt will talk to Ken to resolve. Ken will talk with Joe to provide updates. Steve has taken over this request. Tool to be done in 2 weeks to match SAP document. WPPL is still waiting for the Tool. Stephen W. on 15 Sept 2015: TRF SWRT-A2DFVE for Pin bearing Press Shipment. Question: Should WPPL should hold onto the old version press as a backup until full confidence of the newer one is established for a month or so, and then upgrade the older version locally or ship it back to Milford for upgrade and ship it back to WPPL if needed? Sept 29, 2015: Desmond will ship the adapter from the old press back to Milford for Steve to upgrade, and ship back to WPPL. This was originally assigned to Dick Bersani/Joe Labenski/Steve Wright. Dick and Joe have been removed. | Stephen Wright |  | 357
150002 | Test Engineering | - | 2015-01-07 | TE82736 | - | International | Desmond Ann | CLOSED |  | 2015-05-18 | To setup the 5th tester to increase the final testing capacity of Alliance and support the private label project. | On Jan 21, 2015, Mike Cotter has sent items from the 'Equipment Required' section of Test Procedure WAT271129. One balance in Alliance Line will not be needed for much longer (next version/automated Final Test) per Mike. Have materials been procured for this test expansion yet and anything else Milford can provide to get it done? Milford will purchase items available in Milford and ship to WPPL from the list provided by WPPL. WPPL will coordinate with Milford to set up the system when everything procured. All parts orderd by WPPL. Waiting for PC from Mike Cotter. April 17, 2015: TRF MCOR-9VEHUY for PC, Prim/Purge Module, Carousel Assy Kit, Tray Solvent Assy, and 9 pin Cable. | Steve Hunt | Michael Cotter | 357
150008 | Manufacturing Engineering | - | 2015-01-09 | - | - | International | Desmond Ann | CLOSED |  | 2015-09-29 | There is a confusion in the Tool numbers and we will require to standadise the tool numbers. More details in OESR word doc. | WPPL and Ken shall discuss to come up with a solution. Propose to use 999 as the prefix for all tools. More discussion will be done when WPPL comes to Milford for Waters Conference in June 2015. Jaaffar will discuss with Ken, so Ken can provide a plan and timeline. Sept 29, 2015: Separate OESRs will be issued on demand. Closed. | Ken Voisine |  | 357
150015 | Test Engineering | N/A | 2015-01-19 | 405006712 | N/A | International | Desmond Ann | CLOSED |  | 2015-02-16 | Need Needles use in both Acquity CPSM RTZ Assembly Tester | Model Shop is working on it and expected delivery in early next week. Need shipping address and recipient's info ready. 2 blank needles will be shipped to WPPL when ready. TRK# 509095146267 to TRF CNGN-9TNR | Chau Nguyen |  | 357
150019 | Test Engineering | - | 2015-01-20 | Certificate printer at GES Venture | - | International | Chia Wee | CLOSED |  | 2015-02-02 | Certificate of performance: the alignment seems to be a little low. | GES Venture - Waters Certificate of Functional Performance Program: Administrative action: I have made a negative TOP (vertical) adjustment to the three (3) Certificate print fields = 300/1,440 Twips ( 1,440 Twips = 1 inch ) MODEL SERIAL# ( Human Readable ) SERIAL# ( Bar Code 128B ) Have received confirmation of satisfactory results. ( see attachment ) GES Venture - Waters Certificate of Functional Performance Program: Administrative action: I have made a negative TOP (vertical) adjustment to the three (3) Certificate print fields = 300/1,440 Twips ( 1,440 Twips = 1 inch ) MODEL SERIAL# ( Human Readable ) SERIAL# ( Bar Code 128B ) Have received confirmation of satisfactory results. ( see attachment ) | Bill Irons |  | 357
150023 | Test Engineering | N/A | 2015-01-22 | TE83006 | 2 | International | Mark Jussaume | CLOSED |  | 2015-02-04 | Verify TE83006 | Desmond has confirmed 1 at Flex on Jan 22, 2015. Mark completed his works. Closed. | Desmond Ann |  | 357
150042 | Test Engineering | Various | 2015-02-10 | TE82526 | NA | International | Desmond Ann | CLOSED |  | 2016-07-11 | WPPL is conducting an exercise with Flextronics to have the ASPs and TSPs reflecting to the production practice. Flextronics propose to update the TSPs with more information as per redlined attached. | We can incorporate the changes the next time the document is revised. Let us know if the need becomes more urgent. Works will be done on case-by-case basis and assigned to ME to take care of ASPs and TE to take care of TSPs. WPPL suggest to give TE ECRs for approval. Going forward, changes on ASP and TSP will be submitted on separate OESR(s) instead of bundling into an open ended request. In this OESR, Mike, Mark and Matt will review and ECR the 5 procedures. Mike C.: WAT210700 and WAT271129 have been updated. Sept 29, 2015: Waiting for TSP updates from Mike (735000174), Mark (WAT047159) and Matt (WAT242102). Jan 5, 2016: Mark: TSP WAT047159 has been updated to new format and changes form Joel/Desmond included, and ECR will be submitted. Matt: ECR 9203715 for WAT242102 submitted. Mike: 735000174 was last updated in 2013, so clarification is required on the changes requested to this document; WPPL confirmed in monthly meeting that 735000174 to be ECRed to include changes per request on Feb 12, 2015. 735000174 updated via ECR 9204319 effective 07/11/2016. Spreadsheet (TE82660) was qualified as part of this revision. See TE0082660QX for qualification report. All other changes on this request were strictly administrative/documentation and required no qualification. | Michael Cotter | Mark Jussaume | 357
150050 | Test Engineering | - | 2015-02-16 | TE82586-31 / 999029699 | - | International | Chia Wee | CLOSED |  | 2015-02-22 | PCB intermittent lost connection, replace connector, problem remain. Need replacement. | Made replacement PCBS | Steve Hunt | Eric Turner | 357
150051 | Manufacturing Engineering | - | 2015-02-16 | 999029189 | - | International | Chia Wee | CLOSED |  | 2015-03-23 | The change resulted in the part getting jammed in the fixture 999029189. Please see picture below. 3083 engages the fixture on all 4 sides, but the new part 9792 engages the fixture on the small “tab” feature (circled in red). Please help review and provide an upgrade fixture 999029189. (See Attached for details) | Use this account WP21 615150. TRF: SWRT-9UNJT5 on March 16, 2015 by Steve W. Chia received on 24 March 2015. | Stephen Wright |  | 357
150055 | Manufacturing Engineering | - | 2015-02-22 | 999029418 | - | International | Chia Wee | CLOSED |  | 2015-03-02 | Flextronics alerted that the Flowcell alignment fixture is broken and the production for 186002998, 186299800 and all 2998 Flow Cell spares are currently down since we do not have any spares for this fixture. Please help to advise if you have any spares and when can Milford help to send the replacement to us. Please see attached photo on the broken fixture. | Use this account WP21 615150. Replaced and the fixture is up. | Stephen Wright |  | 357
150065 | Test Engineering | - | 2015-03-06 | 289002219 | - | International | Desmond Ann | CLOSED |  | 2015-03-10 | Shunt 289002219 cannot passed test. The light cannot go through the fibre. Please help send us a replacement. See attachment for the failure mode.. | Desmond have received the shunt and shipped out the faulty shunt attention to Chau. Tracking number is 487737198344. | Chau Nguyen | Eric Turner | 357
150067 | Test Engineering | - | 2015-03-11 | TE83064 | - | International | Rick Davis | CLOSED |  | 2015-05-19 | In order to better understand root cause of test failures Venture was instructed to charcterize failures using the Failure Analysis menu's. I checked FPY from Feb 23 to Mar 11. There are 16 Dynamic Leak Test failures and all are characterized as No Fault Found/Retest. When looking at the test reports in XML format, there are very good reasons for failure including an Abort. The operator can see these test results on the Diagnostic page. Why are they not following instructions? Attaching sample test reports. | Chia has been working with CMs to ensure operators at CMs get proper training to look for the root cause and write it down for Milford to troubleshoot. | Desmond Ann | Chia Wee | 357
150070 | Test Engineering | - | 2015-03-13 | TE82743 | -3 | International | Desmond Ann | CLOSED |  | 2015-03-19 | Test failing when initializing and configuring the Electronic Load module 3. Module confirmed by swapping the affected module with module 4 and the failure follows. Need replacement | Rick sent a replacement Electronic Load Module. Waiting for okay from GES and will close it out. | Richard Davis |  | 357
150073 | Test Engineering | - | 2015-03-13 | TE82743 | -3 | International | Rick Davis | CLOSED |  | 2015-04-28 | Racal system is failing voltage measurements using the VXI DMM. Fails on multiple boards and on multiple fixtures. Tried unpluging the wires to the VXI DMM and measure voltage from the banna plugs using handheld DMM, 0 voltage. | Problem seems to have resolved itself. No problem found. | Edward Jackson |  | 357
150076 | Test Engineering | - | 2015-03-19 | TE83090 | - | International | Chia Wee | CLOSED |  | 2016-01-11 | Please help to update TSP WAT038172 Rev 7 for Column Heater Module to include Part Number: WAT091727. The purpose is to align with the Instrument Final Test Record that eventually we used as calibration report to login to calibration system at Flextronics. | Procedure update to support PM process will complete in early April 2015 by Matt. Matt should complete by the end of June 2015. Matt will provide latest updates since it already past June 2015 as originally planned. Matt on Sept 29, 2015: This has been a low priority, will try to complete in October. Jan 05, 2016: ECR 9202766 submitted on Nov 17, 2015, rejected by quality on Dec 16, 2015, ECR 9202766 just resubmitted on Jan 05, 2015. ECR 9202766 Complete | Matt Brule |  | 357
150077 | Test Engineering | - | 2015-03-19 | Auto Bagger | - | International | Desmond Ann | CLOSED |  | 2015-06-24 | ichport autobagger Sharp SX Serial Number SX-10I-1127-33 is having printing issue. Please help to advise what actions to be done. See attached on the printout sample. | Either a defective ribbon or dirty printer head issue. WPPL to get Datamax service representative to help, and the printhead requires 4 weeks to be procured. The existing Bagger could be at end of life, so WPPL will explore other Baggers at CMs to source locally, and submit to Milford for review and approval. Bill will help sourcing a printhead for the existing Bagger, and also define specs for a new Bagger. CLOSED. | Bill Irons |  | 357
150097 | Test Engineering | - | 2015-03-30 | TE83064 | all | International | Rick Davis | CLOSED |  | 2015-08-26 | Contamination: GES complaining of Columns not lasting very long. They have installed Pre-Column Filters to increase the life of their Columns. Analysis of a Frit and Vial Cap has shown particles of Stainless Steel. The pre-column filter can cause band spreading, peak distortion, carryover and can add to the peak retention time. It is not part of the equipment list and is not part of the IQ. WPPL needs to open an investigation in order to find the source of the contamination. The Pre-Column Filters need to be removed in order for their systems to be compliant. | (Reference on OESR# 140450) CLOSED. | Desmond Ann | Desmond Ann / (Chia Wee) | 357
150132 | Test Engineering |  | 2015-05-11 | TE83084, 999040126 |  | International | Desmond Ann | CLOSED |  | 2015-11-02 | ECO 9192881 is changing the plug tool from 999034958 to the new Plug Tool 999040126 and add in an additional TE83084 FTN BEZEL TESTER. Please help to consigned the items to us. | Plug Tool 999040126 was created by Fred S. who will come back to the office on 26-May-2015 to get it made. Steve Hunt is working on the TE83084 FTN Bezel Tester. Matt Franklin already TRF the Plug tool 999034958 on TRF MFRN-9XZR96WPPL. Steve Hunt will provide updates on Bezel Tester. Received and closed. | Steve Hunt |  | 357
150140 | Test Engineering | - | 2015-05-18 | TE82969 | 1 | International | Desmond Ann | CLOSED |  | 2015-06-24 | 1. RS232 Not Working. 2. TE82969-1 PDA Linearity Tester Not able to read Serial Number of the UUT.refer to email 13-May-15 3. Installed USB-to-Serial Port Cable - No succcess to recover the RS232 Communication to UUT 15-May-15 4. Transferred PDA Linearity Tester HDD to another good CPU: 4.1 After transfer the HDD of PDA Linearity Tester to another CPU, its being read as E:\ Drive. 4.2 Required to reconfigure the PDA Linearity Tester. 5. In-case not able to recoverthis LACE Workstation, pls. send a preconfigured PDA Linearity Tester replacement Workstation. | 1) Each PC provided to CMs is specifically configured and controlled in the Master PC List so Milford can provide support and troubleshoot remotely. Change to a different PC is the last option and need to consult with Milford before doing so, so the new system will be verified to be compatible or not, and also update the Master List for tracking purposes. 2) Please explain what went wrong with original PC and its COM Ports, or the serial cable was connected to the wrong port. Bad montherboard caused COM failure. 3) It might already be too far down the road to reassemble the original system, but we might have to do so unless the new system is compatible. Please put the new system online so Eric can check it and see what can be done. 4) If the new system works, Milford will need to know info for each IP Address/Computer moved/swapped: The Platform The Amount of RAM The Operating System (including Service Pack) The Symantec Antivirus software version. Eric and Joel are able to get the system up and running. Monitoring will continue in the next few days. Specs of new PC will be needed to update the PC Master Database. CLOSED. | Eric Turner |  | 357
150143 | Test Engineering | - | 2015-05-19 | Autobagger | - | International | Chia Wee | CLOSED |  | 2015-05-19 | It seem that Richport was requested to Critical Clean 405020117RevA. The autobagger database do not include this part number. Please email an updated database and I can add it into the Autobagger. | No action taken. The material 405020117 Rev A is NOT released to production ( US31 Plant ), and therefore can not be included in the Waters Production Labeling Database. Waters RDE email consultation attachment (Paul Linderson: New Part and being evaluated). | Bill Irons |  | 357
150144 | Test Engineering | - | 2015-05-19 | TE82972 | 1 | International | Desmond Ann | CLOSED |  | 2015-05-19 | 1. Using the station TE82972-1 to program/configure Part Number 186015032 ACQUITY UPLC PDA Detector but showing on Instrument Final Test Record for Product : Altus A-30 PDA Detector. 2. Firmware Revision 1.60.6169 is following as per ECO 9183698. 3. Affected Serial Numbers : 9 units has been released for Shipment as Firmware revision is correct as per ECO 9183698. | Related to a Text File which contains model information that was out of order. Should already be fixed so please verify and confirm back to Milford. TE82972-1 is now able to run the Auto downloading Station and now able to Program PN: 186015032 to Product Description : PDA Detector. | Eric Turner |  | 357
150145 | Manufacturing Engineering | - | 2015-05-19 | See List | - | International | Desmond Ann | CLOSED |  | 2015-06-24 | Flextronics is requesting the below tools due to the current tools have been broken: 1) 119652 requested under OESR# 150001 dated 5 Jan 2015, still outstanding. 2) 999028048 3) 999032368 4) 9026700 | The duplicate tools were ordered from the shop Today 19May2015. Replacement tools 9026700, 999028048 and 999025111 are being shipped to WPPL on TRF #SWRT-9WNMXD. Fixture 119652, we do not have a spare, the duplicate will be ready in 2 weeks. CLOSED. | Stephen Wright |  | 357
150150 | Test Engineering | N/A | 2015-05-26 | TE82760 | N/A | International | Chia Wee | CLOSED |  | 2016-03-29 | FLUORESCENCE DETECTOR FINAL TEST STATION. Revise TSP 735000615: (1): Add test step B to H in Test Operation 1.5 . (2): Add Test Step 1.6 and 1.7 (3): Add test operation 5 to 7.4 . (4) Add photos | Eric will provide latest updates. Nov 02, 2015: Edwin asked to help putting this through as soon as possible. Jan 5, 2016: In process of making the update and writing the ECR by Tyler Modelski. Jan 11, 2016: ECR submitted to QC. Feb 22, 2016: ECR 9203820 for OESR 150150 (updates to TSP 735000615) is now "Approved Pending Execution". Mar 29, 2016: Approved and Closed. | Tyler Modelski |  | 357
150151 | Test Engineering | - | 2015-05-26 | TE83031 | - | International | Chia Wee | CLOSED |  | 2016-01-12 | 2475 OPTICS BENCH ALIGNMENT STATION. Revise TSP 735000629: Revise Operation 3.1 and add photo. | Eric will provide latest updates. Nov 02, 2015: Edwin asked to help putting this through as soon as possible. Jan 5, 2016: ECR 9203686 by Tyler Modelski. Jan 12, 2016: Per Chia from WPPL, the TSP ECO was processed. Closed. | Tyler Modelski |  | 357
150152 | Test Engineering | - | 2015-05-26 | TE82757 | - | International | Chia Wee | CLOSED |  | 2016-01-12 | 2475/FLR FLUORESCENCE DETECTOR FIRMWARE AND SERIAL NUMBER DOWNLOAD STATION. Revise TSP 735000589: Add photos on page 3, 4 and 5. | Eric will provide latest updates. Nov 02, 2015: Edwin asked to help putting this through as soon as possible. Jan 5, 2016: ECR 9203686 by Tyler Modelski. Jan 12, 2016: Per Chia from WPPL, the TSP ECO was processed. Closed. | Tyler Modelski |  | 357
150153 | Test Engineering | - | 2015-05-26 | TE82774 | - | International | Desmond Ann | CLOSED |  | 2015-11-05 | ALLIANCE COLUMN HEATER/COOLER. Revise TSP 735000698: Add photo on page 5 & 7. | TSP 735000698 is the one to be worked on, not 735000798. Waiting for updates from Ed. Ed J. on August 25, 2015: Other changes are being made on the same TSP, so changes requested by WPPL will be incorporated into a single ECR to be released in Q3/Q4. ECR 9200681 approved. | Edward Jackson |  | 357
150154 | Test Engineering | - | 2015-05-26 | TE83013 | - | International | Desmond Ann | CLOSED |  | 2015-09-11 | HIGH TEMP COLUMN HEATER TEST STATION. Revise TSP 735000615: 1: Add Assy Trough Clip and gasket 289010457 in the equipment list. 2: Update Test Specifications table. 3: Add photos for Test Operation 3.4 | ECR 9200654 submitted and notified on 21 August 2015. Matt will provide latest updates with a completion date if possible. CLOSED. | Matt Brule |  | 357
150155 | Test Engineering |  | 2015-05-26 | 9120000-3 |  | International | Desmond Ann | CLOSED |  | 2015-06-24 | Flextronics have reported having issue on Techniquip Fiber Optic Illuminator Model R150 MV2 from Capra Optical. The light coming out from the Fiber Optic cable is very dim, which makes it difficult to perform the alignment for optic bench. The affected optic bench will be 2489 and TUV model. We will like to request for the replacement cable as seen in the picture attached. | 9120000-3 can not be found on TE List. June 02, 2015: Desmond to place PO to get cable shipped directly from Supplier Capra Optical, Inc in Natick MA, USA per Quotation # WA0526AL. CLOSED. | Paul Deschenes | Chau Nguyen | 357
150156 | Test Engineering | N/A | 2015-05-26 | 210000389 Aquity CM PCBA | 1 | International | Chia Wee | CLOSED |  | 2016-05-13 | Quote"…GES had reject these boards to us on “I-button” or “Heater/Cooler test failed graph” failures at the instrument test. As this 210000389 board under gone ICT test only at CEI, we were unable to verify and diagnose the failure. The boards were thoroughly check, visual and resistance check about the suspected circuitries, but could not determine the cause of failure…." (See Attached email for details) | PCBs ordered, Matt will update when received and troubleshot. Matt on Sept 29, 2015: Have the PCBs, will get installed and tested soon, believe this to be a final test issue. Jan 5, 2016: Ongoing investigation, hope to complete soon with final test upgrade. (150156, 150165), are both an issue with CMA test that needs to be resolved in Asia, this will came with the update to Sigmasure 7, the two OESRs could be combined. M Brule May 12, 2016 OESR Closed May 13, 2016 M. Cotter | Matt Brule |  | 357
150157 | Test Engineering | - | 2015-05-26 | TE82882 | 8 | International | Desmond Ann | CLOSED |  | 2015-06-24 | On 25/5/15,one of the PDA test station TE 82882-8 kept restarting and all of sudden it hangs on blue screen. It was suspected that the hard disk has corrupted. So we managed to reformat and recover with Lace 32 basic setup image but we noticed that the OS is Windows XP instead of Windows 7. We would like to request Milford’s help to reconfigure the test station as per below details: PDA Final Test station : TE82882-8 IP address : 10.249.26.79 Test software : TE82882 rev:3.2 Document number : 735001613 TSP Rev 16 | Closed | Michael Cotter | Eric Turner | 357
150159 | Test Engineering | N/A | 2015-05-27 | PDA Shunts | N/A | International | Desmond Ann | CLOSED |  | 2016-01-26 | Flextronics have reported 7 PDA shunts have shown degradation and will need replacement. Please help to send us the replacement shunts. | CEA from WPPL to build 7 NEW Shunts was approved. New Shunts left over from prior projects or from Refurb are free. 6 Medium Shunts 289002226 and 2 Large Shunt 289002149 can be shipped first. 2 Small Shunts 289002227, and 4 Large Shunts 289006386 are to be built. The supplier in US does not have any main plate P/N of 405012460 to make 289006386 and will not take order of small quantity, so Chau currently work with Model Shop to get them made. How urgent are the 4 Large Shunts 289006386 needed (If able to wait 6 more weeks, we can get 6 plates for the price of 4 for future usage by using more automatic machine) ??? Desmond will ship any broken shunts back to Milford for repairs, and also provide maked-up prints to redesign mounting holes on any shunt as needed. Sept 29, 2015: 6 Medium Shunts 289002226 and 2 Large Shunt 289002149 will be shipped first. WPPL can wait 6 more weeks to get 6 base plates for the price of 4. Nov 02, 2015: 2 Small, 6 Medium, and 2 Large were shipped and received at WPPL. 4 UPC2 Flow Cell to be shipped in December. Jan 06, 2016: Chau has received 4 UPC2 Flow Cells to ship out ASAP and also clear with Finance to close it out. Jan 11, 2016: TRF CNGN A63LRT. When WPPL confirms of shippment received, Finance will cross charge into the CEA. Jan 26, 2016: Received and Closed. | Chau Nguyen | Eric Turner | 357
150163 | Test Engineering | - | 2015-06-02 | PDA rough aligment fixture | - | International | Desmond Ann | CLOSED |  | 2015-06-08 | The cable for the lamp in the PDA rough aligment fixture is broken. Please help to provide spares for the lamp assembly (WAT097323) | It looks burnt, look into the cause if needed. TRF: CNGN-9XAQ3X | Eric Turner | Chau Nguyen | 357
150167 | Test Engineering | - | 2015-06-11 | TE82988 | 3 | International | Chia Wee | CLOSED |  | 2015-10-05 | Tester down when start to test Model 493, Tried to used self test also unable to run the test seems show error ..(tried to run other model also same error ) We attach the print screen for your view , The system online now can access remote. we plug in the Self test fixture for verified the test fixture . | Geotest Chassis power supply failure. Replaced. WPPL still waits for P/N of Custom Power Supply from Milford. Fixture fixed. WPPL still waits for P/N of Custom Power Supply from Milford. June 24, 2015 should be the date to close it out unless WPPL needs further assistance from Milford. Rick D. on August 25, 2015: The OESR is for the GTS, and it is up and running. Rick are recommending the issues with the PXI Chassis and PC to be taken up in separate OESR's so they can be assigned to the appropriate engineers. August 26, 2015: Chia shipped PXI to CA for repair, and also working with Michael Cotter on PC. Per Chia on Oct 5th, 2015: The GeoTest chassis is back from Marvin Test. | Richard Davis |  | 357
150168 | Manufacturing Engineering | - | 2015-06-11 | 999028262 | - | International | Desmond Ann | CLOSED |  | 2015-09-14 | Kindly be informed that the screw used in the fixture already found wear off. Need to replace the screw urgently. Appreciate if you can give us new screw 8pcs Can you also provide a longer handle align key? This will help us to do the job better. Affected parts: 405006788, 495003295 & 405002568 | TRF SWRT-9ZGK45 received. Close. | Stephen Wright |  | 357
150172 | Test Engineering | - | 2015-06-19 | TE82520 | 1 | International | Chau Nguyen | CLOSED |  | 2018-04-02 | WITHOUT making any adjustment/changes in any step to either the Tester TE82520-1 and Golden CCD, go through the PM 735002998 - Rev B - Draft 5 and TSP 735000431 - Rev G to provide feedback on the PM so it can be edited and released since Milford doesn't have the Tester physically. Raise an OESR if discover any problems on the Tester during the evaluation. | Feedback should be provided before July 2015. August 26, 2015: Chau will provide CMM information from Machine Shop so Desmond can try to source similar machine locally. Chau will make 2 new sets (Master Carousel and Carrier) to ship to WPPL along with CMM reports, so Desmond can determine local capability to measure tight tolerance specified in PM because current measurement at WPPL and local source didn't meet the specs. Reference for 2014 OESR 262. Sept 22, 2015: 2 new sets (Master Carousel and Carrier) are being fabricated in Model Shop. Sept 22, 2015: Chau will provide total cost to Chia to raise CEA. Nov 02, 2015: 2 sets made and to be shipped today. WPPL will help reserve CMM for 2 days at CM per planned travel schedule of Paul and Chau in 3rd week of Nov. Jan 06, 2016: Chau to work with Steve Wright and Refurb to upgrade the origininal CCD Tester as a 3rd Tester to realign the misaligned Master CCD shipped back from Flex, and will finalize the 2 Master Sets of Carousel and Carrier to complete PM. Jan 22, 2016: Go through BOMs to determine cost to upgrade the origininal Tester to TE82520-3, and will discuss with Refurb/WPPL to raise CEA. April 24, 2017: To complete settings for 2 spare sets of Master Carrier and Carousel , and finalize PM to release into SAP. Oct 24, 2017: 2 spare sets of Master Carrier and Carousel were shipped. PM for CCD Tester and for Master CCD and 2 Mechanical Gauges are being reviewed by WPPL/CM to release this October to close out this project. 02 April 2018: ECR 9221672 and 9221682 were submitted and approved. No changes to software, hardware, and process, so no qualification required. | Desmond Ann | Chau Nguyen | 357
150173 | Test Engineering | - | 2015-06-22 | TE82577 | - | International | Chia Wee | CLOSED |  | 2015-11-19 | FCT test failed for W1081060. The tester failed at keypad test. | Matt is working on it now, and not sure of problem yet. Chia will provide additional info by the end of the week. July 27, 2015: Tester to be shipped back. Matt B. on August 25, 2015: Actively working on this one and communicating with Chia, we are in a line down situation, I am aware of the urgency, and am giving it as much priority as I can. Edwin will discuss with Planing to determine the need for a new and upgraded Tester. Matt on Sept 29, 2015: Sent an alternative test method this week, such that product can be tested. Will still spend some time to try to get the original tester going as it is not as cumbersome to use. Sept 29, 2015: Chia will get 2487 board to test and notify to close. Nov 02, 2015: No issues so far on using the alternative test method. Waiting for further information from Matt. Jan 5, 2016: Replacement tester sent and they are up and running, just need to complete ECR for procedure to close it (!?!). | Matt Brule |  | 357
150184 | Manufacturing Engineering | N/A | 2015-07-03 | Genealogy Station | N/A | International | Chia Wee | CLOSED |  | 2016-01-26 | Per the ECO9199200, there is additional requirement for inputting the Genealogy for the detector models. Flextronics only have 1 Genealogy station for the BSM line. They will need another system to support this PKI Detector requirement. Please help to configure and provide one more station for Flextronics. | This OESR is requesting for a Genealogy Station. Per Ken, Dick will take on this request from the original assignee Joe Labenski who already left Waters. WPPL has the 2 Arc Detectors coming up @ Flex with Genealogy requirements added on top of the PKI listed on this OESR. WPPL will need this additional station. Per Dick on Sept 22, 2015: Within Milford operations, engineering lets the production supervisor know that scanning equipment is needed on an assembly line and the production supervisor gets it. The scan equipment plugs into a PC and is ready to use. Is this what you need? Do you already have a PC in the work area that the new hardware can be plugged in to? Do you have a preference on manufacturer brand for the scanner? Sept 29, 2015: Milford to provide a complete set of ready PC (Dick/Mike Cotter) and scanner (Bill Irons) to WPPL. Nov 02, 2015: Dick will call a meeting in Milford to go over details. Jan 14, 2016: TRF TFOS A66QCY. | Fred Sharifi | Dick Bersani (Joe Labenski) | 357
150207 | Test Engineering | - | 2015-07-16 | N/A | - | International | Desmond Ann | CLOSED |  | 2015-08-26 | We are unable to access Sigmasure. Please help to check | CLOSED. | Eric Turner |  | 357
150212 | Test Engineering | - | 2015-07-15 | N/A | - | International | Chia Wee | CLOSED |  | 2015-08-26 | The Power Supply of the Fiber Optics Light Source is down. It is used in the Clean Room optics bench. OEM : Techniquip Corp Model : R150 MV2 PN : 596-20-383 In the interest of the production, we identified a replacement for your review. If it is not acceptable please help to provide the same asap. | CLOSED. | Eric Turner |  | 357
150218 | Test Engineering | N/A | 2015-07-27 | TE82882 | N/A | International | Chia Wee | CLOSED |  | 2016-03-29 | TE82882 Additional Tester Request | Eric will review data already provided from Chia to determine the needs to make the additional tester. Sept 29, 2015: WPPL requests for 2 Testers with CEA approved/on the way. Nov 02, 2015: CEA sent and waiting for updates from Eric. Nov 12, 2015: Have all the parts, beginning construction, software loading soon - hope to ship by EOQ4. Jan 06, 2015: Eric working on Packaging and TRF to ship. Feb 24, 2016: Received at WPPL and will follow up with Eric for setup within this month. March 29, 2016: Closed. | Eric Turner |  | 357
150219 | Test Engineering | - | 2015-07-27 | Acquity Arc TUV Final Test | - | International | Mark Jussaume | CLOSED |  | 2015-08-03 | In order to test 186017002, ACQUITY Arc TUV Detectors, you will need to procure some 2489 analytical flow cells to install into the instruments for final test. Waters PN 205001027 2489 Analytical Cell or equivalent will be fine. They can be the tan or blue doesn't really matter. Get however many you will need up to 20, enough to populate all stations. Here is account info: CEA# USYG15002 Account US02 170650 | Per DO, WPPL have received the 20units of flowcells. | Desmond Ann |  | 357
150220 | Test Engineering | - | 2015-07-27 | WATERPNL1 | - | International | Desmond Ann | CLOSED |  | 2015-08-13 | Model " IBM M50 WATERPNL1 No Video Fan - OK Rotating Power Indicator - ON Suspect to be Motherboard faulty. Temporary replace with a white Lace CPU (Windows 2000) decommission from another tester. Will require Milford to configure the replacement CPU. IP address :- 10.249.26.3 Long term : Request for a replacment CPU with the configuration setup in Milford to send Flex. | The PC is already setup and is working now. | Bill Irons |  | 357
150223 | Test Engineering | N/A | 2015-07-30 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-03-30 | We received ECO9199778 on the update SOP WAT999267 and noticed that there is a requirement to validate the backpressure testing station with PN 8854 daily (Page 4 of 16, Operation 5.1). We don’t have this part. Please help to send this part to WPPL and share with us setup and the maintenance procedure (PM) of this part. | Steve R. on 30 July, 2015: We don't have the test sample flow check master #8854 available yet as Engineering is still in the development stages. Once we have the 1st sample produced and here in Milford and calibrated we will sending along the requested sample and information. Chau will ask Steve Rebello and discuss with ME to have Model Shop make the needed tool. Sept 22, 2015: Steve Rebello has received 2 tools and performance evaluation is on-going. The plan is to ship 1 tool to WPPL after evaluation completed and documented, and additional tool can be made upon request. Jan 06, 2016: Chau placed a call to Steve and waiting to hear back from him. Steve R: Still working on the tool with a redesign iteration and the engineer is still sorting out the details. Mar 29, 2016: WAT999267 and 735003758, Chia and Mike Scalon are working on it to get it setup and close it out when notified. Mar 30, 2016: Closed. | Desmond Ann | Steve Rebello | 357
150224 | Manufacturing Engineering | - | 2015-07-30 | - | - | International | Desmond Ann | CLOSED |  | 2015-08-12 | CEI have finished the stocks of the old barrier board 210000255 based on ECO9193655 to use up 210000255 and replace with 210000450. 1) In Sigmasure, under 279002283 assy, it is still stated as 210000255. We need Milford to be change to the new barrier board part number 210000450 in the Genealogy record. 2) ASP 289004472 Rev T has already changed to the new part number 210000450. We need Milford to update the ASP 279002283 to revise to the new part number 210000450. | Chia tested and it works. Closed. | Randy Chin |  | 357
150231 | Test Engineering | - | 2015-08-11 | N/A | - | International | Chia Wee | CLOSED |  | 2015-08-25 | The Server @ GES has a faulty LCD display. No Display. | The server display has been replaced by Chia. | Michael Cotter |  | 357
150253 | Test Engineering | - | 2015-08-20 | Genealogy | - | International | Chia Wee | CLOSED |  | 2015-09-03 | ECO9199200 added added genealogy to 279003530 " 279003530 (ASSY, DETECTOR, PDA, ALTUS, A-10) ASP: 1. Added genealogy scan steps to the procedure " Please add the following personal into the list of operators. Name Badge 1) Koh Huan Leng 246444 2) Wang Qi Ming 246504 3) Kuo Chia Yang 909883 4) Au Weng Onn 246112 5) Kwek Chock Meng 842185 6) Gary Lai Yam Bee 246076 | Completed and Closed. | Bill Irons | Don Farley | 357
150256 | Test Engineering | - | 2015-08-26 | TE82475 | - | International | Mark Jussaume / Joel Deduyo | CLOSED |  | 2015-08-26 | New flow cell P/N 289010022 for Acquity Arc Detector not on TSP WAT025373. They’re tested using the same tester as the 2489 and other TUV cells. ASP 289010022 page 15, step 13.1 instructs operator to use TSP WAT025373. | Procedure updated and linked to part number 289010022. Reference ECR # 9200630. CLOSED. | Mark Jussaume |  | 357
150257 | Test Engineering | - | 2015-08-26 | N / A | - | International | Chau Nguyen / Chia Wee | CLOSED |  | 2016-01-06 | It was decided from today's meeting to close OESR# 140354, 140394, and 140416 to wrap up open OESRs in 2014. This new OESR in 2015 will be assigned to Steve Wright to provide tool drawing/documentation from Thurston, Waters and other vendors. Steve will review the lists from Chia to ensure a complete list of tool is ready to treat it as a real project to raise a CEA, and provide possible timeline for completion. | Reference for 2014 OESR # 140354, 140394, and 140416. | Stephen Wright |  | 357
150266 | Test Engineering | - | 2015-09-08 | 999028612 | - | International | Desmond Ann | CLOSED |  | 2015-11-02 | Flextronics have informed that the Rough Alignment Station is down due to System Freeze / Hang during alignment of the Optic Bench (PDA and eLambda). The IP is 10.249.26.73. We tried to shutdown, reset the wiring but still seeing the same issue after we run the alignment program. | NI Cable P/N 187805C-01 was shipped and received at WPPL on TRF TFOS-A28FX2. Currently request PCI-1410 and PCI-1409 card for spares. Waiting for Eric's feedback on switching to a digital system for long term solution. Sept 29, 2015: No PCI-1410 and PCI-1409 card available for spares. Eric proposes to raise a CEA to make the same Universal Coars and Fine Alignment System as France got in 2014 as the long term solution. The Universal System could test all 3 benches: PDA, Classic 2998, and eLambda. Nov 02, 2015: WPPL will evaluate production space and capacity for the Universal System, and determine to raise a separate OESR after receive BOM+Cost from Eric. Closed. | Eric Turner |  | 357
150267 | Test Engineering | N/A | 2015-09-08 | Ops | N/A | International | Desmond Ann | CLOSED |  | 2016-03-29 | TSP 735000615 is showing Waters ELGA Pure Water in the Disposables required. Will require to change to Waters MilliQ Water as per production practice. | Sept 29, 2015: Chau to ECR the TSP 735000615 to add "Waters MilliQ Water or Equivalent". Nov 02, 2015: TSP revised and ECR submitted: Change# 9202398. Jan 06, 2015: Chau's ECR was approved on Dec 17, 2016. Tyler Modelsky might work on another ECR for the same TSP and to be confirmed. Jan 08, 2016: ECR 9203820 has been submitted for this TSP. Mar 29, 2016: can be closed because it looks like that change was incorporated. | Chau Nguyen | Tyler Modelski | 357
150268 | Manufacturing Engineering | - | 2015-09-08 | 119745 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | Alliance Assy pump w/seal wash enhanced ( Tooling 119745 ) Tool broken see attached, please help to ship replacement. | Sept 29, 2015: Works are in the model shop and should be completed in October 2015. | Stephen Wright |  | 357
150269 | Manufacturing Engineering | - | 2015-09-08 | 999028048 | - | International | Desmond Ann | CLOSED |  | 2015-11-03 | Alliance Assy pump w/seal wash enhanced ( Tooling 999028048 ) Encounter tooling with hair line crack highlight in red. Please help to ship 2 units of replacement. | Sept 29, 2015: Works are in the model shop and should be completed in October 2015. Nov 03, 2015: TRF, SWRT-A3HP5T. 2pcs 999028084 instead of 999041500 listed in the TRF. 2pcs 999026700 | Stephen Wright |  | 357
150279 | Manufacturing Engineering | - | 2015-09-17 | 999040744 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | As per discussed, we will require 3 units of 999040744 for replacement in the Flextronics production. | Sept 29, 2015: Works are in the model shop and should be completed in October 2015. SWRT-A3WS9A | Stephen Wright |  | 357
150280 | Manufacturing Engineering | - | 2015-09-17 | 999040746, 999040756 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | As per discussed, we will require 3 units of 999040746 and 999040756 for replacement in the Flextronics production. | Sept 29, 2015: Works are in the model shop and should be completed in October 2015. Nov 03, 2015: TRF, SWRT-A3HP5T. 2pcs 999028084 instead of 999041500 listed in the TRF. 2pcs 999026700 | Stephen Wright |  | 357
150288 | Test Engineering | N/A | 2015-10-01 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-01-11 | During the Arc Detector Mftg Readiness Qualification Run, it was brought up that the instrument PN is not stated in operation 2.10 of 735001613 Rev R but it is required to enter the Lamp S/N. Will require Milford to update the TSP to add in the instrument PN 186017003, N2971013 and N2971035 into the TSP 735001613 Rev R. | Jan 5, 2016: ECR 9202728 by Tyler Modelski. Approved and Closed. | Tyler Modelski |  | 357
150295 | Test Engineering | - | 2015-10-07 | TE83013 | - | International | Desmond Ann | CLOSED |  | 2015-10-15 | GES have reported that the Window operating system for station TE83013 is corrupted beyond recover. WPPL have verified that the PC is beyond repair and have replaced with a new Lace CPU. The IP address for the test station is 10.249.40.25 please help to configure the station. | Confirmed fix. Closed. | Matt Brule | Michael Cotter | 357
150304 | Test Engineering | N/A | 2015-10-13 | TUV Shunts | N/A | International | Desmond Ann | CLOSED |  | 2016-01-26 | Flextronics have reported 7 TUV shunts p/n 289002450 have shown degradation and will need replacement. Please help to send us the replacement shunts. | Desmond on Oct 21, 2015: Increase order quantity to 12. In process. Jan 5, 2016: Chau received 12 TUV shunts to ship out, and will close after clearing with Finance. Jan 11, 2016: TRF CNGN A63LRT. $5070.72 was charged to WPPL Account WP21 615150, need WPP to confirm. Jan 26, 2016: Received and Closed. | Eric Turner | Chau Nguyen | 357
150305 | Test Engineering | N/A | 2015-10-13 | SigmaSure | N/A | International | Desmond Ann | CLOSED |  | 2015-10-13 | Currently Flex have 2 Tech who are allowed to access to log into Arc SigmaSure. Flex have another 2 trained personnal who are trained on Assy, Detector,2489 W/O flow cell / Assy , Detector 2998 W/O flow cell and request to have access to SigmaSure. 1) 877288 Sun Na 2) 819665 Li Yang | Complete and Closed. |  | Don Farley | 357
150315 | Test Engineering | - | 2015-10-27 | 405006712 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | Needle bend due to wear and tear. Need replacement. Qty: 4 pieces. | To study how needle got bent much more frequent than Milford. To request account to charge for new needles. Nov 02, 2015: 4 reserved blank needles to be shipped today. Chau to observe production usage at Milford with Rick and at CM with Desmond to see if any possible ways to avoid the bending issue. Jan 06, 2016: Closed, but Chau will continue exploring difffent options to make replacement needles with Model Shop. | Richard Davis | Chau Nguyen | 357
150317 | Manufacturing Engineering | - | 2015-10-28 | 9026543 | - | International | Desmond Ann | CLOSED |  | 2015-11-11 | The modified center punch tip head break off. | Tool received by Chia. Closed. | Stephen Wright |  | 357
150321 | Test Engineering | N/A | 2015-11-03 | Sharp AutoBagger | N/A | International | Desmond Ann | CLOSED |  | 2015-11-04 | Key Point can't print the barcode label for the new 10 tubing assy. Refer to attached list for the PN. We would like to request to update the barcode software database. Station #1: Sharp Autobagger Barcode Program: LabelView V8.50 Station #2: Datamax I-4308 Barcode Program is LabelView v.6.07 XLT. | The database is working. Closed. | Bill Irons |  | 357
150326 | Manufacturing Engineering | - | 2015-11-05 | 410002427 | - | International | Chia Wee | CLOSED |  | 2015-11-19 | Request for a supply of screws (Qty 50) of P/N: 410002427. This screw is used in the I2V check valve of the iBSM used in the FTN Final Test Stack. During PM requires the cartridge replacement. the Screw head wear out. Hence will need replacement. It is not available in the production here. | TRF received. Closed. | Stephen Wright |  | 357
150332 | Test Engineering |  | 2015-11-11 | TE82847 |  | International | Chia Wee | CLOSED |  | 2015-11-12 | During operator testing of the 210000422 vista pump PCBA, after 2 pcs board passed test, subsequent board tested failed at last test step QSPI interface. QSPI test connected to J19 and J20 (Port 1 and Port 2), trace the connection inside of the test fixture, the J19 and J20 go through the card with label “Waters QSPI Degasser” . Suspect that the card of “Waters QSPI Degasser” failed due to the Program chip content of U7 corrupted IC U7 marking : MC9086P32 Label on the U7 IC : 6018-0005 V1.05 request replacement QSPI Degasser board from waters engineering. Please help to review the design/test software if it can that be improved upon to eliminate the failure? | On 11, N0v., 2015 we purchased (5), 725000549 ASSY, PCB Degasser, QSPI Multi-Pers boards to ship, gave to Paul D. On 12, Nov., 2015 Paul D. shipped the boards | Paul Deschenes / Tony Fontes | (Mark Jussaume) | 357
150336 | Test Engineering | - | 2015-11-16 | TE82988 | 3 | International | Jason Agbayani-B | CLOSED |  | 2015-11-23 | Test Down when testing 210000478 (Screenshot) attached. | OVT PS was cleaned. | Mark Jussaume |  | 357
150343 | Test Engineering | N/A | 2015-11-20 | TE82988 | 4 | International | Yulharman | CLOSED |  | 2016-03-30 | The GTS1 Computer Cannot turn ON, after Engineer investigate and found the The Burnt is geotest backplane and PSU. geotest module (GX7310 SN.00167) , the capacitor on the backplane measured is 1.4 ohm (Shorted). | Chia will ensure that the replacement PSU was installed properly and recorded. Rick will release a PM to include all neccesary steps to care for GTS going forward. When Milford receives the burnt unit shipped from CM, Milford will try to determine the possible cause of the failure or send it to Geotest in California for analysis if required. Jan 06, 2016: Rick/Chau will contact Geotest in California to get estimate and then ship unit out for analysis/repair. Jan 22, 2016: P.O from WPPL was sent to Marvin to start analyzing failed Geotest Module. Mar 30, 2016: WPPL will keep the repaired Geotest Module as spare, and notify Milford whenever deploy it for any future usage. Closed. | Richard Davis | Chia Wee / Rick Davis | 357
150344 | Manufacturing Engineering | - | 2015-11-24 | 999027952 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | Flextronics have reported that the tool 999027952-1 is broken and requested 2 pcs for replacement. See attached on the photo of the broken tool. | TRF SWRT-A4RQ3R on December 03, 2015. Received and Closed see ECO 9203393. | Stephen Wright |  | 357
150354 | Manufacturing Engineering | - | 2015-12-03 | 999038530 | - | International | Desmond Ann | CLOSED |  | 2016-01-06 | GES feedback the Gap Tool, 999038530 uses in Acquity HTCH-A, 186015042 assembly station is getting thinner going to out-of-specification. We will like to request 2 pcs of spares. According to the Part Drawing on 999038530. Its thickness is 0.020 +/- 0.001 Our measure found that it is currently at 0.01985. | Dec 05, 2015: TRF:SWRTA4VPA4. Jan 06, 2016: Received and Closed. | Stephen Wright |  | 357
150356 | Test Engineering | N/A | 2015-12-04 | TE83013 | 1 | International | Desmond Ann | CLOSED |  | 2015-12-22 | We found that the data for TE83013-1 seems to be missing from the Sigmasure database. We download those database as part of our monthly first-pass-yield analysis. This month we found that the data is missing in our download for the month of Nov, 2015. Can you help check if there are any issues with the database? | Probe was configured incorrectly when CPU was replaced. | Matt Brule |  | 357
150362 | Test Engineering | - | 2015-12-14 | SQT/ORT Station | - | International | Chia Wee | CLOSED |  | 2015-12-18 | SQT/ORT test station is using Empower 3. We tried to use the tester to test to test Alliance. And found the message: License Expiration Reminder message. IP Address: 10.249.26.90 Please help to activate the tester | This test station was setup for Quality, when they want to use it for the full SQT/ORT, ask them procure the license. If usage is minimal, run the carry-over manually seems to be the way to go. If the need is more frequent, Discuss the issue with Quality to pay for the annual subscription of Empower 3. | Michael Cotter |  | 357
150368 | Test Engineering | N/A | 2015-12-22 | TE82526 | 1&4 | International | Desmond Ann / Joel | CLOSED |  | 2016-03-30 | TE82526-1 & 4 Test screen swos pass data but print-out shows failure. Pls. refer to attachment. | Jan 06, 2016: In Process. Mar 30, 2016: Mark knows it requires a code change. If this issue happens at a higher frequency rate, the request's priority would be changed to get it fixed. It's recommended to close the request at this time, and WPPL will monitor the issue to revisit it when required. No risk of passing bad uut as the test record travels with uut. Test record will not allow "0" value as this usually indicates a non valid return string during test so it is failed at test record. | Mark Jussaume |  | 357
150369 | Test Engineering | N/A | 2015-12-22 | TE82888 | N/A | International | Desmond Ann / Joel | CLOSED |  | 2015-12-23 | Clean Room Rough Fine Alignment is down due to interface PCBA TE-82888. Pls. send as spare part. | Built 2 set of boards, also sent parts to build 2 more sets when needed. CNGN-A5GKV5 received at WPPL on Jan 03, 2016. | Eric Turner | Steve Hunt / Eric Turner | 357
150370 | Manufacturing Engineering | N/A | 2015-12-23 | 999029157 | N/A | International | Chia Wee | CLOSED |  | 2016-02-02 | Tool wear and tear after 4 years of usage. | Jan 06, 2016: In Process. Feb 02, 2016: SWRT-A6QH46 | Stephen Wright |  | 357
150371 | Manufacturing Engineering | - | 2015-12-23 | 999026375 | - | International | Chia Wee | CLOSED |  | 2016-01-11 | Tool wear and tear after 2 years of usage | Torque Wrench W/spare head is being sent to Singapore on TRF SWRT-A5VP64. Jan 06, 2016: Waiting for confirmation from WPPL to close. Jan 11, 2016: Parts received: 1) One complete set tool 999026375 2) One piece, custom wrench head without the torque wrench. | Stephen Wright |  | 357
150372 | Test Engineering | - | 2015-12-15 | 999027952 | 3 | International | Desmond Ann | CLOSED |  | 2016-01-11 | In the beginning, we have 3 pieces of Dog Point Screw Tool, Tool number: 999027952-3. Current 2 pieces cannot be use as the screw track has already vanished due to wear & tear. Please help to replenish 2 pieces. | Sending 6 replacement pieces by TRF SWRT-A5VHH5. Jan 06, 2016: Waiting for confirmation from WPPL to close. Jan 11, 2016: Received only 5pcs of the 999027952-3 although the TRF states 6 pcs. | Stephen Wright |  | 357
160737 | Manufacturing Engineering | NA | 2016-04-08 | 999026017 | 00 | International | Desmond Ann | CLOSED |  | 2016-05-19 | The alignment pin 999026017, that Flex have 8 pins and none is within the dimension mention in the drawing released in Mar 2016. We will require to replace the pins. Please help to send us 10 pins and advise the cost. The account number is WP21 615150. | Apr 27, 20106: Parts are in Model Shop Due the first week of May. Steve will ensure: 1) New pins made to have .086" diameter to fit existing hole. 2) Update and Release 999026017 to show pin diameter of .086". May 24, 2016: Tools will be shipped by Friday. June 29, 2016: Waiting for Steve to release updated drawings to close out. This OESR will remain open as the correct drawing is in Rev B in SAP but it is still in Concept, not RG. Released on ECO 9213763. | Stephen Wright |  | 357
160738 | Manufacturing Engineering | NA | 2016-04-11 | NA | NA | International | Desmond Ann | CLOSED |  | 2016-09-07 | CEI have feedback that they have input wrongly one of the serial no. of PS Interconnect PCB (210000385) for C16UPO079H (186015020) and the data has already imported to Sigmasure. We are unable to make amendment for it in Genealogy. Please help to amend the serial number to the correct serial number.The correct serial number should be:030316031C.CEI has printed out the genealogy report and changed the serial number on hard copy, and signed off beside the change. | Apr 27, 20106: Eric Turner and Roland have made several attempts to make the changes or even did a replacement for this item, but it didn't seem to work. When Eric comes back to the office next week, he could look into it again to see if anything could be done in SigmaSure 6 or 7. (Hopefully, this could be done in Sigma 7 when it got deployed overseas). May 18, 2016: Per Eric, the correct serial number is in the record as well as what was in the unit. To discuss with WPPL and close. June 29, 2016: Request to create Sigma log-in account for Edwin Koh, and then close it. July 27, 2016: Edwin Koh to verify and close it. | Roland Poisson | Matt Franklin | 357
160744 | Manufacturing Engineering | NA | 2016-04-13 | NA | 00 | International | Desmond Ann | CLOSED |  | 2016-06-29 | CEI have feedbacked that they need Milford assistance to change the serial number in two units of Heater/Cooler in Sigmasure. As the two Heater/Cooler were used for drop test, WPPL requested CEI to replace with new Heater/Cooler used in the instrument. The serial numbers of the Heater/ Cooler assembly have been input and imported to Sigmasure. CEI is unable to make software amendment for it and will need Milford's help to change Heater/Cooler(279000926) serial no. from C001GM to C001JW for C16CPO598H and change Heater/Cooler(279000926) serial no. from C001GN to C001HU for C16CPO599H.CEI has printed out the genealogy reports and changed the serial numbers on hard copy, and signed off beside the changes. | Apr 27, 20106: Eric Turner and Roland have made several attempts to make the changes or even did a replacement for this item, but it didn't seem to work. When Eric comes back to the office next week, he could look into it again to see if anything could be done in SigmaSure 6 or 7. (Hopefully, this could be done in Sigma 7 when it got deployed overseas). May 18, 2016: Per Eric, the correct serial number is in the record as well as what was in the unit. To discuss with WPPL and close. Closed. | Roland Poisson |  | 357
160766 | Manufacturing Engineering | WAT271087 | 2016-04-20 | 999027699 | 00 | International | Desmond Ann | CLOSED |  | 2016-06-29 | To request for 2 pcs of 999027699 use in ASP WAT271087. Please advise the cost and charge to account WP21 615150. | Apr 27, 20106: Parts are in Model Shop Due the first week of May. May 24, 2016: Tools will be shipped by Friday. Closed. | Stephen Wright |  | 357
160775 | Test Engineering | 735002984 | 2016-04-25 | TE83064 | NA | International | Hong Cheng/ Chia | CLOSED |  | 2016-09-28 | During chemistry test, a error message appear. (Runtime Error Program: C:\Empower\Bin\PROCSRVR.EXE.) When this message appear press the "OK" button. Another message appear (Processing Server Monitor Unable to locate the processing server, PracMon exiting) When this message appear press the "OK" button again, the Unit Under Test will pass the Final Functional testing. This error happen in all FTN test stack from TE83064-11 to TE83094-18. | April 26, 2016: This is currently being investigated. It has been verified that while inconvenient this does not invalidate test results. June 29, 2016: WPPL to keep monitoring since the error message did not occur in Milford. July 27, 2016: On April 26, 2016, power down and restart on test station 14, 16 & 17. The Error appears on test station 16 only and didn't appear on test station 14 & 17. WPPL will keep monitoring the frequency and inform Milford. Sept 07, 2016: Rick will consult with developmet engineer about Empower's issues. Sept. 28, 2016, Mike Cotter has deleted the Empower project and then restored the project. This seems to have fixed the issue. We also tried to delete results from the project, this also seems to have fixed the error. | Richard Davis |  | 357
160805 | Test Engineering | NA | 2016-05-05 | NA | NA | International | Chia Wee Chun | CLOSED |  | 2016-05-11 | The SHARP Autobagger system (S/N: SX-10I-1127-33) at Richport is down. It cannot print. SHARP Autobagger does not have local support. Please help. Line down. | Qualification is not required because I have taken no remedial action. This problem is outside my scope and responsibilities ( labeling software configuration ). However, I have researched the problem, located and provided troubleshooting advice from the manufacturer (SHARP) documentation. I have also provided via an email, links to the manufacturer technical support hotline and service parts. ( see attached email on pg.02 ) | Bill Irons |  | 357
160810 | Manufacturing Engineering | 279003271 | 2016-05-11 | NA | NA | International | Desmond Ann | CLOSED |  | 2017-06-15 | ECO 9199582 made the following changes: 186248900 2489 UV/Visible Detector (blue) Remove: 1 ea of p/n 279002733 ASSY, DETECTOR, 2489 Add: 1 ea of p/n 279003346 ASSY, DETECTOR, 2489, REFRESH ASP 279003271 Rev J (Child part of 279003346) Operation 19.0 page 45/45 and ASP 279003346 Rev A Operation 3.0 page 5/5 request for a genealogy record. But current Genealogy system does not have this project 186248900 in the system. Please help to add this project into Genealogy. | June 29, 2016: TMD 010153 and ECR 9208222 have been submited. July 27, 2016: TE already approved the TMD and it is out for approval from others. Mark will coordinate with Tanya and Desmond to complete ECR with Q4,2016 Target for completion. June 14, 2017: All test procedures and test programs have been updated as necessary to accommodate test related issues resulting from ECO 9199582. A TEV for TE82903 TUV Test System was written and executed. The protocol and report are in SAP at RG status. Also a TEQ was written and performed on all(4) test systems installed at Flextronics. Those reports are in SAP at RG status. The TSPs have been released in SAP to RG status per ECR #9208222. | Roland Poisson | Mark Jussaume | 357
160811 | Manufacturing Engineering | 289004472 | 2016-05-11 | NA | NA | International | Desmond Ann | CLOSED |  | 2016-06-29 | We are finding difficulties in aligning the optics sensor for final test. As an improvement activity for the assembly alignment, we have fabricated an alignment block to help in the alignment assembly. Please see attached for the alignment block drawings and we will require Milford to assign a tool number for this alignment block. We will update the ASP as attached once the tool number is available. | PLACE HOLDER CAD FILE CREATED IN SAP, NO FORMAL DOCUMENTATION EXIST. TITLE: OPTICAL SENSOR ALIGNMENT FIXTURE P/N: 999042576 FILE NAME: 999042576_TLG~00.PRT and DWG file. Cheated drawing files, released under ECO 9222108. | Stephen Wright |  | 357
160815 | Test Engineering | NA | 2016-05-17 | TE82903 | NA | International | Chia, Wee Chun | CLOSED |  | 2016-05-18 | When we compare the Final Test Record printed at the point of use and compare that pulled from Sigmasure We found that the data are completely different for the same UUT Serial Number. ( attached is the 2 copies). Please help to debug. | No qualification required as there was no real issue and no changes made. Explanation was supplied to the Originator. See attachments | Mark Jussaume |  | 357
160817 | Manufacturing Engineering | NA | 2016-05-18 | 999033547 | 00 | International | Desmond Ann | CLOSED |  | 2016-11-30 | We are being informed that there were two styles of fixture in our possession one was a single mark and the other was the dual mark. Milford have retired the dual mark fixture as it caused them problems. Singapore operation is still having the dual mark fixture and it is unable to perform the operation based on the settings used in Milford. We will require Milford to send us 3 sets of the single make punch which Milford is using in the 999033547 rear staking fixture. | June 29, 2016: CEA# WPPL16008 was created by Edwin. Milford Model Shop received request to work on it. Fixture was shipped to WPPL. | Stephen Wright |  | 357
160830 | Test Engineering | NA | 2016-05-24 | NA | NA | International | Gary Lai (Flex)/Chia | CLOSED |  | 2016-05-25 | The Certificate Printing @ Flex has an issue. The Description Is misaligned and missing for 186299800 . (See attached in the follow up email. | Qualification is not required because no changes were made to the Certificate Printing program. I have used the administrative account to inactivate the certificate record for serial number E16998426A, allowing the operator to create a "NEW" record to replace the Certificate that was printed in error. (see attachment email on Page 02) (see attachment procedure on page 03) | Bill Irons | Bill Bates | 357
160831 | Test Engineering | 735003543 | 2016-05-25 | TE82964 | 1 | International | Joel /Chia | CLOSED |  | 2016-05-25 | 1. TUV DOWNLOAD / PRETEST STATION is loaded with old version of A-10 UV Detector.sft for A-10 UV Detector P/N N2971014. As per ECO 9201010 released 04 Nov 15 : Change the revision on Material/Bom from Rev B to C. N2971014 Altus A-10 UV Detector Remove: 1 ea of p/n 361000575 Altus A-10 UV Firmware v1.0.2493 Add: 1 ea of p/n 361000613 Altus A-10 UV Firmware v1.01.2526 | Qualification not required as there was no change to tester or test process. The latest version of firmware p/n 361000613 found in SAP was loaded onto Download Station TE82964 A-10 UV Detector UUT able to get the checksum v1.01.2526 and able to do Final Test. | Mark Jussaume |  | 357
160843 | Test Engineering | 735003657 | 2016-05-31 | NA | NA | International | Gary Lai (Flex) | CLOSED |  | 2016-05-31 | Refer to OESR 160830: There are 3 more similar instances. The Description is not complete. The printout is attached to the email that will follow. Serial Number 1. E1687E702A 2. E1687E756A 3. E1687E757A | Qualification is not required because no changes were made to the Certificate Printing program. I have used the administrative account to inactivate the three (3) certificate records for serial numbers: E1687E702A, E1687E756A and E1687E757A, allowing the operator to create "NEW" records to replace the Certificates that were printed in error. (see attachment email on Page 02) | Bill Irons | Bill Bates | 357
160847 | Test Engineering | N/A | 2016-06-15 | TE82774 | 3 | International | Chia | CLOSED |  | 2016-06-16 | The Alliance CPU board is damaged due to the PCBA 210000400 inside the UUT has a solder short on Connector J12 at pin 7 & 8. This damaged U32, U33 & U26. Both tester is affected. We are able to recover TE82774-2, using an old "spare" board. But we do not have another U26 to replace of the -3 tester. Please ship us Qty 2 Pcs of the U26 (Programmed) | Venture with the help of WPPL was able to repair the CPU boards for the -2 and -3 stations. I (Ed) have put into shipping two more CPU boards to be used for spare parts in case the existing CPU boards need additional future repairs. Both the -2 and -3 stations are running again as have been reported by Venture. No qualification required due to it's a replacement of an identical component. | Edward Jackson |  | 357
160848 | Test Engineering | N/A | 2016-06-15 | TE82774 | 3 | International | Chia | CLOSED |  | 2016-06-15 | During Cool and Heat Test, UUT immediately failed at Test Step 4_2: Checking thermocouple Probe @ Room Temp Measurement: 999.99 Units: DEGC Limits: 18 High: 30 Done: Swapped the Thermocouple Probe with test station TE82774-2. It still failed Checked System NI Card via device manager indicated Working. (See Attached) | I (Ed) sent a spare Omega D1331 and it fixed the problem. | Edward Jackson |  | 357
160850 | Test Engineering | N/A | 2016-06-16 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-06-16 | Flex feedback that while printing certificate description wording extra text as shown below highlight in red. Please assist to correct. See attached for the photo. S/N F16UPD465A. | Qualification is not required because no changes were made to the Certificate Printing program. I have used the administrative account to inactivate the certificate record for serial number F16UPD465A, allowing the operator to create a "NEW" record to replace the Certificate that was printed in error. (see attachment email on Page 02) (see attachment procedure on page 03) | Bill Irons |  | 357
160851 | Manufacturing Engineering | N/A | 2016-06-17 | 999028024 | N/A | International | Desmond Ann | CLOSED |  | 2016-07-15 | GES feedback that there are missing items which they do not have after reviewing the drawing 999028024. Please help to send us a set of the missing items. Email Request from Desmond to change material to delrin or equivelent, sent to Chau 12/06/2016 | Waters Pacific had the parts made locally, Parts do not need qualifying they have no critical dimensions. See ECO 9212543 for drawing specifics. | Stephen Wright |  | 357
160852 | Manufacturing Engineering | N/A | 2016-06-20 | 999034552 | N/A | International | Desmond Ann | CLOSED |  | 2016-10-06 | Venture feed back Item Number 2: 999034555, fixture, support, front , 30 cm ,CHA at wrong orientation shown in 999034552 drawing. See attached for details. | Drawing corrected and released per ECO9206827Revised per ECO9214044 | Stephen Wright |  | 357
160853 | Manufacturing Engineering | N/A | 2016-06-20 | 999034559 | N/A | International | Desmond Ann | CLOSED |  | 2018-03-20 | Venture feed back that the drawing 999034559 does not have all the parts on the fixture. See attached for details. | Two OESR's for the same fix. Still investigating. Issue corrected see ECO 9206814. Apr 25, 2017: Per WPPL's request, reopen for Steve to double check. Tool was made by Automated Machine. Apr 25, 2018: Steve will get dwg #, and Desmond will update ASP. | Stephen Wright |  | 357
160854 | Test Engineering | N/A | 2016-06-21 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2016-07-15 | Implement password change to all Windows XP PCs of Testers and Genealogy Stations at CMs to enhance computer security. NO Passwords shall be changed on the Labelers, Certificate Systems or Servers !!! | June 29, 2016: Chau to discuss with Mike Cotter about Genealogy Systems' issues per Chia's email to provide instruction to complete password changes at CMs. July 15, 2016: Completed. | Wee Chun Chia | Desmond Ann | 357
160855 | Test Engineering | 279001962 | 2016-06-22 | TE82880 | 2 | International | Chia | CLOSED |  | 2016-06-28 | 1. Rough Alignment Station TE-82880-2 (10.249.26.73) unable to adjust M2 Mirror alignment as per operation 6.7 of ASP #279001962 as the white dot just freeze there no matter how we adjust. 2. Able to move the mouse around as it shows that the computer is working fine. 3. Check all cables and wires , everything ok. 4. Symptom Intermittent. Pleae help to check if there any software, driver etc issue(s). | The target is clean daily and the surface that shown is the material from the opposite site of the target. Rough Alignment is working fine now. | Eric Turner |  | 357
160879 | Test Engineering | 735003657 | 2016-07-11 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-07-13 | 1. Existing printer for WATERSCOFP1 HP Laser Jet 4250 reach wear and tear for roller and no serviceable spare part available. 2. New Printer for WATERSCOFP1 Certificate HP LaserJet Enterprise M604 need set-up configure as Flex does not have administrative privilege to install driver. 3. Location of driver , C:\HP Laser Jet Enterprise M604. 4. WATERSCOFP1 IP address is 10.249.26.4 | No qualification is required because ALL printer output related changes were applied local for this single station and have no broader impact on the Certificate application or data. The printer driver installation application required Microsoft .net. software. Additionally, numerous Windows XP updates were installed requiring multiple restarts. After a final required reboot, the printer driver was installed. The default Windows printer changed, and the Certificate program was print tested. An adjustment to the local Certificate layout was required to align the three (3) printed elements to the proper location on the pre-printed forms. See email confirmation of successful problem resolution on Pg. 02 of this request. | Bill Irons | Bill Bates | 357
160880 | Test Engineering | NA | 2016-07-12 | TE82956 | 2 | International | Chia/Hong Cheng | CLOSED |  | 2016-07-13 | TE82956-2, IP address: 10.249.40.20 is having a problem. Failure symptom is “Thermistor Error”. one UUT is attached to the tester detail as below: Product Part Number: 186015010 UUT Serial Number: G16UPH324G Work Order: 1607090 (Details in email ) | Investigated the reported problem. I found the tester to be fully functional. Chia visited Venture and tested 4 UUTs and found no problems. The test sequence was not current. I updated them to the latest test version as documented in the test procedure. | Richard Davis |  | 357
160897 | Test Engineering | NA | 2016-07-14 | NA | 0 | International | Desmond Ann | CLOSED |  | 2016-07-15 | WPPL is currently retrieving the data from the SigmaSure to the CM for FPY compilation. As per proposal from Edwin and discussion with Eric, moving forward, we are proposing to setup one account for each CM. Below are the details for the account holder to be setup. Flex - Joel (JoelMamalateo.Deduyo@flextronics.com) GES - Kok Leong (kokleong_loo@venture.com.sg) CEI - Summer (zhuxx@cei.com.sg) . | Set up all 3 CM user accounts for SigmaSure 6 with the emails listed below. The password is 'waters' for each user. You should be able to access SigmaServer 6 with an IP address of 10.2.240.189. If someone could pass that along to them, that would be helpful. | Bill Bates | Eric Turner | 357
160919 | Manufacturing Engineering | 289002755 | 2016-07-18 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-10-27 | ASP 289002755 Rev C Assy W/Zipconia Mixer Cartridge, 50 uL. Operation 3.0 Tooling 50 uL Description M & O Perry, Vacuum Wand need to be replaced. Please help to send us a replacement. Please see email for photo of the tool. | PO placed waiting for shipment. Dec 06, 2016: WPPL received the shippment. | Stephen Wright |  | 357
160926 | Manufacturing Engineering | NA | 2016-07-21 | NA | NA | International | Desmond Ann | CLOSED |  | 2016-07-25 | One child unit of Heater/Cooler needs to change the serial no. in Sigmasure. The serial number has been input wrongly and imported to Sigmasure. Please help to change Heater/Cooler(279000765) serial no. from 86904 to 8604 for C001UE (279000926) for instrument S/N G16CPO665H. CEI has printed out the genealogy reports and changed the serial numbers on hard copy, and signed off beside the changes. | We can't change imported data. | Roland Poisson | Eric Turner | 357
160927 | Test Engineering | 735002056 | 2016-07-21 | TE82984 | 1 | International | Desmond Ann | CLOSED |  | 2016-07-21 | Production Operators feed back UUT completed testing the Test Record cannot auto print out and an error message appear on the screen. See screenshot in email. Test Station ID: TE82984-1. IP address: 10.249.40.21 Remark: All the UUT have completed testing and release for tomorrow shipment. Don't have any more UUT. | Directory on server that is required for writing legacy genealogy files was removed. Folder has been replaced and tester is OK. No qualification is required as nothing was changed on tester. | Matt Brule |  | 357
160931 | Manufacturing Engineering | 279002695 | 2016-07-25 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-02-22 | Flex have been using the curing fixture since transfer but was not stated in the ASP. Understand that the fixture was fabricated by Flex previously. Will request Milford to create a fixture number for the curing fixture and see attached on the fixture drawings and redline ASP to include the curing fixture for Michael Dassau to include into the ASP. | New tool # 999042850 Drawing needs to be generated and ASP updated. Requested sketches to produce detailed drawings Aug 2016. Insufficient info on supplied sketches waiting for details from WPPL Nov 2016. SW. Completed and closed . | Stephen Wright |  | 357
160934 | Test Engineering | 735001613 | 2016-07-27 | TE82882 | 13 | International | Desmond Ann | CLOSED |  | 2016-07-26 | Tested 1 golden unit on tester. However, after reaching 88 to 89 percent of running the noise, it will automatically start back to zero percent and status will show loss comm and then reconnecting attempt 1 will appear. This will carry on until 5 cycle and tester will stop the test. We have placed 2 units on TE82882-13 port 1 and port 4 p/n:186299800, 1 unit on TE82882-14 port 1 p/n:186002998 to be verified. IP address for TE82882-14 is 10.249.26.85. IP address for TE82882-13 is 10.249.26.84 | Increased Empower table space size. The table space was full, so the test program could not write the results to Empower, which resulted in a Lost Comm. As this is purely an administrative task, no qualification is necessary. A UUT was subsequently tested, did not lose communication, and passed test. | Eric Turner |  | 357
160943 | Test Engineering | NA | 2016-08-02 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2016-09-07 | GES feedback that the connector on the Racal test fixture for 153 PCBA is broken. See attached on the photo. Will require a replacement for repair. | Sept 07, 2016: Replacement connector shipped to GES on 05 Aug 2016, so this can be considered closed. Closed. | Paul Deschenes | Edward Jackson | 357
160958 | Test Engineering | NA | 2016-08-05 | NA | NA | International | Desmond Ann | CLOSED |  | 2017-01-18 | CEI reported that the fans at the mainframe HP5020A for Racal was not working. The fan is EOL. Does Milford have spare Mainframe to loan to CEI and fan to repair the Racal tester. | Sent our spare to CEI purchased replacement fans on ebay installed replacement fans tested power up sent unit back to CEI Replacement component has been reviewed by the technical expert and is identical to the original. (specification of replacement or PN match, etc). Technical expert observed correct installation. Test system function wasn't compromised by component replacement. | Paul Deschenes | Edward Jackson | 357
160966 | Test Engineering | 735002510 | 2016-08-11 | TE82865 | 5 | International | Desmond Ann | CLOSED |  | 2016-08-17 | One of 279002260 XTS ASSY is down due to cable fault (wire wear and tear). The cable has been used for more than 3 years, Please help to send 3 pieces for replacement and spares. The cable part#: 441001128 REVC | Purchased 3 Cables PN:441001128 from stock, and shipped on TRF- PDES-ACVG6U, on 17, Aug., 2016 No qualification required, this was a cable replacement, specification and part number match | Tony Fontes | Richard Davis | 357
160979 | Test Engineering | WAT046538 | 2016-08-16 | N/A | N/A | International | Matt Brule | CLOSED |  | 2016-09-07 | Please implement ECR 9208662. TMD010198 allows until Week 1 in October to complete equipment setup. Please get back to me with any questions. | Sept 07, 2016: Implementation was completed. Matt will incoporate changes related to errors found by Desmond, and this can be considerred as closed. | Desmond Ann |  | 357
160984 | Manufacturing Engineering | WAT271290 | 2016-08-19 | 999120083 | NA | International | Desmond Ann | CLOSED |  | 2016-10-26 | Flex do not have the tool listed as Detail 2 in tool drawing 999120084. See email highlighted in Red box. As per Steve's advised, we will require Milford to fabricate 2 units of the missing tool for use with tool 999120083 and 999120084. We will also need advise on how is the tool being used in the assembly of the Alliance CCD drive. Please charge to account WP21 615150. | Parts ordered from M-Shop 17-Sep-2016. Parts released per ECO 9212291 and parts were shipped to Singapore see TRF SWRT-AF4JXE. Modification made to part for clearance see ECO 9212291. | Stephen Wright |  | 357
160988 | Test Engineering | NA | 2016-08-22 | TE82743 | 3 | International | Desmond Ann | CLOSED |  | 2017-08-07 | GES have reported that the spare card for PXI-2460 series failed to perform calibration, Failure symptom it’s the card cannot detect. See attached photos for more details. Please advise if there are any solution to resolve or to send back to Milford for repair. | Apr 24, 2017: out for repair. Tra-cal repaired the card and Paul shipped it back to Jason on TRF MPDE-G0-14KN. | Paul Deschenes | Edward Jackson | 357
161002 | Test Engineering | 735001613 | 2016-08-26 | TE82882 | 13 | International | Desmond Ann | CLOSED |  | 2016-08-26 | Flex feedback there is issue on the TE82882-13 tester. Tested 1 golden unit on tester. However, after reaching 88 to 89 percent of running the noise, it will automatically start back to zero percent and status will show loss comm and then reconnecting attempt 1 will appear. This will carry on until 5 cycle and tester will stop the test. The tester have a message "Available tablespace (9.312500) MB) below minimum level to continue processing. Processing aborted". No rogue IP address observed to the station when we go to DHCP Config ID icon. Please help to also advise how can we clear the tablespace from our end. | Increased table space in Empower. Verification consisted of running a unit on the tester. The symptom of this error was the inability of the instrument to communicate with Empower. Therefore, running a unit and verifying it is communicating with Empower is evidence that the action taken was sufficient to resolve the issue. Attached is Test Record and Error Log. Eric to add instruction to PM to clean out Memory. | Eric Turner |  | 357
161005 | Manufacturing Engineering | NA | 2016-08-29 | 999033793 | NA | International | Desmond Ann | CLOSED |  | 2017-04-03 | While comparing the released tool drawings with the physical part, the hole location is found to be different. Please help to correct the drawing per attached. | Updating fixture DWG 20-Sep-2016, see ECO9204438. ECO Stuck 22-Feb-2017 S.W. Released on ECO 9219438 | Stephen Wright |  | 357
161006 | Manufacturing Engineering | NA | 2016-08-29 | 999029571 | NA | International | Desmond Ann | CLOSED |  | 2017-03-03 | While comparing the released tool drawings with the physical part, the hole size is found to be different. Please help to correct the drawing per attached. | Updating drawing to reflect tool. See ECO 9210430 | Stephen Wright |  | 357
161010 | Manufacturing Engineering | 279001852 | 2016-08-30 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-16 | In preparation for the localization of the 279001852 Assy Drive Injector Housing with Venture GES to have them self support the assembly drive injector housing requirements from 2017 onwards. Below is the list of tools we required to help support for the project. Milford to provide the cost and lead-time to duplicate another set. In addition, there are some queries. Is there anyone from your side that could help advise on those queries attached. 1. 999028070 Encoder Bend / Assy Tool 2. 999028183 Holding Fixture 3. 999028250 Encoder Gap Tool 4. 999029204 MOTOR ASSY FIXTURE 6. 999029498 Grease Remover 7. 999029024 Motor Assy Fixture Below is the list of testers we require for the localization: To be taken of by OESR 161048. | Sept 07, 2016: Edwin to raise CEA to build 4 Testers first per OESR 161048. Another CEA to build tools/fixtures will be raised when getting all info from Frank. NOTE: This project is on hold pending further instructions per Ken Voisine. OESR closed because project was cancled. | Stephen Wright |  | 357
161035 | Test Engineering | NA | 2016-09-05 | TE82510 | NA | International | Desmond Ann | CLOSED |  | 2016-09-09 | 1. The PCBA model 210000384 is not able to pass the testing. Error message prompted "DOOR SWITCH FAILED! WITH THE INPUT OPEN". And display on the test fixture did not appear during testing. Note : This tester is cater for model 210000411, 210000444, 210000416 and 210000425 PCBA testing too. All of them may pass, but abnormality observed at the 2690 Display (Refer excel file individual model details). Suspect that CPU (WAT056370) in the 2690 PCB TESTER defective. Will require Milford for advise and also send us the replacement and spare CPU board (WAT056370). | The tester was repaired by Sua Sin Jein at Flextronics. It was found to be bad receiver contacts on Receiver Block 3 where the J15 CPU ribbon cable connects. | Edward Jackson |  | 357
161037 | Manufacturing Engineering | NA | 2016-09-05 | 999032941 | NA | International | Desmond Ann | CLOSED |  | 2016-12-06 | GES have feedback that the physical tool and the drawing is different. The drawings does not indicate the details highlighted in red which is different from the physical tool. Please advise. | Under advisement. Closed per WPPL . | Stephen Wright |  | 357
161042 | Manufacturing Engineering | NA | 2016-09-07 | 999025959 | NA | International | Desmond Ann | CLOSED |  | 2017-01-27 | GES feedback that for the tool 999025959, the tool does not consist of item 3, 4 and 5. See attached for details. | Looking into Reason for tools...3,4 and 5 are not present or required. Requested thru E-Mail to Desmond to provide justification why tools are not required. SW. Created drawings and had model shop make the parts. See ECO 92211942. | Stephen Wright |  | 357
161043 | Manufacturing Engineering | NA | 2016-09-07 | 999026597 | NA | International | Desmond Ann | CLOSED |  | 2017-05-16 | 1) GES feedback that they do not have physical tool item 2: 999042639 and item 7: 410000664 mention on the parts list table in 999026597, Assy Fixture Lamp 2475. Please help to advise how is item 2 and 7 used if required. 2) We notice that the gauge, item 2 : 999041215 of tool assembly 999026585, Fixture Lamp Adjustment Housing is identical as item 3 : 999042640 of tool assy 999026597, Assy Fixture Lamp 2475. GES don’t have this gauge item 2 : 999041215 of tool assembly 999026585, Fixture Lamp Adjustment Housing. ASP 279000320 Assy XE Lamp & Holder at page 8 of 9 it states the gauge as 9026587. The physical part we are using is item 3 : 999042640 of tool assy 999026597 shown in the photo attached with engraved 9026597. Please help to update the ASP. | Under advisement items hold lamp in vertical position for ease of assembly. Drawings is released so WPPL can have parts made locally. See ECO 9208514. | Stephen Wright |  | 357
161049 | Test Engineering | 735003612 | 2016-09-09 | WAT027629 | NA | International | Desmond Ann | CLOSED |  | 2016-09-22 | Venture feedback that the WAT027629 10 ml Syringe crack after drop. Will need help from Milford to send a replacement. | Sept 19, 2016: Desmond received 2 syringes Chau shipped on Sept 13, 2016 through TRF CNGN - ADRRMF. | Chau Nguyen |  | 357
161052 | Manufacturing Engineering | NA | 2016-09-13 | 999120082 | NA | International | Desmond Ann | CLOSED |  | 2016-12-06 | Flex feedback that the tooling 999120082 Linear Motor To Top Plate . 999040708 Base Linear Motor to Plate (Thickness not tally to physical ) Drawing Show 0.375 and Physical measure 0.525. Please help to amend the drawings. | Updating fixture DWG 20-Sep-2016. Closed per WPPL. | Stephen Wright |  | 357
161053 | Manufacturing Engineering | NA | 2016-09-13 | 999119906 | NA | International | Desmond Ann | CLOSED |  | 2016-11-16 | Flex have feedback that the tool drawing 999119906 is missing a 8mm socket. See attached for details. Please help to update the drawing. | Updating fixture DWG 20-Sep-2016 See ECO 9211722 for drawing changes. | Stephen Wright |  | 357
161072 | Manufacturing Engineering | 279001991 | 2016-09-16 | 999038298 | NA | International | Desmond Ann | CLOSED |  | 2016-12-06 | GES feedback that the ring spring on the FTN need replacement. Please help to send us some spares. Is there any special requirement on these rings or we can get normal chain rings locally for replacement. | Need more information. Sent e-mail to Desmond. Ring to be purchased locally at WPPL. | Stephen Wright |  | 357
161074 | Manufacturing Engineering | 279010045 | 2016-09-16 | 999027952 | 01 | International | Desmond Ann | CLOSED |  | 2016-10-28 | GES feedback that they don't have item 1, 999041598, Pin Gauge .084 DIA plus and item 7, 999041617, Wrench, Spanner of the tool 999027952. Please help to send us the tool. | Parts ordered from M-Shop 17-Sep-2016, parts in RG status ECO 9203393. Item received on Nov 25, 2016. | Stephen Wright |  | 357
161095 | Manufacturing Engineering | NA | 2016-09-28 | 999026620 | NA | International | Desmond Ann | CLOSED |  | 2018-02-01 | Flex have reported that the plate 999040910 and 999040911 are different from the drawing. See attached for details. 410005817 which is part of 999026620 is still in concept and not release in SAP. | - Updating drawing to match existing tool. - May 30, 2017: Waiting for verification from Singapore Feb. 2017 SW - August 30: WPPL approved changes and is waiting for the official release into SAP of 3 drawings: 99902660 Rev B, 999040910 Rev B, and 999040911 Rev B. - 01-Feb-2018 - ECR 9220723 submitted to release tool 99902660 Rev. B and associated children. | Stephen Wright | Dan Guertin | 357
161096 | Manufacturing Engineering | NA | 2016-09-28 | 999025528 | NA | International | Desmond Ann | CLOSED |  | 2017-02-22 | Flex feedback that the physical tool 999025528 Pad, Holder, Flat Wire Clip in Flex do not have item 3, (999040766 Pad , Wire Clip Holder). Please help to remove item 3 from the 999025528 drawing. | Under advisement. same fixture used in Milford and clip/spacer is used. prefer to have clip/spacer. SW | Stephen Wright |  | 357
161103 | Manufacturing Engineering | 279000886 | 2016-09-30 | 999038952 | NA | International | Desmond Ann | CLOSED |  | 2016-12-06 | Flex feedback that there is an ECO 9208194 that uprev 279000886 and added a tool 999038952. Flex did not receive this tool. Please help to send us a tool 999038952. | Updated drawing 999039008 opened up interior cavity from 2.25 to 2.270. Part will be made locally by WPPL. | Stephen Wright |  | 357
161112 | Test Engineering | na | 2016-10-03 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-10-04 | Kindly grant access to SIGMASURE application for both rev 6 & 7 to access record in serial tracing & test records for the below WPPL user. wesley_wong@waters.com | Access added for Sigmasure 6/7 for above employee on 4-Oct-2016 | Bill Bates |  | 357
161113 | Test Engineering | 735002132 | 2016-10-03 | NA | NA | International | Desmond Ann | CLOSED |  | 2016-10-05 | Sigmasure is already implemented in Milford. WPPL is creating the access for Sigmasure 7 account. 1) ID 850596, Chen Xue, email xue.xhen@flextronics.com from Flex have resigned. please remove her account and also delete her badge number from the operator list in Sigma 7. 2) To update the document of 735002132 Genealogy that to use the manual document 730002048 as an alternative genealogy instruction in the event that the sigmasure server is down. | Removed Chen Xue's access to Sigmasure 7 and removed her badge number from the operators list. OK to close. Part 2 of the request to be completed on OESR 161127 and managged by Burt Bliss | Bill Bates |  | 357
161117 | Test Engineering | 279002655 | 2016-10-04 | 9029221 | 2 | International | Desmond Ann | CLOSED |  | 2017-05-30 | Flex feedback that the lamp use in the 2998 fine alignment station is not working. I will need Milford to help purchase 2 units of lamp WAT097323 to be sent to us as replacement and spare. The account to charge is WP21 615150. | 2 lamps were shipped. Parts drawn from stock for WPPL to keep as spares, so no qualification required. | Chau Nguyen |  | 357
161121 | Manufacturing Engineering | NA | 2016-10-04 | 999026259 | NA | International | Desmond Ann | CLOSED |  | 2016-11-01 | Flex feedback that the tooling 999026259 Fixture Installation ,O-ring to Seal in Flex only consists of items 10, 5, 7, 6, 9 and 4 of the drawing. 1) 999041065 Plate ,O-Ring to Seal Housing (Item 4) 2) 999041066 Adapter ,O-Ring to Seal Installation (Item 5) 3) 999041067 Locator ,O-Ring to Seal (Item 6) 4) 999041068 Punch ,O-Ring to Seal ( Item 7) 5) 999041071 SHCS , 1/4-20 UNC-2A X 1.75 LG , Reworked (Item 9) 6) 410005888 Knob , Ball , CL-852-PB (Item 10) ( Without Drawing ) Please help to amend the drawing accordingly and also drawing 410005888 is still in concept. | Up dating drawings to match existing fixtures. See ECO 9211576 | Stephen Wright |  | 357
161132 | Manufacturing Engineering | NA | 2016-10-07 | 999025481 | NA | International | Desmond Ann | CLOSED |  | 2016-12-06 | There is a typo error in the drawing 999025481 for part 999032667 in the part list. Need to amend McMaster-Carr from 5147157 to 5147A57. | Correction is being made to the drawing. Tool qualification is not required documentation only (SEE ECR 9212036) | Stephen Wright |  | 357
161134 | Manufacturing Engineering | 279003400 | 2016-10-10 | 999038140 | NA | International | Desmond Ann | CLOSED |  | 2016-11-21 | The head of the torque wrench and wear and will slip during tightening. See photos for more details. | Modified wrench to have a thicker head, doubled griping surface for better engagement. Wrenches were calibrated at Waters Milford prior to shipment. See TRF SWRT-AFYHTC for shipping details. | Stephen Wright |  | 357
161146 | Manufacturing Engineering | NA | 2016-10-12 | 999120114 | NA | International | Desmond Ann | CLOSED |  | 2016-10-21 | Flex feedback that the fixture 999040782 drawing is having a missing hole. Fixture 999040782 is item 1 of 999120114 Fixture drawing. See attached for the dimensions. | Hole in plate has no physical purpose, drawing will remain as is. | Stephen Wright |  | 357
161148 | Manufacturing Engineering | NA | 2016-10-12 | 999120003 | NA | International | Desmond Ann | CLOSED |  | 2017-09-21 | Flex feedback that the fixture 999120003 is having issue. Item 1, 999041035 drawing is different from the drawing. Item 3, 999042737 is different from physical tool. Actual tool dimension is according to drawing 999041039 as per attached. Drawing 999041039 is still in concept. See attached for details | Looks like Fixture was cut in half, changed drawings to match existing fixture. Drawing package is on hold for 3+ weeks. Singapore has had no response from multiple e-mails from Mike. The dimensions in question are related to the base and have no influence to the operation of fixture. Waiting for verification SW 25-Feb-2017. August 30, 2017: WPPL is waiting for the official release into SAP of 3 drawings: 999120003 Rev B, 999041035 Rev B, and 999041039 Rev A. Apr 25, 2018: On list to be ECO . | Stephen Wright |  | 357
161150 | Manufacturing Engineering | NA | 2016-10-12 | 999027250 | NA | International | Desmond Ann | CLOSED |  | 2017-03-07 | Fixture 999027250 drawing is not different from the actual fixture. See attached for dimensions and photo. | Making modifications to sub assembly parts. The provided sketch dimensions did not line up, WPPL to provide correct dimensions. Fixture part number was changed to 999043279 because it was modified in Singapore and we use the old version in Milford Waters. See ECO 9213540 | Stephen Wright | Michael Dassau | 357
161162 | Test Engineering | 735002984 | 2016-10-14 | TE83064 | 15 | International | Wesley Wong | CLOSED |  | 2016-10-17 | During Chemistry test System Hang loss communction with BSM and UUT. | Provided detailed troubleshooting steps to try. Replacing the CPU2000 board restored ethernet communications | Richard Davis |  | 357
161169 | Manufacturing Engineering | NA | 2016-10-19 | 999025235 | NA | International | Desmond Ann | CLOSED |  | 2016-11-16 | The tool drawing 999025235 is different from the physical tool. Drawing 999025235, 999041665 and 999041667 need to be corrected. See attached for details. | Under advisement. See ECR 9212039 Drawing updated per request and released. | Stephen Wright |  | 357
161170 | Manufacturing Engineering | NA | 2016-10-19 | 999027250 | NA | International | Desmond Ann | CLOSED |  | 2017-04-25 | The tool drawing 999027250 is different from the physical tool. Drawing 999027250, 999039688 and 999039689 need to be corrected. See attached for details. | Under advisement. Drawing was released 2 years ago prior to tooling project. WPPL wants to change 999039688, modification could have been made and not documented. Still under advisement. See OESR 161150 for details. | Stephen Wright | Michael Dassau | 357
161182 | Manufacturing Engineering | NA | 2016-10-25 | 999026505 | NA | International | Desmond Ann | CLOSED |  | 2018-03-20 | Tool 999026505 drawing is different from the actual tool and not located in any part of the physical tool. See attach for the photo of the current tool 999026505. | Fixture was made Automated Machine Design and is part of a different drawing package. Provided Quote from manufacturer for drawings, waiting for capital from WPPL Nov 2016 SW. Apr 25, 2017: Per WPPL's request, reopen for Steve to double check. August 30, 2017: Per WPPL, Tool drawings are being worked on by 3rd party Automatic Machine. 3 drawings were provided so far and there is no timeline for completion and require follow-ups from WPPL. | Stephen Wright |  | 357
161183 | Manufacturing Engineering | NA | 2016-10-25 | 999028067 | NA | International | Desmond Ann | CLOSED |  | 2018-04-25 | Tool 999028067 drawing is different from the actual tool. Child part 999042548 dimension is different from the drawing. See attach for details. More details is required for the press used. | There are two versions of this fixture, Need more info to see who made Singapore's version. No Response. June 28, 2017: Steve will investigate. August 30, 2017: On Priority 3 list to be worked after complete Priority 1 and 2 Lists. Tool drawings are provided by third party and are listed on another list.. SW 25-Apr-2018 | Stephen Wright |  | 357
161199 | Test Engineering | 735003678 | 2016-10-28 | TE83013 | 2 | International | Wesley Wong | CLOSED |  | 2016-10-28 | According to PM procedure on High Temp Column heater tester, "Test Step 5. Search for directory SENDBUFFER. Open all SENDBUFFER folders. If there are any files in those directories, notify test engineering for resubmission." Attached is the directory screenshot below. | Old files removed from directory. No Qualification required, as the files in question were invalid data files that needed to be removed, no affect to the test hardware software or process. | Matt Brule |  | 357
161205 | Test Engineering | 735002271 | 2016-11-02 | TE82995 | xx | International | Wesley Wong | CLOSED |  | 2017-01-09 | Problem Statement: "R/Theta Positioner Test >> Z Pierce" Test Failures (Marginal readings on LSL) on RTZ Testers at ~45% failure rate uptick during late July/early August 2016 at Venture Findings/Action Taken: i) Venture reported that once it fails to "R/Theta Positioner Test >> Z pierce", production operator select retest and the UUT (S/N: 1101161100) passed. ii) "R/Theta Positioner Test >> Z Pierce" is failing marginally at -11 UUT (S/N: 1101161100) compared to the LSL: -10. iii) The marginal fail readings occurs on both testers at Venture. iv) iv) The sudden uptrend in these fails were seen from late July/ early August till now. | ECR 9212730 released Test Software 16. This version increased the Lost Steps to +/-15. | Richard Davis |  | 357
161212 | Manufacturing Engineering | NA | 2016-11-03 | 999033528 | NA | International | Desmond Ann | CLOSED |  | 2017-02-10 | Current drawing does not specify the stopper used in the physical tool fixture. Current tool does not have the magnet specified in the drawing. Please correct the drawing. See attached for more details. | WPPL mentioned this part in not used to assemble these parts. Waiting for verification why this tool isn't required. Drawings were updated and release on ECO 9212403 | Stephen Wright |  | 357
161218 | Test Engineering | 735003029 | 2016-11-04 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-11-04 | During RMA | Issue was resolved by WPPL and CM. | Wesley Wong |  | 357
161235 | Test Engineering | N/A | 2016-11-15 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-12-06 | Kindly grant access to view Test Records in Sigmasure7 for the below users: - wesley_wong@waters.com - desmond_ann@waters.com Please map account privileges from reference user: tyler_modelski@waters.com | Tried using Chrome browser, and the tree automatically displays upon navigation to the "Supply Chain" tab. | Eric Turner | Bill Bates | 357
161236 | Test Engineering | N/A | 2016-11-15 | TE82720 | 2 | International | Wesley Wong | CLOSED |  | 2016-11-15 | Encountered 100% test failures on sample/reference noise failures for universal FE test platform for pcba: W1279001098. Resolution: After replacing x2 custom-triax cables for the sample/reference signals, the units are passing. | Resolution: After replacing x2 custom-triax cables for the sample/reference signals, the units are passing. Replacement component has been reviewed by the technical expert on Nov 15, 2016 and is identical to the original. Technical expert observed correct installation. Test system function wasn't compromised by component replacement. Verified on a UUT after the solution implemented. | Richard Davis |  | 357
161247 | Test Engineering | N/A | 2016-11-25 | 2286I-9817 | 00 | International | Wesley Wong | CLOSED |  | 2016-11-29 | During ICT open/short test, a PCB with an open trace between J9_PIN43 and U49_PIN29 prompts user with "Continue testing (Y/N) ?". If the operator presses "Y", the test will continue and bin as a PASS, instead of a FAIL. Venture team evaluated the *.tpg script, suggest to add line "TEST PINOPENS FAIL(6); BRANCH FDONE;". By adding this line, the PCB with the open trace will FAIL test indicated on the GUI. | program was corrected and set to fail Solution satisfying Test system service request has been assessed by OERS Coordinator. Considerations has been given to the potential impact on the health and safety of the operator as well as performance of the system . | Paul Deschenes |  | 357
161248 | Test Engineering | 735004006 | 2016-11-29 | TE82988 | 03 | International | Wesley Wong | CLOSED |  | 2017-02-07 | Problem Statement: Unable to locate “TE82988 Selftest.seq” from C:\GTS Selftest\ on TE82988-03 at Flex-M'sia Findings: i) According to procedure in CAL docu 735004006, section 2.16, unable to locate “TE82988 Selftest.seq” from C:\GTS Selftest\ ii) In the sub folder C:\GTS Selftest\Rev J\, the sequence file "TE82988 Selftest rev J.seq" was seen. iii) Ran this sequence file and it failed to "Test Pickering Relay MUX", typical failed measurement ~1.7ohms (LSL: 1.5ohms) | Installed the latest version Rev K. on TE82988-3, Flex Maylasia. Please close this OESR and submit 2 new OESRs. One for the file name descrepancy and another for the test failure. | Richard Davis |  | 357
161249 | Test Engineering | 735001356 | 2016-11-29 | 186015022 | N/A | International | Wesley Wong | CLOSED |  | 2017-01-25 | Problem Statement: TUV Golden Unit Box build (S/N:K05UPT479M) is experiencing intermittent connection at Final Test for testing on Optics Bench spares Observations: i) LED on TUV golden unit is not stable prior to testing - Blinking green ii) LED on optic bench UUT is unable to calibrate - Blinking red Actions Taken : Swapped the following, i) Ethernet Cable, ii) RS 232 Cable & iii) Power supply but problem still persists. Suggestion: May need replacement for the following, i) personality board (p/n: 210000258), ii) CPU (p/n:210000282) & iii) lamp power supply (p/n:081115), for further troubleshooting. Note : Attached last calibration report for your reference. | Dec 06, 2016: Mark recommends to try new board first, lamp power supply next to troubleshoot. CPU replacement should be the last resort. Jan 24, 2017: Feedback from WPPL based on recommendation from Mark that it was a Setup error and has been resolved. Closed. | Mark Jussaume |  | 357
161250 | Test Engineering | WAT271129 | 2016-11-29 | N/A | N/A | International | Mike Cotter | CLOSED |  | 2016-12-07 | Please provide the degasser pressure test (PSIA) results for the 50 most recent Alliance instruments. | Data provided by Joel. | Desmond Ann |  | 357
161255 | Test Engineering | WAT271129 | 2016-11-30 | N/A | N/A | International | Mike Cotter | CLOSED |  | 2017-04-05 | A new version of Alliance firmware is nearly ready for release. It hasn't changed since early July, and we've verified it since then. We should check it again, just to be safe. I copied the firmware to the firmware folder on the Autoloader-2 station. To verify, run a known-good unit through test (but with the new firmware) confirming that it passes test and that the new firmware is build # 3.03.739. After the test, install the released firmware before shipment. I don't anticipate any issues, but it's best to be safe. Let me know if you have any trouble. | Joel provided evidence that the Alliance firmware v3.03 still functions fine under final instrument test conditions. No Changes made; no change to qualify. | Desmond Ann |  | 357
161264 | Manufacturing Engineering | NA | 2016-12-01 | 999043035 | NA | International | Desmond Ann | CLOSED |  | 2016-12-16 | In OESR 161053, ECO 9211722 was raised to create 999043035 to replace 999119906. After reviewing the tooling 999043035, we realized that dimension highlighted in red attached in the email should be 1.15" instead of 1.50". Please help to amend. Thank you. | Dimensions were supplied by WPPL when part was originally released, this OESR request is to changed dimension from 1.5 to 1.15 see ECO 9212742 | Stephen Wright | Michael Dassau | 357
161267 | Test Engineering | N/A | 2016-12-02 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-12-13 | Venture is currently unable to access sigmasure weblink -> http://sigmaserver2:8888/sq/ at their facility. I tried to access within WPPL location, and is able to do so. Kindly help check the DNS linkage from Venture server to the main server at Milford please. | The IP address of sigmaserver2 has changed to 10.249.190.189 so they have to edit the C:\WINDOWS\system32\drivers\etc\hosts file to point to the new IP. A ticket into our IT department in Milford to try and resolve. Updating the host file on each PC. | Bill Bates | Eric Turner | 357
161268 | Manufacturing Engineering | NA | 2016-12-02 | 999025895 | NA | International | Desmond Ann | CLOSED |  | 2017-02-22 | The current 999025895 fixture in Flex is without item 4 Pad Skid. The 999025895 fixture without Pad Skid, when tried to use on the assembly, it will not be able to slot the Heater Cooler into the Alliance as the height of the heater cooler to the chassis is too low, there is a misalignment in height of about 0.06in. If we are to add the pad skid 0.24in height, it will cause the fixture to be too high and unable to slot the heater. Our proposal is to add a 0.06" height pad at the base of the fixture and not using the 405000371 Pad Skid. See photo in email for more details. | Discussed option with Flex and waiting for response 12Dec2016. Created mark-up and sent to WPPL to have parts fabbed. See ECO 9213573. | Stephen Wright |  | 357
161269 | Test Engineering | WAT271129 | 2016-12-05 | N/A | 01 | International | Wesley Wong | CLOSED |  | 2017-01-05 | Problem Statement: Work station CPU (IBM M50) of Alliance Final Tester 1 (s/n: KKPAAOH) unable to boot up. History: 16 Nov 2016 – Alliance IBM M50 Workstation FT1 able to Power On but not able to boot-up to Windows. - Replaced Alliance IBM M50 Workstation Main Power Supply - Able to Power ON and boot-up to Windows and able to use until 28 Nov 2016 28 Nov 2016 – Alliance IBM M50 Workstation FT1 able to Power On but not able to boot-up to Windows. - Inspect the whole CPU and found on mother board having bloated capacitor - Not able to boot-up to Windows suspected to be due to leaky capacitor in Motherboard. - Flex does not have M50 Motherboard spare - We used image of Alliance FT 5 (Lace 32) and used LACE32 s/n:C12L32255W as Alliance FT 1 Alliance FT 1 using Lace 32 Next Steps: - Request for Michael to check on the tester PC setup (from M50 to Lace32) environment after re-imaging. | System is up and running well. Closed | Michael Cotter |  | 357
161275 | Test Engineering | N/A | 2016-12-06 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-12-13 | Problem Statement: Test records from Venture, CEI & Flex are not uploading to Sigmasure6,7 since ~01.Dec.2016 Findings: i) Checked on sigmaqueue folder and found test records *.xml files are still residing in this subfolder on FTN stations (Venture) ii) Checked sigmasure web on CPSO (CEI) and Acquity BSM (Flex) have recent records reflecting up till 29.Nov & 30.Nov respectively - There were productions units being ran from December onwards, but they are not being uploaded. Attached pictures on findings. | Current works and pending solution in OESR# 161267 might help resolving this rerequest. Updating the host file on each PC. | Eric Turner | Bill Bates | 357
161281 | Test Engineering | N/A | 2016-12-07 | OESR | N/A | International | Wesley Wong | CLOSED |  | 2016-12-07 | Please help to create an OESR account for the below user. i) Siew Yee Yeap E-mail: Siew_Yee_Yeap@waters.com | Added an ASSIGNEE account for sgsyyea Edited the OESR value list = ASSIGNEE to include "Siew Yee Yeap" | Bill Irons |  | 357
161295 | Test Engineering | WAT271129 | 2016-12-12 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-01-05 | A new version of Alliance firmware is nearly ready for release. During verification, Flex is encountering issues and require Milford advise. See attached email for details. | Provided .txt files for use with the new firmware. Flextronics verified/qualified that these files work successfully. Files have been updated on all test systems. | Michael Cotter |  | 357
161298 | Test Engineering | N/A | 2016-12-12 | N/A | N/A | International | Chau Nguyen | CLOSED |  | 2017-01-05 | PC host file update for all PCs at all CMs : A TE Computer List will be sent along with using the instruction below to update all PCs. Here are the instructions: Login as an administrator Locate the hosts file C:\windows\system32\drivers\etc\hosts Edit existing or add to the host file the following lines: 10.249.190.189 sigmaserver2 10.249.190.127 sigmaprod Save hosts file Reboot PC Login as a user If PC is a test system using SigmaProbe 6, navigate to the sigmaqueue directory C:\Program Files\SigmaQuest\SigmaProbe\SigmaQueue Verify there are no xml files present. If there are notify the systems test engineer. If PC is a test system using CamstarProbe 7, navigate to the camstarqueue directory C:\Camstar\Probe\CamstarQueue Verify there are no xml files present. If there are notify the systems test engineer. No special instructions for genealogy | PC host file update for all PCs at all CMs was complete. | Desmond Ann | Wesley Wong | 357
161308 | Test Engineering | N/A | 2016-12-13 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-12-13 | Please reset below user's Sigmasure7 account and enable tree view on product/station view, applicable for Venture only. - kokleong_loo@venture.com.sg | Account reset - no qualification required to reset an account | Bill Bates | Eric Turner | 357
161322 | Test Engineering | 735003612 | 2016-12-16 | TE82943 | 1 | International | Wesley Wong | CLOSED |  | 2016-12-16 | The interface tubing connecting to the Input Filter port on the TE82943-1 test box has break away. Please help to send in a replacement of the Input Filter port. Attach photo for reference. | Sent Wesley the part number for the in-line filter that needs to be replaced WAT035190, so the part can be ordered locally. This will be a like for like replacement and will not require qualification because the tester will not be compromised by the filter replacement | Paul Deschenes | Edward Jackson | 357
161326 | Test Engineering | N/A | 2016-12-19 | TE82882 | 11 | International | Wesley Wong | CLOSED |  | 2016-12-22 | On TE82882-11 (ip_addr:10.249.26.82), it prompts "Database Connection Failed". See picture Kindly check if there is program dependencies to the database that will affect testing. | Added Computer's IP to hosts file - fixed issue per Yelena Yashtar. No qualification required as this problem was preventing Empower access, now Empower can be accessed. | Eric Turner |  | 357
161327 | Test Engineering | N/A | 2016-12-19 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2016-12-19 | Sigmasure7 web access is down as reported by CMs (Flex & Venture). See snapshot. Please assist as we require to rollout the procedure for Sigmasure7's genealogy. | Reset Tomcat service. Sigmasure7 came up immediately after. | Eric Turner | Bill Bates | 357
161333 | Test Engineering | WAT271129 | 2016-12-21 | N/A | N/A | International | Mike Cotter | CLOSED |  | 2017-05-30 | There is a request to study the carryover performance of current Alliance instruments, as shipped, beginning in January, 2017. As we don't routinely perform carryover testing in production, we’ll need to audit a number of instruments after the completion of final test. It may be simplest to perform the testing at Flextronics, post-test and pre-shipment. At Flextronics there is a 'QA Workstation' (basically an Empower Workstation and 2489 detector). This would be a good station for this testing. To collect accurate data, it would be best to use the 2695/2489 SQT project. There is a copy on the QA Workstation, and on the Test Engineering\Alliance SQT_Rev G\ folder on the Flextronics Manufacturing server. The SQT documentation can be found in the same folder, but we'll basically need (in addition to the 2489 detector-equipped test station): · HPLC Grade Water · HPLC Grade Methanol · 70:30 HPLC Grade Water: HPLC Grade Methanol, pre-mixed for needle wash · Sample solutions kit (Part Number 700003741) · Column (Part Number 186003031) To collect accurate data, we'll need to ensure that the system is in good working order (recent detector lamp with < 1,000 hours), fresh, clean solvents (new water daily) and fresh, clean samples and clean vials (don't keep any pierced vial for more than one week). When creating individual projects for the testing, use the naming convention SQTSerialNumber (the letters SQT followed by the UUT's serial number). Once the project is created and the system is set up/primed, it would be best to perform the System Precision test (020 System Precision SSM) first. This small amount of time would be well spent to ensure that the system is capable of producing good, valid results. After the system precision test runs successfully, run the Carryover test (040 Carryover SSM). Keep a record of the carryover test results on the Alliance Carryover Data 2017 spreadsheet. Keep the system precision and carryover test reports for each instrument (we may want them for post-test analysis). Once the testing is complete, the instrument can be prepped and shipped. We won't need to route all of production through the carryover audit, but it would be best to keep the station busy and gather as much data as possible. Once the testing begins, we’d like to review the data weekly, so either send the spreadsheet to me or save it in an accessible location (e.g., on the QA workstation or on the server). Flextronics has performed this testing before (Joel and/or David would be good resources), and WPPL likely has some knowledge in-house, but let me know if you have any questions. | May 30, 2017: Data collection complete and closed. No qualification is required on collecting data for the study of carryover performance on current Alliance Instruments. | Desmond Ann | Wesley Wong | 357
161334 | Manufacturing Engineering | 289001994 | 2016-12-22 | 999120015 | NA | International | Desmond Ann | CLOSED |  | 2017-03-07 | There is feedback that the WAT031956 seal wash tube was too short. After investigation, it was observed that there is a chance of over travel which may compress the WAT031956 seal wash tubing causing it to be shorter. Will require Milford to release the drawing of the seal tube press 999120015 and add a limiter into the SAP. See email for more details. | Sent information to purchase stroke limiter, drawing requires update followed by an ECO. 3-Jan-2017. Jan 24, waiting for Desmands feedback on PDF's. S.W. Ordered parts because Desmond said the can not order them sent out on TRF SWRT-AHWJYY. Changes released on ECO 9214143 | Stephen Wright |  | 357
161340 | Test Engineering | NA | 2016-12-27 | NA | NA | International | Yeap Siew Yee | CLOSED |  | 2016-12-27 | Please help to create SigmaSure 6 & SigmaSure 7 account for the user below: 1) Yong Fook Cheong Email: YongFC@cei.com.sg | Access added for both SigmaSure 6 & 7 for Yong Fook Cheong. Username and password for SS6 is email + 'waters', Username for SS7 is email + whatever you set it as (you should receive an email detailing setup). Recommended password is 'waters'. No qualification needed for adding access to SigmaSure. | Bill Bates |  | 357
171352 | Test Engineering | WAT271129 | 2017-01-03 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-01-06 | In ECR9209687, it is only applicable for new FW 3.03 for two products: 186269505 & 186269506. The corresponding setup.txt file has changed from the last OESR#161255 raised. However. the PKI model (N2971010, FW 1.xx) is not compatible on same setup file. Could we suggest to have the TSP procedure to direct operator on the setup file usage between the PKI & xxx505/506 models? This is for clarity in operations when performing testing. | TMD010268v00 issued and approved Jan 05, 2017. (ECR 9213029 effective Apr 05, 2017) | Michael Cotter |  | 357
171359 | Test Engineering | N/A | 2017-01-04 | TE83064 | N/A | International | Wesley Wong | CLOSED |  | 2017-01-09 | Problem Statement - During CPSM-FTN testing, the timestamp of a failure is recorded before the actual test start time. Findings i) Simulate fail on CPSM (S/N: xxxx270G) and checked the current timestamp from tester GUI -> 2:27pm, 12/29/2016 ii) Check on SigmaSure7 web, and noticed the START timestamp to be -> 2.06pm, 12/29/2016 iii) The fail record for START timestamp in SigmaSure7 is before the actual start of test -> 2.16pm, 12/29/2016, which is not possible. | This is intentional. It was a fix for the way Sigmasure computes FPY. It only occurs on failing test records and has no impact on statistical parametric data. | Richard Davis |  | 357
171367 | Manufacturing Engineering | NA | 2017-01-09 | NA | NA | International | Yeap Siew Yee | CLOSED |  | 2017-01-19 | Problem Statement: Unable to print Geneology Report in SigmaSure7. Observation: -Login SigmaSure7 with Firefox (Version 30). -Select 'File->'Print Preview'. Report content showed blank page. -Select 'File->'Print'. Report printout showed blank page too. | Actions taken at CEI, Singapore. i) Install IE8 on the genealogy stations at CEI (previously from IE6) ii) Checked out the printing -> okay iii) Walked through with CEI the procedure in usage of SigmaSure7 genealogy -> no major concerns | Roland Poisson | Eric Turner | 357
171368 | Test Engineering | N/A | 2017-01-09 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-01-24 | Problem Statement: Serial Number printing out of sequence for iBSM & arc-QSM at Venture Findings: i) Checked on manual records during printing of serial numbers on iBSM -> Serial number generation was observed to be: a) "xxx202G" generated on 27.Dec.2017 b) "xxx178G" to "xxx184G" generated on 06.Jan.2017 c) "xxx203G" generated on 08.Jan.2017 ii) Checked on manual records during printing of serial numbers on arc-QSM -> Serial number generation was observed to be: a) "xxx416G" generated on 24.Dec.2016 b) "xxx001G" to "xxx013G" generated on 06.Jan.2017 | Waters Instrument Serial Number station ( VENTUREISNL1 ) was prematurely mapped to the "NEW" VENTUREMFG server ( approx. 09-JAN-2017 ). The "NEW" server Data Base(s) had NOT yet been fully synchronized with the CURRENT manufacturing data on the "OLD" server. The "NEW" server has now been brought on-line and the labeling tables have been updated. I have re-mapped the VENTUREISNL1 station to the "NEW" Server and successfully tested same. There will be NO ACTION taken to address the gap in serial numbers since unique Waters Instrument Serial Numbers is mandatory, and changing the Model Code counters might risk a serial number duplication. | Bill Irons |  | 357
171370 | Test Engineering | 735002984 | 2017-01-10 | TE83064 | N/A | International | Wesley Wong | CLOSED |  | 2017-09-25 | In section 2 of test procedure (TSP735002984, rev K), it does not list down the PKI model (N2971031). See snapshot attached. Is PKI model (N2971031) supposed to be listed in this section? | There is no Functional Specification Document (FSD) for the PKI model so It cannot be referenced for specifications. When an FSD becomes available the test procedure can change. RDE has produced an FSD however the manufacturing test specifications are incomplete and incorrect. The documents will require change. 9/11/17 - PKI part number has been added to specifications section with H-Class part numbers in tsp 735002984 Rev. L. Which will be released after IQ activities are complete ECR 9218193 | Richard Davis |  | 357
171385 | Test Engineering | N/A | 2017-01-16 | TE82536 | 2 | International | Wesley Wong | CLOSED |  | 2017-01-17 | The ICT Test Fixture TE#82536-2 used to test model 210000425 needs to be repaired due to air leakage during suction in testing. It was found out that the PCBA rubber guide was hardened and its physical condition has deteriorated possibly from wear and tear. Attached is the quotation on its job activities and pictures of the current ICT test fixture condition. We would like to run by Milford TE owner on its proposal to perform this repair job. Can you let us know if its good to go ahead? | Email from Wesley, Dated 01/17/2017. Hi Paul, It would be using the molded seal as the original, as it is a replacement activity. Yes, WPPL will be paying for this as refurbishment. I sounded off to Flex on a possible vacuum box, to see if they can explore with their local vendor on its availability. However, seeing that the fixture is residing in Malaysia and the PCB testing volume is relatively low, we would pursue the current reburbishing proposal as quoted. Thanks for the response. Hi Mike/Chau, please close this OESR. Thanks. | Paul Deschenes |  | 357
171389 | Test Engineering | N/A | 2017-01-16 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-01-24 | Please reset below user's SigmaSure7 account. - YamBeeGary.Lai@flextronics.com | Password reset to 'waters' for email address YamBeeGary.lai@flextronics.com No qualification needed for account updates. | Bill Bates |  | 357
171392 | Manufacturing Engineering | N/A | 2017-01-17 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-02-22 | Summary of discrepancies seen at Venture on SigmaSure7 Genealogy:: I) Product 18601500, there are no fields to input "Side A/B" and "Primary/Accumulator" in SigmaSure7 for its subassemblies. ii) Product 186015043, there is a missing sub-assy -> 210000520 in SigmaSure7. iii)Product 186247500, there is a missing "Build" selection on "186247500 2475 Detector" iv) Product 186015046(FTN-I-Class), the sub-assemblies are incomplete after comparing to 18605017(FTN-H-Class), knowing they are identical in its sub-assemblies. | Completed and closed . | Roland Poisson |  | 357
171403 | Manufacturing Engineering | N/A | 2017-01-19 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-02-21 | Summary of discrepancies seen at Flextronics on SigmaSure7 Genealogy: i) For products with more than 1 sub-assy of the same part, the index number is running for SigmaSure7. Compared to SigmaSure6, the index number it is grouped for products with more than 1-subassy of the same part. ii) Product 186002487 & 186002998 are missing as selection for SigmaSure7 genealogy iii) For product 186015034, it is missing 2 sub-assemblies: - PCB, CPU 210000520 - Lamp Power Supply 279001105 | Per Wesley: "I've gotten Flex to check on SigmaSure7. They feedbacked and its okay, please close out this OESR." | Roland Poisson | Eric Turner | 357
171420 | Manufacturing Engineering | NA | 2017-01-25 | 999120062 | NA | International | Desmond Ann | CLOSED |  | 2017-03-10 | Flex have feedback for tool 999120062 under priority 1 toolings: Wear n Tear ( Request for replacement from Milford) 1) 999040999 Gear, Modified Syringe Drive Assembly Found Discrepancy on below drawing 999040999, please amend the drawing. 1) 999040999 Gear, Modified Syringe Drive Assembly a) Physical part do not have .075 opening ( Highlight in red ) b) Physical part with Gear and drawing never indicate ( Highlight in red ) Note :- Flex don’t have below parts as shown in drawing 999120062, please amend the drawing to remove the following items. 1) item 5 - 999040975 Screw , Shoulder , Modified , .250 OD X 1.75 L 2) item 6 - 999040978 Block , Support , Syringe Drive Assembly . 3) item 9 - 999040995 Shaft , tool , Syringe Drive Assembly 4) item 14 - Knob, Oval Tapered, Carr Lane P/N CL-66-PTK 5) item 11 - 999041000 Plier, Round Nose. 7’’ Long , Modified | Discussed option with Flex and created mark-up and sent to WPPL. Waiting for response 18-Jan-2017 Changes were released on ECO 9214322 | Stephen Wright |  | 357
171421 | Manufacturing Engineering | NA | 2017-01-25 | 999025231 | NA | International | Desmond Ann | CLOSED |  | 2017-03-08 | Flex have feedback the below for tool 999025231, please help to amend the drawings. Drawing 999025231: 1) 999025231 Drawing never indicate Toggle Clamp 2) 999042271 Dowel Pin, 0.1875 Dia X 1.00 LG height protrusion is 0.49 instead of 0.45 3) 999042271 Dowel Pin, 0.1875 Dia X 1.00 LG is a tapper head and not round head. Drawing 999042271: 1) Dowel pin has a cut out and it is not shown in the drawing as highlighted in blue. 2) Tapper face up in the physical tool as shown in the photo. | Discussed option with Flex and waiting for response 12Dec2016. Created mark-up and sent to WPPL. Waiting for response 8Jan2017 S.W. Changes released on ECO 9214374 | Stephen Wright |  | 357
171422 | Test Engineering | N/A | 2017-01-25 | TE82903 | 9 | International | Wesley Wong | CLOSED |  | 2017-02-02 | Problem Statement: No boot, suspect PC crashed on TE82903-09 Actions Taken: i) Replaced to a LACE 11 PC and clone HDD from reference Lace11 PC ii) Re-assign IP to 10.249.26.46 Request below to Milford to perform recovery of the O/S and test program in TE82903-09 i) Config/activate Windows XP ii) install drivers for PCI Device iii) Unknown device iv) Install VGA controller v) Check if program is the latest revision | Updated drivers on the missing items in Device Manager, only 1 remains (unknown device). Wesley had Flex test several UUT on this test station and all appears to be functional now. OK to close this OESR. No qualification needed, only the PC has changed. Updated TE PC list. | Michael Cotter | Bill Bates | 357
171431 | Test Engineering | WAT046538 | 2017-01-25 | N/A | N/A | International | Bill Bates | CLOSED |  | 2017-04-24 | New instructions created in WAT046538 Rev 12 for validating the Hi-Pot testers with a 'test box'. Singapore CMs should have a number of these, that will need to be reworked per the attached instructions. ECR is currently being written to make this rev 12 official. Resistors for this project are to be shipped over today from the US. | Boxes have all been updated. Verification will be performed via calibration of boxes. | Desmond Ann | Wesley Wong | 357
171435 | Test Engineering | WAT271129 | 2017-01-26 | TE82736 | NA | International | Desmond Ann | CLOSED |  | 2017-05-30 | WPPL received feedback from the field that there were 16 Alliances report 'Missing vial" most of them report at vial# 60, one or two days later 5 of other Alliances(not first 16 fault Alliance) report "Missing Vial". All serial numbers reported were build in Nov 2016. Action taken by the service engineer is to adjust the lower vial sensor alignment. There is 1 Alliance which encountered intermittent with 2 times missing vial error at location 37 out of 14 times RV test observed in the factory. Action performed so far, i) Material: Inspected parts that affect both the bottom and top sensors, i.e the holders x 2, bracket mounting ii) Method: Look at the assembly process, alignment of the carousel and perform puncture test measurement analysis. Understand that RDE engineers in Milford are helping to investigate the issue, and hope to find the root cause to provide a resolution. This OESR is raised for this "Missing Vial", so that we can use this OESR to follow up on the actions performed. | Feb 22, 2017: WPPL to work with CM to investigate the findings related to "Missing Vial" Issue per Alex from RDE in Milford. March 23, 2017: Carousel Adjusment Instruction was provided to FSE. WPPL/CM is performing PM to bring both CCD Testers back to normal production. Apr 24, 2017: Both CCD Testers are back to normal production. Bracket issue will be followed up and completed by WPPL. Apr 26, 2017: WPPL will ship one Go/No-Go Gauge for the bracket to Milford for evaluation, and OESR will be closed when Bracket Issue is resolved. May 30, 2017: Waiting for updates from WPPL to close. WPPL proposed to close this request, and no qualification is required. | Chau Nguyen | Michael Cotter | 357
171453 | Test Engineering | 735004006 | 2017-02-01 | TE82988 | 3 | International | Wesley Wong | CLOSED |  | 2017-04-20 | Problem Statement: Performed self-test with TE83004 fixture on GTS (TE82988-3), it failed to NI-PXI-2588 test sequence. Actions Taken: i) Mount TE83004 self-test fixture on GTS ii) Run "C:\GTS Selftest\Rev K\TE82988 Selftest rev K.seq", still fail iii) Remove and re-seat TE83004, still failed iv) Attached failed log | There were numerous problems causing selftest to fail. 1. multiple relay failures were resolved by cleaning the connectors to the relay cards. 2. I found Instrument LAN configuration was incorrect (changed). fixed. 3. The printer port cable to the PC's LPT1 port was disconnected. As reported by Venture selftest now passes | Richard Davis | Paul Deschenes | 357
171461 | Test Engineering | N/A | 2017-02-06 | N/A | N/A | International | Yeap Siew Yee | CLOSED |  | 2017-02-06 | Problem Statement: Database Connection Failure Description: Flex is experiencing Database Connection Failure which might due to license renewal or authorization access issue. IP address of the affection SQT station: 10.249.26.90 | Per Desmond Ann, the database connection error has been resolved. No changes were made, as the error resolved itself. Perhaps restarting the computer resolved the issue. If the error recurs, we'll dig into the root cause, but for now, all seems well, per Desmond. | Desmond Ann |  | 357
171465 | Manufacturing Engineering | N/A | 2017-02-09 | N/A | NA | International | Desmond Ann | CLOSED |  | 2017-05-12 | Milford to create a drawing for the modified 13mm socket. Sketch is shown in the email. This tool is used in GES Singapore for Column Manager assembly CM002. | Created drawing and is released on ECO 9214029. | Stephen Wright |  | 357
171472 | Test Engineering | N/A | 2017-02-13 | N/A | N/A | International | Bill Irons | CLOSED |  | 2017-02-27 | The VENTURE - Certificate Of Functional Performance computer drive map was prematurely changed to the "NEW" server, before the server was ready. This change resulted in a loss of Certificate data entries from January 9, 2017 - January 17, 2017. During that time, 417 Certificates were printed and are now missing from the history stored in the Waters Certificate System. The "NEW" server has been updated, the drive map has been corrected and it is now necessary to manually enter the missing data to have a complete historical record. I have identified and listed the missing data in both; an excel spreadsheet (see page 2) and a pdf (see page 3). Please use the pdf to scan the barcode data into the system. | Bill confirmed that 100% of the missing 418 certificate records have been entered into the VENTURE COFP System | Wesley Wong | Desmond Ann | 357
171476 | Manufacturing Engineering | WAT045248 | 2017-02-15 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-04-25 | We conducted a compliance audit with Flex and would like to make the following amendments to the ASP for more clarity. Below is the proposed change to the this ASP WAT045248. 1) Add p/n WAT073289 for lithium grease. | Feb 22, 2017: If ECRs written from WPPL, who will be the reviewer from ME for approval. Request for list of Reviewers for Product Lines. | Roland Poisson |  | 357
171477 | Manufacturing Engineering | 289001994 | 2017-02-15 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-31 | We conducted a compliance audit with Flex and would like to make the following amendments to the ASP for more clarity. Below is the proposed change to the this ASP 289001994. 1) Pg 2 of 6 listed methanol in the tool list, but in operation 1.0 and 1.1, there is no usage of methanol in the assembly process. Please remove the methanol from the tool list. 2) Pg 4 of 6 listed 119850 spacer press tool in the tool list, but Flex do not have this tool and was not used for operation 1.3. Please remove the tool 119850 from the tool list. | Feb 22, 2017: If ECRs written from WPPL, who will be the reviewer from ME for approval. WPPL requests for list of Reviewers for Product Lines. May 30, 2017: WPPL will submit ECR and use the list of Approvers provided by Ken (via. Jaaffar.) | Roland Poisson |  | 357
171478 | Manufacturing Engineering | 279000925 | 2017-02-15 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-31 | We conducted a compliance audit with Flex and would like to make the following amendments to the ASP for more clarity. Below is the proposed change to the this ASP 279000925. - Pg 4 of 11, remove 999029150 fixture from tool list, it is not used in page 4, it is used in page 3. - Pg 5 of 11, add WAT027033 cable tie, qty 1 to the part list. - Pg 6 of 11, remove the last sentence "Use cable tie to secure the cable onto the bracket as shown. and WAT027033 cable tie, qty 1 from the part list." as this is already performed in operation 3.1 page 5 of 11. - Pg 7 of 11, add Tweezer (Optional) to part list. - Pg 10 of 11, add operation 6.2, turn the injector assy with tubing facing you and bent tubings per photo. We will raise the ECR once the changes is accepted by Milford. | Feb 22, 2017: If ECRs written from WPPL, who will be the reviewer from ME for approval. WPPL requests for list of Reviewers for Product Lines. May 30, 2017: WPPL will submit ECR and use the list of Approvers provided by Ken (via. Jaaffar.) | Roland Poisson |  | 357
171505 | Test Engineering | 735000431 | 2017-02-21 | TE82520 | N/A | International | Wesley Wong | CLOSED |  | 2017-11-27 | Proposed to add below in TSP, 735000431_RevG i) - Pg 6 of 6, add step 5.6.2 if carousel is not aligned, repeat step 5.3 to step 5.6. | 21 APR 2017 Recommend reassignment to Chau, I am pretty sure he is currently working on and responsible for this test system. May 30, 2017: Chau will review and follow up. June 27, 2017: There is no process change because operators has been doing it, so no validation required. TSP will be revised to include such clarification steps for both Carousel and Carrier alignment. July 27, 2017: WPPL is reviewing the TSP, and Chau to release into SAP. Aug 30, 2017: WPPL is okay with changes in TSP, and ECR to be submitted. Sept 26, 2017: Chau is working on the ECR. Oct 24, 2017: ECR 9218806 out for approval. Nov 27, 2017: Approved and effective on Nov 10, 2017. | Chau Nguyen |  | 357
171507 | Test Engineering | 735003390 | 2017-02-21 | TE83084 | N/A | International | Wesley Wong | CLOSED |  | 2017-03-23 | Please add below on coversheet to include part number that is tested in Venture, 735003390_RevA. i) From "Part Number Affected: 279001730" to "Part Number Affected: 279001730, 279003265" | the TSP has been updated to add the altus part number ECR# 9214171 is just waiting on one signature to be released we do not need a qualification no change was made to the procedure just added the altus part number to the cover sheet | Paul Deschenes |  | 357
171508 | Manufacturing Engineering | NA | 2017-02-21 | 999028068 | NA | International | Desmond Ann | CLOSED |  | 2017-03-13 | GES feedback that there is a typo error in the Part List on DWG No: 999028068, Actuator Torque Housing Fix. –VISTA- Item 4, Part Number Should be 999042075 and not 999042074. Please see attached. | Corrected typo on drawing, see ECO9214184 for correction. SW 3-Mar-2017 | Stephen Wright |  | 357
171509 | Test Engineering | N/A | 2017-02-21 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-04-24 | There was an ECR 9213478 raised that will result in a design change of Assy, PCB, 2690 Fluidic Rohs (210000416) from rev H to J. ECO team in WPPL received a response from Bryan Deguire that FA articles are not needed in Senai-Flex and it was evaluated by Waters Test Engineering team with fixtures in Milford Could i ask for a clarification for the above that this is the understanding? | Apr 24, 2017: Assigned to Paul to follow up with Wesley. May 30, 2017: Desmond to review and clarify to reopen or keep it closed since Paul sent an email to clarify Wesley's concerns and got no response. | Paul Deschenes |  | 357
171519 | Test Engineering | 735002010 | 2017-02-24 | TE82943 | N/A | International | Wesley Wong | CLOSED |  | 2017-03-29 | Proposed to change below section in TSP, 735002010_revG a)On Pg7, step V: - Change "password=tester" to "password=TE82943-X", where X is the pc number, to allow current practice. OR - Remove the login credentials information | TSP 735002010 updated, approved, RG, to Rev.H. (removed references to login and password). March 2017 | Edward Jackson |  | 357
171521 | Test Engineering | 735004006 | 2017-02-24 | TE82988 | 02 | International | Wesley Wong | CLOSED |  | 2017-08-30 | Problem Statement: GTS Self-Test@Flex-M'Sia - Marginal Fail on Test Pickering Relax MUX Findings: 1) Ran rev K self test program x3 times, all failed to Test Pickering Relay Mux 2) Attached failed reports, with marginal readings averaging to 2.7ohm, (passing limit < 2ohm) | Apr 24, 2017: Desmond to take over the request? April 26, 2017: Desmond will investigate if problem still exists. May 30, 2017: Desmond will provide feedback. June 28, 2017: WPPL/CM is waiting for parts to repair and then test to confirm the problem, and will notify Milford. Aug 30, 2017: Notified to close by WPPL, and no qualification required. | Desmond Ann |  | 357
171540 | Test Engineering | 735000698 | 2017-02-28 | TE82774 | NA | International | Yeap Siew Yee | CLOSED |  | 2017-10-19 | Description: To change the sentence below, as per documented in TSP Doc No: 735000698 Rev L 1) Page 3 of 16: To change 'HP laser jet printer' to 'printer' Justification: Samsung printer is currently used in Venture/GES Singapore instead of HP printer | Apr 24, 2017: Will combine this change with introduction of new P/N instrument to TSP. Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI Oct. 2017: 735000698 Revision M corrected OESR 171540. | Edward Jackson |  | 357
171541 | Test Engineering | 735004052 | 2017-02-28 | TE82733 | N/A | International | Wesley Wong | CLOSED |  | 2017-03-13 | Proposed to change below section in CAL document 735004052_RevA, i) On 1.1, Change statement "..module into 110 volt outlet" to "..module into AC outlet" This is to make it generic as Singapore AC supplies are at 240V, instead of 110V in Americas. | The procedure was changed and the new version is released in sap Change #9214398 | Paul Deschenes |  | 357
171542 | Manufacturing Engineering | 279001990 | 2017-02-28 | 999027058 | NA | International | Desmond Ann | CLOSED |  | 2017-10-03 | While reviewing the 279001990 ASP, we saw that there is an addition of the tool 999027058 in page 14 of 42 operation 2.8 during ECO9208661 to change the ASP from 279001990 Rev AB to AC. Currently GES is using hand to insert the cables to the AC Power Receptacle and we do not have the drawing for 999027058. Please help to send us the tool 999027058 and put tool 999027058 as optional. See attached for details. | Created drawing file Mar 2017 sent to Desmond to have parts fabricated in Singapore. Waiting for feed back. June 28, 2017: WPPL placed order for parts being made in US with 2-3 month leadtime. (Goind forward, Milford would consider ordering parts to ship to WPPL to facilitate the process on case-by-case basis based on production urgency.) August 30, 2017: GES received the tool. Waiting for Steve to release dwg into SAP. Apr 25, 2018: ECO # 9218615 | Stephen Wright |  | 357
171554 | Test Engineering | N/A | 2017-03-06 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-03-06 | With regards to ECR 9213583 attached, there is a requirement to change the printout format (WAT024406) on section D for the instrument serial labels. We would like to check with Milford team, if there are plans to modify the template on the serial printing stations at all CMs? | No Qualification is necessary due to no action taken. ECR 9213583 is currently "Submitted to QC". Once the ECR is Approved, Released and converted to an ECO then the label templates will ALL be revised to reflect the changes noted. The change will impact labeling both; here in Milford & CM's in Singapore. I ( Bill Irons ) have participated in the proposed changes and will phase-in the changes to correspond with the ECO effectivity date. ( ECR 9213583 attached; see Pg. 02 ) | Bill Irons |  | 357
171557 | Test Engineering | N/A | 2017-03-07 | N/A | N/A | International | Bill Irons | CLOSED |  | 2017-06-15 | Having successfully installed the "NEW" server and transitioned same, it is now time to retire the "OLD" server and ship it back to Milford. Please power down, disconnect, package and ship the "OLD" server to the following address: Waters Corporation 34 Maple Street Milford, MA USA 01757 Attn: Bill Irons M/S: REL | Apr 24, 2017: Desmond to take over the request. Apr 26, 2017: Bill will close the request when old server received. 20170615 Server has been received. No back-up tapes were returned. Bill Irons | Wesley Wong | Desmond Ann | 357
171573 | Manufacturing Engineering | 279001841 | 2017-03-13 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-03-15 | To add in SigmaSure7 genealogy for new records in the below, i) Select Top level Product -> 186015042.. (in get assembly process steps) ii) Select Assembly Process Steps -> Build 186015042 ... -> Replace 210000493 ... | The options are now available. | Roland Poisson |  | 357
171582 | Test Engineering | N/A | 2017-03-17 | N/A | N/A | International | Wesley Wong | CLOSED |  | 2017-03-20 | CEI Employee Account Info: Yong Fook Cheong ( YongFC@cei.com.sg) Please add below operator id in drop down list when creating/replacing records: - XEM13293 | Added operator 'XEM 13293' to operator table. | Bill Bates |  | 357
171583 | Test Engineering | 735002510 | 2017-03-17 | TE82865 | N/A | International | Wesley Wong | CLOSED |  | 2018-07-02 | Proposed to change below section in TSP, 735002510, Rev D: a)On Pg3, step1.3: - Change "password=tester" to "password=TE82865-X", where X is the pc number, to allow current practice. OR - Remove the login/password credentials information | Apr 24, 2017: Login credentials are to be removed from all test procedures, per TE policy. May 30, 2017: Will close if WPPL has no further questions. | Richard Davis |  | 357
171588 | Test Engineering | 735002130 | 2017-03-20 | N/A | N/A | International | Yeap Siew Yee | CLOSED |  | 2017-03-21 | Propose to change the following in TSP 735002130 Rev B: (i) Test step 3.7 (Page 3 of 6): To change software name 'ACQUITY TUVII Detector.sft' to 'ACQUITY TUV Detector.sft' | No change made to TSP as it correctly states name of file. TUVII is the correct description for Acquity TUV detector that is currently being manufactured. I verified that the firmware file on download PC is correct and the latest rev. I renamed th file to match TSP and SAP description. No qualification necessary as their was no change to tester hardware or process. See attached screen shot from SAP that will show file name as TUVII firmware. | Mark Jussaume |  | 357
171590 | Test Engineering | N/A | 2017-03-20 | N/A | N/A | International | Yeap Siew Yee | CLOSED |  | 2017-03-22 | Problem Statement: Database Connection Failure/ Node configuration Failure on SQT station Flex is experiencing database connection failure on SQT station (IP add: 10.249.26.90) Refer to screenshots for error messages. | Connected to system and Empower seems to be running without issue. If error recurs, please submit another OESR, but it seems fine for now. No qualification required: No changes were made, so there is nothing qualify. | Michael Cotter |  | 357
171602 | Test Engineering | WAT024406 | 2017-03-24 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-03-28 | GES have feedback that the new template of the physical printout of the serial number label WAT024406 is different for the drawing. Please see email for more details and correct the template. | No qualification is required because no action was taken. The printout "difference" is not sustained or reproducible. Operators have been instructed to print preview and monitor label creation. If problem returns, they will raise another OESR. | Bill Irons |  | 357
171621 | Test Engineering | WAT271129 | 2017-04-03 | TE82736 | - | International | Mike Cotter | CLOSED |  | 2017-05-30 | Please provide the System Leak Test results (spec: <= 60 PSI) for the last 50 instruments shipped. Thanks, Mike | April 26, 2017: Mike will close when data received. May 30, 2017: Closed. Only collect data so no qualification required. | Desmond Ann | Siew Yee Yeap | 357
171628 | Test Engineering | NA | 2017-04-05 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-04-05 | GES feedback they have made an error in printing the PKI certificate. Missing an “A” in the Model description. The unit serial numbers affected are: C17P10021G &C17P10024G. See attachment for details. Will require Milford to help to change it at the database for GES to reprint it. | Qualification is not required because the two (2) certificate records that had data entry errors, have now been moved to the "INACTIVE" table. Allowing GES to create "NEW" records for those instruments. Both records have been moved to the "INACTIVE" table. GES can now print "NEW" Certificates of Functional Performance" for the following serial numbers: C17P10021G C17P10024G | Bill Irons |  | 357
171645 | Test Engineering | N/A | 2017-04-11 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-04-11 | Flex feedback that some of the Flex technician forgot password and some of them encounter after change the password few month later unable to log-in. Will require Milford to reset the password for attached name list. Thank you. | all passwords for employees on attached list have been changed to 'waters'. This was changed in SS6 only. An additional list of reset passwords and inactivate requests came through on 19-Apr-2017 and were reset/inactivated in SS7. Note that any reset password requests in SS7 will get an email explaining how to do so. No qualification needed. Closed 11-Apr but editted here on 21-Apr to include the SS7 updates. (BB) | Bill Bates | Michael Cotter | 357
171646 | Test Engineering | NA | 2017-04-11 | 999037997 | NA | International | Yeap Siew Yee | CLOSED |  | 2017-04-28 | Replacement needed for Thumb Screw used on Flywheel (P/N 999037997) due to wear and tear after 3 years usage. | April 26, 2017: Mike is getting the screws made in shop. April 28, 2017: Two replacement parts were shipped on April 28th. These are identical (like-for-like) replacements. | Michael Cotter |  | 357
171669 | Test Engineering | 735000887 | 2017-04-21 | TE82817 | N/A | International | Desmond Ann | CLOSED |  | 2017-05-24 | While reviewing the TSP 735000887, we observe the following: 1) Step 1.7 to change to “Plug tester into 120/ 240 volt AC outlet.” 2) Step 1.9 to change to “On desktop, double click on appropriate icon “TE82817 PCM II” or “TE82817 SFO”.” 3) Step 4 : The tester is not able to perform firmware download from the main menu. GES is using the loader to download the firmware. See attached for details. Please update the TSP and resolve the firmware download issue. | The instrument serial number is entered when the firmware is downloaded using Waters loader program. The TSP has been updated to reflect this. It has been approved and released on ECR 9215744. Qualification not required as there was no change to test or product. | Mark Jussaume |  | 357
171670 | Test Engineering | 735000888 | 2017-04-21 | TE82818 | N/A | International | Desmond Ann | CLOSED |  | 2017-05-24 | While reviewing the TSP 735000888, we observe the following: 1) Step 1.7 to change to “Plug tester into 120/ 240 volt AC outlet.” 2) Step 4 : The tester is not able to perform firmware download from the main menu. GES is using the loader to download the firmware. See attached for details. Please help to update the TSP and resolve the firmware download issue. GES is currently using Loader.exe from desktop to download the firmware into UUT. See attached excel for picture illustration on the steps performed. | The instrument serial number is entered when the firmware is downloaded using Waters loader program. The TSP has been updated to reflect this. It has been approved and released on ECR 9215744. Qualification not required as there was no change to test or product. | Mark Jussaume |  | 357
171689 | Test Engineering | 735002271 | 2017-04-26 | TE82995 | NA | International | Yeap Siew Yee | CLOSED |  | 2018-07-02 | Proposed to change the below in TSP 735002271 Rev G: 1) Page 3 of 4, section 1.0, step 1.4: To change window password from 'tester' to 'TE82995-#' , # refer to PC number or to remove login password from TSP | This change will be incorporated when document is revised. See OESR 171440. | Richard Davis |  | 357
171715 | Manufacturing Engineering | 735003655 | 2017-05-05 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-16 | Flex feedback that for ASP 735003655 Instrument Serial Number Printing -STD 1. Oper 1.3 instruction is for Milford plant only, Propose to remove "Singapore" for instruction icon "Singaporestandard.lbl" in Flex serial number station. Refer to attached picture. Only Flex icon is different, CEI and GES icon do not have the "Singapore" on the icon. 2. Oper 3.2 instruction to enter Part number,Mfr code and Factory order number where as physcial (screen) prompts to enter only partnumber. Proposed to update work instruction to "Not Applicable for Singapore Site) for Mfr code and Factory order number. Refer to attached picture1. | The query for the OESR have already been answered to go by the ECR process where WPPL will update the procedure and raise the ECR for approval. | Desmond Ann |  | 357
171718 | Manufacturing Engineering | 735003656 | 2017-05-05 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-16 | Flex feedback for ASP 735003656 Instrument Serial Number Printing -MAN 1. Oper 1.3 instruction is for Milford plant only. Propose to remove "Singapore" for instruction icon "Singaporestandard.lbl" in Flex serial number station to be consistent and same as other 2 site (GES and CEI). Refer to attached picture. Only Flex icon is different, CEI and GES icon do not have the "Singapore" on the icon. 2. Oper 3.2 instruction to enter Part number, Factory order and Serial number where as physical (screen) prompts to enter only part number and serial number. Proposed to update work instruction to "Not Applicable for Singapore Site" for Factory order number. Refer to attached picture. | The query for the OESR have already been answered to go by the ECR process where WPPL will update the procedure and raise the ECR for approval. | Desmond Ann |  | 357
171722 | Manufacturing Engineering | WAT077886SO | 2017-05-08 | TE82625 | 1 | International | Desmond Ann | CLOSED |  | 2018-02-14 | In ASP WAT045248 REV L, step 6.2, it is stated that "TEST COMPLETED ASSEMBLY IN ACCORDANCE WITH SOP WAT077886SO (SAMPLING INSPECTION/ TESTING). USE TEST PROCEDURE 735000287." But the procedure WAT077886SO is already obsolete. Please help to reset the procedure WAT077886SO back to active. See attached for more details. | July 27, 2017: Desmond to confirm of taking over the request from Mike. August 30, 2017: 999119592 dwg to be released by Steve into SAP, and PRO doc can be completed. Feb 14, 2018: Desmond took this over and processed ECR 9217360 to adjust instruction detail within assembly PRO document WAT045248. His change became effective on 01/11/2018. | Mike Aubin |  | 357
171748 | Test Engineering | WAT024406 | 2017-05-19 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-05-31 | We are seeing intermittent errors in the instrument S/N label printout in GES. There is also a field complaint being feedback on the error in the S/N label. There is also an OESR 171602 previously in Mar 17. Yesterday, we encountered again on the error of the printout. Below are some product serial number that are printed with the wrong format. 1: 186179100. Unit serial number: E17SMC088G TO E17SMC139G. 2: 186001791. Unit serial number: E17SMC086G TO E17SMC087G. 3: 186002449. Unit serial number: E17PC2394G TO E17PC2397G. 4: 186015011. Unit serial number: E17UPX319G TO E17UPX330G. 5: 186017008. Unit serial number: E17VCH639G TO E17VCH646. 6: 186017009. Unit serial number: E17VHC920G TO E17VHC931G. 7: 186015029. Unit serial number: D17UPF065G TO D17UPF074G. 8: 186247500. Unit serial number: D17475564G TO D17475574G. Will require Milford to help check and advise. See attached on the history of OESR 171602 and field feedback on the printout error. Samples of the error printout are attached in email. | No qualification is necessary, the equipment was NOT used as intended, and created unsatisfactory results. As previously suspected on OESR 171602, and now confirmed by testing for OESR 171748. The issue of bitmap images being misplaced on the label, is 100% attributable to having multiple label templates open simultaneously. It was recommended in March to instruct Venture operators to open one (1) label template at a time. Unfortunately, that recommendation was obviously not followed. | Bill Irons |  | 357
171760 | Test Engineering | 735001613 | 2017-05-26 | TE82882 | 2 | International | Desmond Ann | CLOSED |  | 2017-09-06 | Flex informed that the tester TE82882-2 testing for Spare OB PN 700002787 for old PDA, when using Golden Unit SN : E17UPD00002 and TE82882-2, the UUT PASS but Not able to print test report. Problem : TOPAZ Signature Pad - not able to detect by TE82882-2 and unable to do electronics signature. Action Taken: a. Restart CPU - retest and UUT PASS but not able to sign-off b. Power OFF CPU - unplug and plug-in TOPAZ signature pad - Power ON CPU , retest and UUT PASS but not able to sign-off. Will require Milford to help to check and advise on the issue. TE82882 IP Address : 10.249.26.78 is online for Milford access. | May 30, 2017: Eric is reviewing the issue. June 28, 2017: I (Eric) think we can close this - Flex is using PDA II chassis to test bench per procedure so they can use newer test software as indicated on procedure without sig pad. Desmond will check and notify to close. Aug 30, 2017: Notified to close by WPPL, and no qualification required. | Eric Turner |  | 357
171763 | Test Engineering | NA | 2017-05-30 | TE82943 | 1 | International | Yeap Siew Yee | CLOSED |  | 2017-05-30 | Problem statement: Error message 'Instrument error. Invalid IP address' for tester TE82943-1 Problem Description: During test steps "Test Loader Download". Monitor screen show "Instrument error. Invalid IP address", testing cannot continue. Action taken: 1) Remove existing LAN cable connecting from PC work station to UUT and replaced by 2 new LAN cable to test, result still FAIL. Unplug the UUT and move to other Test Station to test, result PASS without changing any part from the UUT. Use another new UUT to test, result still FAILS. Move the UUT to the other test station to test, result PASS. Suspect the Intel® 82567LM-3 Gigabit Network card is faulty. PC work station model is LAC/E 32 Acquisition Server. 2) Exchanged network card with another identical network card from spare PC. Same error message was observed. | The scheduled task that clears IP addresses was not running. It was configured to run as TE-1, and it has not run since the password was changed. I configured it to run as System* and it should be fine now. *All scheduled tasks should run as System, otherwise they'll cease to run when a computer is renamed or a password changed. | Michael Cotter |  | 357
171778 | Test Engineering | 735003655 | 2017-06-06 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-06-06 | WATERSISNL1 PC work station locked by administrator, cannot login. Attach print screen capture (in email.) Refer to OESR Pg. 02 | ALL SET ! Logged the administrator account out. We have been performing patches to the OS, due to the recent world wide virus outbreak. That PC must not have rebooted properly. | Bill Irons |  | 357
171780 | Test Engineering | NA | 2017-06-07 | TE83013 | 1 | International | Yeap Siew Yee | CLOSED |  | 2017-06-07 | Problem Statement: Error message was observed (Refer to screenshots) when trying to export test data for 186015042 from sigmasure 6. | This was caused by an issue with the Sigmasure server that has been corrected. BSG/IT request ID AMXP2L addressed this issue. To verify, WPPL confirmed with the CM that they are now able to export the data. | Bill Bates | Matt Brule | 357
171781 | Test Engineering | NA | 2017-06-07 | TE82882 | 2 | International | Yeap Siew Yee | CLOSED |  | 2017-06-09 | Problem Statement: SigmaSure 6 not able to show test results for TE82882-2 (PDA tester). Refer to screenshot. Will need data from 16Jan2016 to recent date for Acquity PDA Final Test System TE82882-2 | BT group fixed database connection problem and restored application. No qualification is required as this is a run/not run situation. It is up and running per BT request AMXP2L | Eric Turner |  | 357
171782 | Test Engineering | 735001613 | 2017-06-07 | TE82882 | 2 | International | Desmond Ann | CLOSED |  | 2017-09-06 | While going thru the issue on the Topaz signature pad issue, it was found that the TSP 735001613 Rev T stated the spare optic bench 700002787 as one of the part number affected in the TSP and the test software stated in the TSP is 3.2 but the tester TE82882-2 is using test software rev 2.0 to test the spare optics bench 700002787. Attached is the screenshot of the tester and the TSP and also the history of the communication on this issue. | June 28, 2017: I (Eric) think we can close this - Flex is using PDA II chassis to test bench per procedure so they can use newer test software as indicated on procedure without sig pad. Desmond will check and notify to close. Aug 30, 2017: Notified to close by WPPL, and no qualification required. | Eric Turner | Tyler Modelski | 357
171797 | Test Engineering | NA | 2017-06-13 | N/A | N/A | International | Yeap Siew Yee | CLOSED |  | 2017-08-07 | Problem Statement: Test data files in 'SigmaQueue & SendBuffer' folders are not sent out: 1) TE83013 -1 2) TE83013 -2 3) TE82943-1 4) TE82943- 19 5) TE82943- 21 6) TE82943- 22 and 7) TE82943- 24. | June 28, 2017: Reassigned to Bill Bates and he is investigating the issue. ET These were all addressed more recently in OESR# 171908, except for TE82943-1 - this SENDBUFFER folder has nothing in it currently. OK to close. No qualification required to address Sigmasure issues. BB | Eric Turner | Bill Bates | 357
171801 | Test Engineering | N/A | 2017-06-14 | TE82743 | 3 | International | Desmond Ann | CLOSED |  | 2017-06-21 | GES feedback that they are having issue on the 21000021820 test fixture. When the board is loaded, it have a constant sound when the test start instead of beep and it is showing the error msg as attached. Fixture Details: FCT tester ( RACAL FUNCTIONAL TEST SYSTEM - TE82743-3 ), the test fixture was sharing between two models, for PCBA 2100000460 using W21000021820 test program and for PCBA 210000456 using W210000218 test program. The two model using same test fixture but using different CPU module. Verification: 1) We tried using the golden board and encounter the same failure. 2) We tried running PCBA 210000456 using W210000218 test program and the CPU module for testing 210000456. The test passed. 3) We measure the below : - continuity of the cable connection between the fixture and the CPU board - passed. - continuity of the serial loopback continuity and resistance - passed. We suspect the failure is due to the CPU module use for testing PCBA 2100000460 using W21000021820 test program. Please help to advise if there are any other areas we can troubleshoot or if you can send us 2 pcs of spare CPU module. | Supplied new 182 PCB with TCM firmware. Supplied new rs232 cable. Supplier confirmed test is working | Mark Jussaume |  | 357
171802 | Test Engineering | N/A | 2017-06-14 | TE82865 | 6 | International | Desmond Ann | CLOSED |  | 2017-06-14 | CEI reported that the tester TE82865-6 is unable to download the firmware and it shows on the note as Invalid Ethernet address. We have tried 2 units on this tester it still show the same error. We have tried to verify the IP address with the other tester. The IP address is the same as TE82865-5 and we can see the Lan Cable is connected. Please help to advise. | The system ran out of IP addresses because the task that clears them was not running. The task didn't run because it was configured to run as user te-1. When the te-1 password changed, the task stopped functioning. The task has been updated and is running now, addresses have been cleared, etc. It should be all set. As a note. tasks should all be configured to run as system. Configured in this manner they will continue to run, even if the user password(s) change. This should have been done when the password was changed, but we must've missed that step on this one. Nearly all systems were updated properly. This must have just been an exception. | Michael Cotter |  | 357
171810 | Test Engineering | WAT024406 | 2017-06-20 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-06-21 | CEI feedback that they found the ASO serial no. printing, there was a skip from E17UPO176H ( printed out on 30th May 2017) to F17UPO181H (printed out on 20 Jun 2017). In between 177H to 180H serial no. has confirmed with product leader that he did not printed out at all. Please help to advise. | No action taken. The serial number counter for Model Code = UPO is currently set to print F17UPO183H. At least two (2) labels (181 & 182) have been printed since the discovery of the missing serial numbers: 177, 178, 179, 180. The missing labels are unaccounted for. NOTE: On occasion, serial number incremental gaps may occur for various reasons such as: 1. Operator printed more labels than needed and discarded extra. 2. Printer ran out of Label or printer ribbon. 3. Labels were mis-printed and discarded 4. etc. | Bill Irons |  | 357
171816 | Manufacturing Engineering | N/A | 2017-06-21 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-06-21 | ECO 9213776 have changed the PCB 210000420 to 210001161 for CPSO and ASO assembly, but the sigmasure for 186015014, 186015020, 186015020IVD and 186015021 are not updated yet. Please help to update the sigmasure genealogy for 186015014, 186015020, 186015020IVD and 186015021 to change 210000420 to 210001161. | Complete and no validation is required. | Roland Poisson |  | 357
171849 | Test Engineering | 735002984 | 2017-06-30 | TE83064 | NA | International | Yeap Siew Yee | CLOSED |  | 2017-07-28 | Problem statement: FTN unit serial number not reflected on the footer of the test records. Refer to attached test records for reference. | Test Stations were updated to latest test version which addressed this issue. Monitor and close if no issues found. | Richard Davis |  | 357
171850 | Manufacturing Engineering | N/A | 2017-06-30 | 999043622 | N/A | International | Desmond Ann | CLOSED |  | 2018-01-16 | We have received the release drawing for 999043622, but found that the screw for item 4 is countersunk and not flat screw as stated in drawing. There is also an item which is listed in the drawing. Will require Milford to update the drawing and create a drawing for the part. | August 30, 2017: Waiting for drawing 999043622 to be updated. released under ECO9218648 | Stephen Wright | Mike Aubin | 357
171853 | Test Engineering | N/A | 2017-07-05 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-09-20 | (Prior updates available in prior monthly meetings' minutes.) Jan 24, 2017: Parts and instructions will be sent to WPPL this week (Jan2017) for upgrades. Feb 21, 2017: Boxes are modified and being calibrated, getting quotes for a cable improvement for WPPL. March 28, 2017: Procedure in process, need to make sure all test boxes are calibrated and ready to go, and decide on a logging scheme. April 24, 2017: Siew Yee to provide status of calibration completion of test boxes before Matt can implement the updates. April 2, 2017: Calibration completed at Venture, CEI. Flex should be done in first week of May. May 30, 2017: Matt is actively working on it. June 27, 2017: Still in process, procedure updates under way, still exploring best way to log daily test activities. WPPL suggests to consider moving it onto 2017 List. This OESR is created in place of OESR 140328 to close out 2014 List. | 27-JUL-2017 Procedure updates in process 25-SEP-2017 Ordering tags for tracking the testing 30 JAN 2018, have tags and will be completing procedure and roll out soon 27-FEB-2018 Look for procedure update in March 2018 24-APR Labels delivered, hardware in place, ECR in system, will be approved mid to late may after Milford ISO audit 25, June 2018, had to hold off again due to FDA inspection, putting together training package so that both Milford and WPPL can complete training to new procedure, then will release. Expect to complete training materials this week. 20, AUG 2018 training materials will be sent soon. 24-SEP-2018 Closed and implemented. No verification required as this IS a verification daily of production test equipment using calibrated equipment | Matt Brule |  | 357
171854 | Test Engineering | N/A | 2017-07-07 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-07-12 | Existing printer for WATERSCOFP1 HP Laser Jet Enterprise M604 (unknown symbols is seen on the right of the certificate and continuous printing of more than 1 copy is coming out from the printer) 1. Re-start system/printer and encounter (Problem occur 2 hour later) 2. Printer Test Page (No issue - Refer to attachment) 3. Try to print from notepad ( No issue ) | I suspected that the intermittent "unknown symbols" are the result of either a malfunctioning printer, or a corrupted printer driver. After remotely rebooting the PC, I was unable to print anything ( errors ). I used Windows XP control panel "Printers and Faxes" to administratively remove the printer driver. When I attempted to re-istall the driver, the software could not find the printer. The printer was powered OFF. After emails requesting that the printer be powered ON, I was able to re-install the printer driver for the HP LaserJet M604 PCL.6. Successfully performed a couple of test prints. Email feedback from Flextronics & WPPL indicate that the system is now functioning correctly. I will attach the email confirmation on Pg. 02 & 03 of this document. | Bill Irons |  | 357
171855 | Test Engineering | N/A | 2017-07-07 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-08-10 | Flex reported that the Sigmasure data encounter addition text. ASP 279001834 Rev N Instrument Serial Number F17QSM487A Phyiscal scan part number 289003741 serial number 3741UEWW Print-out report show same part number with 2 different serial. 270517453M Part Number 289003741 3741UEWW Part Number 289003741 Refer to attached picture | Sigmasure data cannot be changed once saved, this can only be red-lined on the DHR documentation. No qualification needed as nothing can be done at this point to Sigmasure. August 30, 2017: Ask Tyler to take a quick look when updating SigmaSure | Eric Turner | Bill Bates | 357
171873 | Test Engineering | N/A | 2017-07-14 | TE82625 | N/A | International | Desmond Ann | CLOSED |  | 2017-08-28 | Due to the obsolete document of the WAT077886SO, there is a capacity crunch for the HPMV tester TE 82625 in Alliance production line. All HPMV sub assembly will have to be tested 100% instead of sampling before assemble into the instrument. We will need to request for a new HPMV tester to support production. Currently production is running daily overtime to support the production numbers. Do advise me the leadtime, cost and BOM and I can raise a CEA for the tester. With reference to Alliance and HPMV UUT Forecast until end of the year 2017, below are daily balance loading requirement: Alliance = 20 unit * 2 HPMV = 40 HPMV’s SPARE WAT-S-WAT045438 = 8 HPMV’s Total Daily HPMV’s = 48 units Test Station Capacity = Test Time (10 mins) / 90% (FPY) = ~11 mins Unit per Hour = 60 mins/11 mins = ~5 units of HPMV Total Daily Tester Capacity = 7.5 Hrs * 5 = ~38 units per day Flex will build 10 units during OT as tester utilization is now overcapacity at 126.32% | No change to Test system. All production HMPV get tested at final instrument test. All spares tested at sub assembly level This gives 100% test coverage for all valves. | Mark Jussaume |  | 357
171891 | Test Engineering | N/A | 2017-07-20 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-07-20 | Flex feedback that while printing certificate description / Serial with misaligned as shown below :- Please assist to re-align the format highlight in red. Thank you. Description :-2998 Photodiode Array Det Serial : G17998508 See attached for photo of Certificate of Functional Performance. | No qualification required. Operator error. I have assigned the misprinted certificate as "INACTIVE". Please instruct FLEX to reprint a correct certificate for serial number = G17998508 email chain attached on Pg.02 | Bill Irons |  | 357
171892 | Test Engineering | N/A | 2017-07-20 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-09-30 | GES feedback that for the testing on this PCBA # W1210000204 in Racal Tester but they are not able to test because an error seen during software loading. They are suspecting that one of the electronic load module # 4 ( model #: 66103A ) is faulty resulted to not able to complete the initialization. Please kindly help to check and advise if you have spare part to replace the said module. Attached is the screenshot of the error for your reference. | Paul D. sent replacement electronic load modules. Tester repaired. One for one replacement. No validation required. | Edward Jackson |  | 357
171896 | Test Engineering | N/A | 2017-07-20 | N/A | N/A | International | Bill Irons | CLOSED |  | 2017-08-25 | There are several errors recorded in the FLEX-Singapore Certificate of Functional Performance system. Need WPPL to sort out mistakes and recommend corrective action(s). I have attached a table on page 02 and an Excel file on page 03 of this document, which outlines the errors discovered. | If there is any error during the Certificate Printing, where the SERIAL NUMBER was the actual ERROR. and a different SERIAL NUMBER ( the correct one ) was used to print a NEW certificate, they will have to notify WPPL for you (Bill) to help INACTIVATE the record with the error. No qualification is required because of no change in process or software. | Desmond Ann |  | 357
171908 | Test Engineering | NA | 2017-07-26 | N/A | N/A | International | Yeap Siew Yee | CLOSED |  | 2017-08-03 | Problem statement: Data files pending in sigmaqueue/sendbuffer folder: Affected testers as below: -TE83026-1 -TE83026-2 -TE82956-2 -TE83013-1 -TE83013-2 -TE83013-2 -TE82943-19 -TE82943-24 -TE82943-21 -TE82943-22 | TE83026-1 - Matt (10.249.40.27) TE83026-2 - Matt (10.249.40.28) TE83013-1 - Matt (10.249.40.25) TE83013-2 - Matt (10.249.40.24) - are all covered under a separate OESR (#171694) and will be dealt with there. TE82956-2 - Rick (10.249.40.20) - fixed TE82943-19 - Ed (10.249.40.9) TE82943-24 - Ed (10.249.40.14) TE82943-21 - Ed (10.249.40.11) TE82943-22 - Ed (10.249.40.12) - fixed No qualification required for pushing files to Sigmasure. (BB) | Bill Bates | Eric Turner | 357
171945 | Manufacturing Engineering | N/A | 2017-08-08 | 999043134 | N/A | International | Desmond Ann | CLOSED |  | 2018-01-05 | Based on RZZ SPACER, 999043134 drawing. Our fixture supplier feedback that they have purchase the Shoulder Screw from CARR LANE, Part Number: CL-5-SS. Parts have arrival and perform measurement, the screw head size is at 0.1855”. Next he have to fabricate the CROSS BAR, 999043243. Base on drawing, there are 6 hole for the shoulder screw to pass through and tighten with JAM NUT 10-24. The drawing show hole size at 0.156” which is too small for the Shoulder Screw to pass through. See attached in email for more details. Will Milford be able to advise if there is a typo error on the drawing? | August 30, 2017: Steve acknowledges the issue and working on DWG updates. Drawing updated to correct the issue, .1855 diameter hole was opened up to .221 to accept 10-24 screw. See ECO 9218782 | Stephen Wright |  | 357
171946 | Test Engineering | 735001613 | 2017-08-08 | TE82882 | 14 | International | Desmond Ann | CLOSED |  | 2017-09-06 | 1. TE82882-14 (IP 10.249.26.85) having below failure: TE82882-14 position 2 - Error - Serial doesn't match. Pls retry to reprogram. In unsucessful again, reprogram manually. Serial Number Return: String Return: TE82882-14 postion 3. Error - 2147220989 occurred at Exception occures in MilleniumToolkit. Insrument: 2. Manually Clear DHCP Server 3. Reload each position 2 and 3 of 186299800 - TE82882-14 Post 2 - same error UUT SN: UUT SN H17998672: OB 31H6P033 TE82882-14 Post 3 - same error -2147220989 UUT SN H17998671A: OB 31H6P032 4. Check on RS232 and Ethernet Connectors, did not find any abnormality. 5. All IPN is resonding when press manual. The units are placed online on the tester. Please help to check and advise. | Aug 30, 2017: Flex has reported the issue intermittent, and will keep monitoring it. Closed and no qualification required. | Eric Turner |  | 357
171955 | Test Engineering | 735001613 | 2017-08-10 | TE82969 | 1 | International | Desmond Ann | CLOSED |  | 2017-08-14 | Flex feedback on the below: 1. The TE-1 and tester password was reset to the default password. 2. TE82969-1 Not able to "Launch Console" / "Not able to detect UUT. 3. Clear Waters DHCP Server and Stop Acquity Console result is the same. 4. Using stand alone "Acquity Console", its not able to detect UUT 186299800 during "Instrument Scan". 5. Using "Instrument Configuration', its not able to detect 2998/PDA Detector. 6. "PDALin Folder" had been added last 08 Aug 2017. Please see screenshots in email and help to advise. | Looks like some kind of disk repair was attempted as all current files were moved to drive E: and the files in C: were not complete. Test software is required to run from drive C so installed new copy to C; Ethernet card was found to be unconfigured so instrument communication was impossible. Configured card. Also, Waters Instrument control software for the PDA and 2998 was not on the C drive so installed ICS. Ran test software and showed to work correctly. Unit was run and a test report was generated to qualify. | Eric Turner |  | 357
171969 | Test Engineering | N/A | 2017-08-16 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-09-01 | CEI have feedback that the below items are out of specs and will require advise from Milford if these equipment are acceptable to be continue to be use for production. Does Milford have spares on the affected instrument to loan and we can return this equipment to Milford to help to perform repair as there is no repair support locally in Singapore. Currently production is on hold pending for the advise if the equipment can be continue to use for production. Please see attached on the certificate for the below items with some range out of spec: 3 certificates for Electronic load from Techsystem 2 certificates for Keithley 236 from Tektronix This 5 equipments are used in Racal situated in CEI Batam. Boards currently tested on the Racal:- 1) 210000853 REV C RAVEN_RF_DC_TOP_MS1 (RoHS) PCBA [ WA9930-0000-1695-G] 2) 210000855 REV C RF_DC_TOP_BRD (RoHS) PCBA [ WA9930-0000-1697-G] 3) 210000955 REV A RF DC TOP (RoHS) PCBA [ WA9930-0000-1795-G] 4) 210000941 REV A RF_DC_TOP (RoHS) PCBA [ WA9930-0000-1793-G] 5) 210000323 REV G PERS. 2200 AUTOSAMPLER (RoHS) PCBA [ WA9910-0000-3362-G] 6) 210000420 REV B PERSONALITY 2200 ELEVATOR (RoHS) PCBA [ WA9910-0000-3969-G] 7) 210000339 REV P PREP PUMP (RoHS) PCBA [ WA9910-0000-6471-G] 8) 210000365 REV K PREP PUMP 2545 (RoHS) PCBA [ WA9910-0000-6472-G] 9) 210000422 VISTA PUMP (RoHS) (210000308) [ WA9910-0000-6546-G] 10) 210000422 VISTA PUMP (RoHS) PCBA [ WA9910-0000-6430-G] Please help to advise if the affected range is used for testing the above models and any risk for the production to continue. | Aug 30, 2017: WPPL will ship parts to Paul for repairs. Referencing the CEI provided Excel shipment chart: The following boards are referenced: 210000422 210000323 210000339 210000365 210000420 The Racal instruments listed out of calibration, were the Keithley 236's and the HP Electronic load modules. (From your earlier email). Looking into each Racal Functional Board test program, we concluded that only the 210000420 is affected, by the HP Electronic Load. The other boards listed above do not use the electronic loads, and only use the Keithley 236 in constant voltage mode (which passed calibration). Looking deeper into the 210000420 board, the HP Electronic Load was used for simulating a solenoid and a Heater/Cooler load. The Calibration Report is missing the High Resistance Range results. However, if it is shown to be off as little (~<1%) as what is shown on the Low and Middle Ranges, then we are not concerned. The intent of the test is to provide a load; then the board is checked for it's ability to handle the current needed, to sustain the applied load. After reading the full reports, we can conclude now that the 210000420 board is not affected. Also, the RD DC TOP board, (210000853, 210000855, 210000955, 210000941), uses the Keithley 236, but does not use the current source mode, so this board is also not affected. Paul has sent spares to repair the Racal. | Edward Jackson | Paul Deschenes | 357
171977 | Manufacturing Engineering | 279003346 | 2017-08-21 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-12-06 | Flex feedback that the ASP 279003346 operation 1.3 ask to scan the flow cell serial number into SigmaQuest. But there is no field in the SigmaQuest to enter the serial number. See attached. | The field is now added. No qualification since there is no change in document, software, and hardware. | Matt Franklin |  | 357
171984 | Test Engineering | N/A | 2017-08-23 | TE82807 | N/A | International | Desmond Ann | CLOSED |  | 2017-08-25 | GES feedback that they are encountering failure at ( Nebulizer Test ) for PCBA 21000204 PCBA test ( TE82807 ). The tester looking a data after the solenoid turning on > 22 but it was failed because the tester measuring only -0.094. Will require Milford to help advise and provide schematic for troubleshooting. | Emailed all documentation for troubleshooting. WPPL/CM manage to make the nebulizer test and board passed after changing the CPU connector. | Richard Davis | Paul Deschenes | 357
171992 | Test Engineering | 279001834 | 2017-08-24 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-08-25 | Lable printer encounter missing Waters Logo . ASP 279001834 Rev N Instrument Serial Number F17QSM787A to F17QSM789A (Qty 3) Refer to attached picture (in email) | No qualification required. System restart corrected this one-off problem. <see page 2 of this OESR> | Bill Irons | Bill Bates | 357
171997 | Test Engineering | N/A | 2017-08-29 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2017-08-29 | GES reported that they have an issue on 1 piece of Certificate of Functional Performance. The correct serial number should be H17BUR770G but the print out copy is “H37BUR770G”. Please help to INACTIVE the unit serial number “H37BUR770G” in the record. | I have moved the H37BUR770G record to the INACTIVE table. <email attached to Page 02> | Bill Irons | Bill Bates | 357
172005 | Test Engineering | 735001613 | 2017-08-31 | TE82969 | 1 | International | Yeap Siew Yee | CLOSED |  | 2017-08-31 | Problem Statement: 1) TE82969-1 NOT able to detect UUT (18601.5032) during 'Instrument Scan' via 'Acquity Console'. Error message 'Communication Failure' was shown 2) Cleared Waters DHCP Server and Stop Acquity Console result is the same. 3) Checked cable connection (RS232 and internet connection), no issue found. 4) Tried using "Instrument Configuration', not able to detect 2998/PDA Detector. Refer to photo for failure screenshot. | They (Flex) shutdown the CPU and unplug all connecting cables and leave it for 10 mins. After that reinstalled all cables and power up the CPU. The issue go resolved and Tester is back to normal operation. No qualification is required since there is no changes in process, software, firmware, and hardware. | Siew Yee Yeap |  | 357
172006 | Test Engineering | N/A | 2017-08-31 | TE82988 | 3 | International | Desmond Ann | CLOSED |  | 2017-10-25 | GES reported that the Agilent 34411A 6(1/2) Digit Multimeter of the GTS is faulty. We check with the manufacturer and found that they do not support repair and only perform a 1 to 1 exchange with a refurbished unit. The current unit Agilent 34411A is obsoleted and it was replaced with 34465A Digital Multimeter, 6½ Digit. We have now 2 options either to change to a refurbished unit or to a new unit of 34465A Digital Multimeter, 6½ Digit. Please help to review and advise if there are any concerns to change 34411A 6(1/2) Digit Multimeter to 34465A Digital Multimeter, 6½ Digit. | Oct 25, 2017: WPPL purchased an exact multimeter refurbished from supplier to replace the faulty unit. No validation required because it is a direct replacement with no change in test process, software, and hardware. | Desmond Ann | Richard Davis | 357
172019 | Test Engineering | N/A | 2017-09-07 | N/A | N/A | International | Bill Irons | CLOSED |  | 2017-09-13 | We are currently unable to connect to the server @ 10.249.40.1 to perform required updates. Please check with Venture/GES and advise. **UPDATE: I am able to ping the sonicwall ( Firewall ) but still can NOT access the Server or network. **UPDATE: BT Request ID: AR6JSK Remote Singapore Server @ Contract Mfg: | Waters BT was able to restart the Venture Firewall and reconfigure to address connection problems. ( see page 2 of this OESR 172019 for email chain ) | Desmond Ann | Bill Irons | 357
172020 | Test Engineering | N/A | 2017-09-07 | N/A | N/A | International | Bill Irons | CLOSED |  | 2017-09-11 | We are currently unable to connect to the server @ 10.249.55.1 to perform required updates. Please check with CEI and advise. **UPDATE: Network service was restored on Sep. 8, 2017, Database updated. This OESR 172020 can be closed. ~Bill Irons | Network service was restored on Sep. 8, 2017, Database updated. This OESR 172020 can be closed. | Desmond Ann |  | 357
172050 | Test Engineering | 735003877 | 2017-09-20 | TE82903 | NA | International | Desmond Ann | CLOSED |  | 2017-10-26 | Flex feedback that they are building the spare 700009348 which is the optic bench 279002656 that is reporting to 279002733 (for 2489X or 186248900) Based on the ECO9208222, part number 279002733 is removed from the TSP 735001356 and TSP 735003877 is released for part number 186248900. Since the BOM structure of 186248900 have changed and the optic bench part number 279002656 is not in the 186248900 BOM. We will need Milford to add part number 700009348 into the TSP 735003877 2733in order for Flex to use the TSP 735003877 to start testing the spare bench 700009348. Please help to advise when can the TSP can be updated as Flex have a shipment requirement for this spare bench 700009348. | Bench part number added to Test procedure. ECR 9218818 No Qualification required as this was only a documentation change. | Mark Jussaume |  | 357
172055 | Test Engineering | 735000438 | 2017-09-21 | TE82705 | 2 | International | Desmond Ann | CLOSED |  | 2017-11-28 | Flex feedback that the display is having issue on the weighing balance of the 1500 Series Pump Final Tester (TE82705-2). They tried to change to a new balance (same model) and encountered error message "Error -1073807339 occurred at Error Reading Balance (Serial Port read)". They suspect it is the driver issue. Will you be able to help reinstall the new driver or advise what could be the issue? Flex have assigned David to be in the factory today till 8pm Singapore time if you need any remote assistance. | Oct 25, 2017: Chau will review information from Desmond to determine assistance required to buy new or repair bad balance with lower cost in US. Nov 22, 2017: Pricing on a replacement balance. a. A&D Model HR-200…….210 g x 0.1 mg…….$ 2495.00 b. A&D Model HR250A…..250 g x 0.1 mg…….$ 1995.00 Lead time: 1 Week ARO Terms: Visa/MC/Amex or N/30 FOB – Origin UPS shipping is not included ISO 17025 Calibration CERT - $ 60.00 *includes hand delivery of new unit to Waters-Milford Copy of manual and instructions for RS232 option have been sent. Per Desmond on Dec 1, 2017: For your info, we will not be moving back the original balance from the 1500 series tester. We will be putting a new balance for the BSM station TE82943-14 and perform the IQ once it is ready. | Mark Jussaume |  | 357
172102 | Test Engineering | WAT271129, WAT210700 | 2017-10-05 | TE82736 | N/A | International | Mike Cotter | CLOSED |  | 2017-11-29 | For the following (six) Alliance instruments: (E17SM4829A, E17SM4821A, D17SM4754A, E17SM4810A, E17SM4813A, E17SM4817A), please provide the following data: 1. Final Test Record including including Empower Reports 2. Pumptray Test Records Thanks, Mike | Oct 25, 2017: Desmond will follow up with Flex to gather data as requested. Nov 27, 2017: Data provided. No qualification required because it's only a data collection. | Desmond Ann |  | 357
172122 | Test Engineering | 735002155 | 2017-10-12 | TE82943 | N/A | International | yeap siew yee | CLOSED |  | 2017-12-29 | TSP 735002155 Rev F showing software version of 1.5. However, tester is running version 1.4. See attached screenshot. | Nov 27, 2017: Actively being worked on. DEC. 2017: The release of ver.1.6 for the CSS7 project was installed across all pump testers Globally. Validation was not required due to the changes, but was performed as part of our new equipment design process. EDV and IQ qualifications are in DIR TE0082943QX. Project completed. | Edward Jackson |  | 357
172128 | Manufacturing Engineering | WAT045248 | 2017-10-16 | 999027590 | NA | International | Chan Yee Hwa | CLOSED |  | 2017-10-23 | Tooling 999027590 Motor to Housing Assy Fixture. 1) Motor to Housing Assy Fixture control panel emit white smoke and the machine malfunction. Please refer to attachment in email. | Ordered PCB's on 16 Oct. 2017. The equipment replacement parts were purchased in Singapore and the equipment is up and running per Desmond Ann (10/20/17). | Stephen Wright |  | 357
172159 | Test Engineering | NA | 2017-10-27 | TE82865 | 2 | International | Yeap Siew Yee | CLOSED |  | 2017-10-30 | Problem Statement: Few testers PC are experiencing a prompt message for symantec anti-virus protection. Tried to update but failed. Refer to attached for error message screenshot and excel list that includes those affected PCs. | updated Symantec successfully on 3 of the 6 PCs affected by this OESR - TE82865-2, TE82865-5, and TE82865-6. (BB) Bill Irons: I fixed the SAV Endpoint issue on the following PC's: CEIISNL1 10.249.55.2 Instrument Serial Number label CEIPNL1 10.249.55.3 Part Number label CEICOFP1 10.249.55.4 Certificate of Functional Performance CEIGEN1 10.249.55.5 CEIGEN2 10.249.55.6 No qualification required for updating Symantec. | Bill Bates | Bill Irons | 357
172163 | Manufacturing Engineering | N/A | 2017-11-01 | 999025111 | N/A | International | Desmond Ann | CLOSED |  | 2018-01-02 | Flex feedback that the tooling 999025111 Socket Driver, Degasser. 1) Missing Drawing 410005715 Pin, Spring,0.063'' X0.75'' Long 2) Currently flex using Torque Wrench with 2 in/lb and not 999040494 Nut Driver, 1/4 xcelite ,Modified. 3) Drawing missing extender 999040496 Nut Extender, Flangeless, Modified. | Replied to Desmond Nov-1-2017 Hello Again, Looking further at the drawing of 999040496 Nut Extender, Flangeless, Modified, it calls out the Tool Assemblies that the wrench is used on in the notes. 999029341 assembly is to be used with the torque wrench. The pin 410005715 will not be released at this point. Best Regards, Stephen Wright Tool/Production Engineer [T] +1 508 482-3330 [F] +1 508 482 3373 [W] www.waters.com [E] Stephen_Wright@waters.com Waters Corporation 34 Maple Street Milford MA 01757 USA | Stephen Wright |  | 357
172180 | Test Engineering | WAT271129 | 2017-11-14 | TE82736 | N/A | International | Mike Cotter | CLOSED |  | 2018-04-26 | Please provide the Chemistry Area Precision test results for (50) units shipped recently. Please provide the Area %RSD result for the retained peak (Butyrophenone) for FPAC and the FPBD sample sets. Chromatograms are not required. A spreadsheet or text file would be fine with data formatted as below (or equivalent): Serial Number AC %RSD BD %RSD A17SM4123A 0.25 0.125 B17SM7321A 0.22 0.28 C17SM4666A 0.125 0.35 Ideally, we'll have this data quickly (by November 17th), but let us know when we might expect to have it. If a subset of the data is available (e.g. 35 units worth), that could be helpful. Let us know when we might expect the data, and how much we can expect. Thanks, Mike | Apr 26, 2018: Data was provided. No qualification required for Data Collection for studies. | Desmond Ann |  | 357
172202 | Test Engineering | 735000445 | 2017-11-21 | TE82723 | N/A | International | Desmond Ann | CLOSED |  | 2017-11-28 | GES feedback that the tester hang after downloading firmware/ See attached screenshot. Performed the following step and restart the tester PC, Step 1:Services. Step 2: Waters DHCP Server and stop the service.Step 3: Delete DHCP.bat. Step 4: Waters DHCP Server and Start the service. Step 5: start BCH test program. Step 6: System hang. Verified with other UUT and restart the tester PC, tester still hang at Firmware download step. Test station IP address: 10.249.40.26 we have connect a UUT to the test station for Milford to access. | Erased firmware from UUT and tried again. Test functioned properly. GES indicates problem resolved. No qual required as nothing was changed. | Eric Turner |  | 357
172204 | Test Engineering | 735003612 | 2017-11-22 | TE82943 | 24 | International | Desmond Ann | CLOSED |  | 2017-12-15 | GES feedback that test station TE82943-24 is showing error "Test Gradient Start Input SHORT" after Firmware download. After checking, it was observed that the Tester box cannot turn to remote control mode when test software launched. In normal working condition, the Tester box will switch to remote control mode when the test software is launched. Checked on the tester PC's device manager and didn't find any abnormality on it. Shut down and restarted the PC, the symptom remains. Please help to advise. | Replaced USB-1408 module with new one. Tester repaired. One for one replacement. No Validation required. | Edward Jackson |  | 357
172209 | Manufacturing Engineering | n/a | 2017-11-30 | n/a | n/a | International | CHEAH Su Fern | CLOSED |  | 2017-11-30 | Would like to request to have "Manufacturing Test View" tab for Sigmasure 6 and 7 for below accounts: Email: 1. sufern_cheah@waters.com 2. Yeehwa_Chan@waters.com | added tabs for both accounts - "supply chain" for Sigmasure7 and "Mfg test" for Sigmasure6. No qualification needed for Sigmasure priviledges. Also added "supply chain" tab for Edwin Koh, as expressed in a later email. | Bill Bates |  | 357
172214 | Test Engineering | WAT271129 | 2017-12-01 | TE82736 | - | International | Mike Cotter | CLOSED |  | 2017-12-06 | This is a request to determine the total number of Alliance final test systems available for use at Flextronics, Singapore. Looking remotely, Test Stations 1-5 appear to be present, up and running. Each of these could be capable of supporting (4) Alliance test systems. Are they each capable of supporting (4) Alliance tests systems now? That would be 4 stations x 5 systems per station = 20 test systems. Originally, there may have been a 6th test station. Is this, or any other test capacity still available? Basically, how many test systems do we have (how many Alliances can we test simultaneously)? Thanks, Mike | Request for info only. No qualification required for this request. Apr 25, 2018: Mark, Mike, HuaZheng and Desmond are investigating to use QA system at Flex as the 6th system. May 31, 2018: Shipment is on its way. | Desmond Ann |  | 357
172232 | Manufacturing Engineering | N/A | 2017-12-11 | N/A | N/A | International | CHAN YEE HWA | CLOSED |  | 2017-12-18 | Tooling 999119592 Housing Assy Fixture arm unable to return to original position. Refer to email for video and photo for each components. See following link for video: https://files.waters.com/?ShareToken=A550FF5DC67F44736461A5927D0FEB8B40AF0466 | 12-Dec-2017: WPPL found a local representative from Fabco-Air, and purchased the same piston to replace it. No qualification is required due to there is no change in function and performance. | Stephen Wright |  | 357
182276 | Test Engineering | 735001356 | 2018-01-03 | TE82903 | x | International | Joel | CLOSED |  | 2018-01-16 | There is error in Traceability Matrix for Acquity TUV. | Corrected with ECR # 9220170 No Qualification required as only a document change to correct typo in TSP. | Mark Jussaume |  | 357
182296 | Test Engineering | 735002056 | 2018-01-12 | TE82984 | 1 | International | Cheah Su fern | CLOSED |  | 2018-01-16 | Symantec Antivirus definition files for each type of protection is >30 days. Definition files cannot be updated need Test Engineer to upgrade anti-virus software. | 15-Jan-2018 BB - Cannot connect to this PC via VNC, notified above engineer on 15-Jan-2018 by email to turn on PC and/or connect to network. 16-Jan-2018 BB - updated definitions manually from the Symantec website. Now up to date, no issues. OK to close, no qualification needed for virus definitions. | Bill Bates | Michael Cotter | 357
182297 | Test Engineering | 735002056 | 2018-01-12 | TE82987 | 1 | International | Cheah Su Fern | CLOSED |  | 2018-01-16 | Symantec Antivirus definition files for each type of protection is >30 days. Definition files cannot be updated need Test Engineer to upgrade anti-virus software. | Could not do a Live Update even logged in as TE-1. Had to manually install definitions downloaded straight from Symantec website (definitions are current as of 15-Jan-2018). No qualification needed for updating virus definitions. | Bill Bates | Michael Cotter | 357
182298 | Test Engineering | 735002229 | 2018-01-12 | TE83026 | 2 | International | Cheah Su Fern | CLOSED |  | 2018-01-15 | Symantec Antivirus definition files for each type of protection is >30 days. Definition files cannot be updated need Test Engineer to upgrade anti-virus software. | Could not do a Live Update even logged in as TE-1. Had to manually install definitions downloaded straight from Symantec website (definitions are current as of 15-Jan-2018). No qualification needed for updating virus definitions. | Bill Bates | Michael Cotter | 357
182304 | Test Engineering | 735002064 | 2018-01-15 | TE83013 | 2 | International | Edwin Koh | CLOSED |  | 2018-01-19 | Each UUT test pass cannot print test record follow test station hang no response.Shut down the test station and re-test problem re-appear. No application running, test station task manager show CPU usage at 100%. | Uninstalled a bunch of unnecessary printers and associated software, all is OK now. From Venture: The TE83013-2 is in good working condition. Production have completed testing on 3 units. No qualification required, no change to the test hardware, software, or process, simply removed unnecessary printers from Windows. | Matt Brule |  | 357
182331 | Test Engineering | N/A | 2018-01-19 | TE82988 | 12 | International | Chau Nguyen | CLOSED |  | 2018-05-18 | Build a new GTS and ship to GES Venture per CEA WPPL 16030. Expected delivery in Q2, 2018 | March 26, 2018: In process GTS is complete and shipped May 2018 | Paul Deschenes |  | 357
182335 | Test Engineering | WAT043815 | 2018-01-22 | TE82503 | - | International | Mike Cotter | CLOSED |  | 2018-04-25 | Please verify that Flextronics is able to test the following subassemblies on their Injector/Sealpack test stations: New, Redesigned Alliance injectors New, Redesigned Alliance sealpacks | Apr 25, 2018: Closed and no qualification was required since there is no changes to the system. | Desmond Ann |  | 357
182336 | Test Engineering | 735002984 | 2018-01-23 | TE83064 | 11 | International | Edwin Koh | CLOSED |  | 2018-01-23 | At chemical test a error message appear. It show "No systems accessibe". @23rd Jan 2018 : Issue fixed Hi Edwin, The tester is all set now. Very strange problem, I have never seen this before. All Chromatographic Systems were gone. There are supposed to be three, but Venture only requires one. I don't know why or how this could happen on its own. Anyway I created and configured a new system and the tester is now running. regards Rick | Very strange problem, I have never seen this before. All Chromatographic Systems were gone. There are supposed to be three, but Venture only requires one. I don't know why or how this could happen on its own. Anyway I created and configured a new system and the tester is now running. No qualification required due to there are no changes in software, firmware, hardware and testing procedure. | Richard Davis |  | 357
182341 | Test Engineering | 735002064 | 2018-01-23 | TE83013 | 2 | International | Chan Yee Hwa | CLOSED |  | 2018-01-23 | Each UUT test pass cannot print test record follow test station hang no response.Shut down the test station and re-test problem re-appear. No application running, test station task manager show CPU usage at 100%. --Happened on 15th Jan; Resolved on 17th Jan-- | This OESR is a duplicate of OESR 182304 closed 19 Jan 2018. Unnecessary printer drivers were removed from the PC. No verification required as no changes were made to test hardware, software, or test process. | Matt Brule |  | 357
182342 | Test Engineering | NA | 2018-01-23 | N/A | N/A | International | Chan Yee Hwa | CLOSED |  | 2018-01-23 | System show it is not able to connect to the US Server from Venture, suspect the Server is down. Result in unable to print Serial Number and Instrument Certificates. --Happened on 22 Jan, Resolved on 23 Jan-- | OESR-182342 is raised to capture the issue unable to connect to US server from Venture Singapore. Drive S was disconnected on both Venture stations. Had to reboot the PC here in the US that is currently running the labeling systems for the drives to connect. Drives needed to be clicked on to re-engage the links. Issue has been resolved. | Bill Bates |  | 357
182351 | Test Engineering | N/A | 2018-01-30 | TE82882 | N/A | International | Desmond Ann | CLOSED |  | 2018-09-10 | There is an issue on the test report for the tester TE82882. Shunt/ Cell Energy 3 ms 80000 – 250000 counts Result : – FAIL Overall Result – PASS The UUT had been retest and pass, This is serious issue and is highly possible that we may base on “Overall Result PASS” and release Instrument. This can result in Flex has to recall units from ADC if we did not detect / retest the Instrument (as we shipped fail instrument). There will be issue answering to the auditor too on the test report. Please help to rectify the issue on the test report. | New test software released and validated, ECR 9224731 | Eric Turner |  | 357
182352 | Test Engineering | N/A | 2018-01-30 | TE83064 | 10 | International | Desmond Ann | CLOSED |  | 2018-02-07 | FTN tester TE83064-10,11 and 13 are having "Runtime Error" during Carryover test at step of 4000ppm@300nm. Based on the popped up dialog box, application name called "PROCSRVR.EXE" requested the Runtime to terminate it. Please refers to attached file. When such dialog box is appeared, we have to click "OK" button to continue the test run, else Carryover test failed. Runtime error is also happened under diagnostic mode, after Carryover test failed. Also need to click "OK" button under diagnostic mode. Test system is shut down and restart, but still encountered this runtime error. Please help to advise. | Bug in Carryover Sample Set method. Step Clear Calibration requires the same Method Set as the Calibration Standard. | Richard Davis |  | 357
182353 | Test Engineering | WAT271129 | 2018-01-30 | TE82736 | - | International | Mike Cotter | CLOSED |  | 2018-03-26 | I am wondering how often (if ever) the following tests fail: 1. Z-Ram Height test. This test was critical at one point because of the sample compartment design. That design has been updated for better performance and I am wondering if this test ever fails. This test is difficult to automate (requires operator intervention), so it would be a nice test to eliminate, but only if it makes sense. 2. Pressure Ripple Test This test is performed at the pumptray level and repeated at the top-level instrument level. Are all pumptrays still tested 100%, or only spares? If all pumptrays are tested, and if the test doesn't fail, then it doesn't add much value at final test. This test doesn't take any extra time (performed during AC chemistry), so we don't gain much time by eliminating it. I am curious though. | No change to tester or test program. It's just a data collection for analysis. no qualification required. No encounter of the fail daily but sampling would be good if to eliminate this testing. Pump Tray and Spare Pump Tray are tested 100% at Pump Tray Tester. | Desmond Ann |  | 357
182358 | Test Engineering | WAT271129 | 2018-01-31 | TE82736 | - | International | Mike Cotter | CLOSED |  | 2018-03-26 | Much like my recent OESR (182352), I am seeking data on Alliance test performance. A number of years ago, the CCD tester was revised/upgraded for enhanced performance/reliability. Based on the earlier/old test design, we needed to run more tests at the final instrument level. What I'm wondering about is the performance of the 'RV' and 'Test CCD' test steps. Do these tests fail often/ever? With the new CCD tester, perhaps we don't need to be as stringent at the final instrument level. But, based on the data, and on Flextronics' experience, perhaps we should maintain this level of testing, moving forward. I'm basically looking for process capability/test capability/general performance of these two tests. | No change to tester or test program. It's just a data collection for analysis. no qualification required. After implementing the new CCD Alignment Station, RV and Test CCD, no encounter of fail on this step recently. Previously, we do see some RV test failures during the Alliance chassis issue. Test CCD is 1 unit per day and we measure the Angle Deviation using Smart Scope. | Desmond Ann |  | 357
182377 | Test Engineering | NA | 2018-02-08 | TE82988 | 3 | International | Chan Yee Hwa | CLOSED |  | 2018-03-22 | Test Box 279002953 were recently setup at TE82988-3, to test QDA DCC Board (PN: 210001115; Person In Charge: Sadiq Mohsin). Upon completion of the FCT, a label will be printed automatically. However, the label only can be print if window is log in as Administrator; Label cannot be print if window is log in as tester (operator mode), with "Access Denied" error for printing. Suspect tester login does not has "Write Access", and causing error during label printing. Need to enable the printing of label upon completion of FCT using tester login. | None taken, Sadiq Mohsin has significantly reconfigured GTS TE82988-3 without notifying anyone from Milford Test Engineering. | Richard Davis |  | 357
182380 | Test Engineering | WAT043815 | 2018-02-08 | TE82503 | - | International | Mike Cotter | CLOSED |  | 2018-04-25 | The injector injector/sealpack tester performs a test of the injector assembly, and reports pass/fail, but does not save data or generate a test report. This is a request to gather the following data from (10) production injector assemblies: | Apr 25, 2018: Closed and no qualification was required since there is a data collection for studies. | Desmond Ann |  | 357
182381 | Test Engineering | WAT043815 | 2018-02-08 | TE82503 | - | International | Mike Cotter | CLOSED |  | 2018-03-26 | TE82503 is used to test the Injector and Sealpack assemblies, but does not store data or generate test records. For (10) production units, please provide the following data: Down Count (Stream) Seal Width Based on current projects/workload, it would be ideal to have this data quickly (by end of week – February 9th). | No change to tester or test program. It's just a data collection for analysis. no qualification required. | Desmond Ann |  | 357
182383 | Test Engineering | N/A | 2018-02-09 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-05-21 | GES feedback that the serial number does not match on the serial number label for 2 cases. See attached. | I have reviewed the e-mail (attached Pg.03), and have concluded that Venture experienced a temporary loss of connection to the server directories (while hosted in Milford). This is based on the "ERROR 1118:Unable to open file manufacturer.bmp" dialog. Closing/re-opening the program would attempt to re-connect to the server, as was observed. The mis-printed labels appear to be a one-off result of this connection loss. The Venture Server is back on-line and no longer hosted in Milford. There have been no more reports of this error in the past three (3) months. No action taken. ~Bill Irons | Bill Bates | Bill Irons | 357
182395 | Test Engineering | 735001613 | 2018-02-13 | TE82882 | 14 | International | Desmond Ann | CLOSED |  | 2018-03-06 | TE82882-14 is showing error message and IP address conflict. Please see attached and advise. | There is no fix for this issue, the system in question is not the one with the IP address conflict, it is just reporting that there are two systems on the network that are conflicting. | Eric Turner |  | 357
182399 | Test Engineering | NA | 2018-02-14 | N/A | N/A | International | CHAN YEE HWA | CLOSED |  | 2018-05-29 | Certificate Printing Error- WAT207000 515HPLC Pump (Serial Number B18515131A) Wrong description selected while printing certificate. Should be 515 HPLC instead of 515 HPLC M.K7. | Email sent to Desmond to destroy that label and reprint a new one, we will deactivate this number once Bill is back from STD. Entry deactivated by Bill Irons once back. This should complete the mis-transaction circle. No qualification required, this is a database correction only. | Bill Bates |  | 357
182401 | Test Engineering | N/A | 2018-02-14 | TE82882 | 14 | International | CHAN YEE HWA | CLOSED |  | 2018-03-06 | TE82882-14 is showing Error causing Loss Comm. Pls. refer to attached photo as netword ID address conflict another error. | System appears to be running fine. Network error is external | Eric Turner |  | 357
182408 | Test Engineering | n/a | 2018-02-20 | N/A | n/a | International | Bill Bates | CLOSED |  | 2018-02-23 | Date: 02/01/2018 01:27 PM Subject: Backup Exec Alert: Media Remove (Server: "VENTUREMFG2") (Job: "Venturemfg2.corp.waters.com Backup 00002-Full") (Server: "VENTUREMFG2") (Job: "Venturemfg2.corp.waters.com Backup 00002-Full") Please remove the media from the drive. SolarWinds email notification was sent with above message. Please swap tapes in the backup tape drive of the main server at Venture at your earliest convenience. Also, please advise on the number of tapes available so we can send a few over if needed. Thanks ~Bill | Tape changed - error was due to Chinese New Year and having no one in the office. No qualification required for this. | SuFern Cheah | Desmond Ann | 357
182412 | Manufacturing Engineering | N/A | 2018-02-22 | N/A | N/A | International | Cheah Su Fern | CLOSED |  | 2018-02-22 | The US Server is down both Certificate and Unit Serial Number label printing station cannot print out.. See attach print screen capture. | Drive disconnected on both PCs for unknown reason, cannot re-connect manually. Rebooting the server here in US (Engserver). Now able to connect both PCs to server without issues, also able to open both CoFP and serial number std. Ok to close, no qualification needed for a server reboot. | Bill Bates |  | 357
182413 | Test Engineering | NA | 2018-02-23 | TE82526 | 1 | International | CHAN YEE HWA | CLOSED |  | 2018-03-06 | TE82526-1 Pump Tray Tester is keep on restarting on its own. Found bloated capacitor on Computer Motherboard. New PC has been replaced. Configuration needed for this new PC. Details: Login ID: TE-1 Login PW: testeng IP Add: 10.249.26.26 | Added software folder to C: drive --> TE82526. Created a copy of .exe of program and dragged to desktop. Configured comm ports to use Comm 1 & 2, 1 for the tester and 2 for the scale. Tester now working and several trays tested with success. No further qualification needed. | Bill Bates | Mark Jussaume | 357
182425 | Test Engineering | N/A | 2018-03-05 | N/A | N/A | International | CHAN YEE HWA | CLOSED |  | 2018-03-12 | Test procedure WAT047159 requires update to reference this new 279003947 as per stated in ECO-9220840. This change needed urgently as we have missed the spare shipment. Thanks. | TSP has been revised to include new part number. see ECR 9221367 No change to tester or test program. Part number add to TSP only. no qualification required. | Mark Jussaume |  | 357
182432 | Test Engineering | WAT047159 | 2018-03-05 | N/A | N/A | International | Chris Vandevoordt | CLOSED |  | 2018-03-06 | Per Flex feedback, Test procedure WAT047159 need to be updated as PN 279003947 or 700011708 is not cover by TSP WAT047159 Rev 8. | Duplicate of OESR 182425 | Mark Jussaume |  | 357
182436 | Test Engineering | N/A | 2018-03-08 | TE82865 | 5 | International | Desmond Ann | CLOSED |  | 2018-03-22 | CEI reported that one of the tester TE82865-5 unable to download the firmware and it shows on the note as Invalid Ethernet address. CEI have tried 2 units on this tester it still show the same error. | Purged Empower files that retain serial numbers and Ethernet addresses. | Richard Davis |  | 357
182440 | Test Engineering | N/A | 2018-03-11 | TE82760 | 5 | International | Desmond Ann | CLOSED |  | 2018-03-13 | GES reported that they are seeing error on the tester TE82760-5. See attached file for your reference. | Like for like replacement of USB to serial port device requires no qualification | Eric Turner |  | 357
182457 | Test Engineering | WAT046538 | 2018-03-21 | N/A | N/A | International | Matt Brule | CLOSED |  | 2018-10-19 | I am working on revision 12 of the Electrical Safety Test procedure WAT046538. Upon release it will require daily verification of the electrical safety testers using the S14-03 load boxes already provided, this will require a tag to be attached to each tester to track this activity. I am going to send two boxes of tags to Desmond to be delivered to each CM (Flextronics, and Venture). | Apr 25, 2018: ECO not released and waiting for Matt to release. Implemented 19 September 2018, ECR 9221675. No qualification necessary this implements a daily verification of the Hi-Pot testers using a calibrated test box. | Desmond Ann |  | 357
182459 | Test Engineering | 735003612 | 2018-03-23 | TE82943 | 1 | International | Desmond Ann | CLOSED |  | 2018-03-23 | GES feedback that abnormal loud noise coming out from the TE82943-1 Test Box. Suspect Wash Pump is faulty. Note: see attached photo. Please help to advise if this Wash Pump is used in any of our instruments or the Waters part number for the Wash Pump. | Provided information via email. Information only, no changes made, no qualification required. | Michael Cotter |  | 357
182468 | Test Engineering | WAT271129 | 2018-03-28 | TE82736 | N/A | International | Mike Cotter | CLOSED |  | 2018-04-26 | Please provide the DHR for the following instrument serial number: K17SM4227A I will be out of the office for a bit, so please provide the data to Michael Ross and Michael Cook. Thanks, Mike | Apr 25, 2018: Closed and no qualification was required since there is no changes to the system. | Desmond Ann |  | 357
182472 | Test Engineering | 279003844 | 2018-03-29 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-04-23 | Flex feedback that the PDA Chip Alignment Station (Starlite) is down and below is our observation : 1. New PDA Chip Alignment Station used in 279003844 was down due to X-axis not moving when turning the knob clockwise and counter-clockwise. 2. Noticed that the red stopper on the other end of X-axis knob was loose and need to retighten back. 3. Need to request for the drawing on how to disassemble and assemble back. 4. Pls .refer to attached photo for the observation and advise. | Any service done on the starlite needs to be done by the QVI (http://www.smartscope.com.sg/) local service organization and then they should re-calibrate the starlite. The custom plate comes off of the starlite by the 4 screws on top. Once you reinstall the plate, you will not have to recal as the distance from the reticle to the tooling pin did not change. | Stephen Wright | Eric Turner | 357
182479 | Test Engineering | N/A | 2018-04-04 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-04-04 | We will need to create an OESR account for our new Test Engineer Hua Zhen. Email Address: huazhen_luo@waters .com. Ref. Sap ID = SGHZLUO | Name User ID Privilege Set Added to Assignee First Last Drop down listing Luo Hua Zhen SGHZLUO ASSIGNEE 04-Apr-2018 | Bill Irons |  | 357
182482 | Test Engineering | N/A | 2018-04-04 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-04-04 | GES feedback that the Sigmasure 6 server is down. Please help to check and advise. | Rebooted and is up and running, no qual required NOTE: IP Numbers 10 249 40 6 VENTUREGEN1 VENTURE GENEALOGY 1 10 249 40 7 VENTUREGEN2 VENTURE GENEALOGY 2 | Bill Bates | Eric Turner | 357
182491 | Test Engineering | N/A | 2018-04-09 | N/A | N/A | International | Luo HuaZhen | CLOSED |  | 2018-04-10 | Need a new account for SigmaSure 6 & 7. Please assist to create one by using email: huazhen_luo@waters.com. Thanks | Created accounts for Hua Zhen, emailed him with the login credentials. Should be all set. | Bill Bates | Eric Turner | 357
182492 | Test Engineering | N/A | 2018-04-09 | TE82988 | 3 | International | Desmond Ann | CLOSED |  | 2018-04-13 | Tester or system not able to power on. This is due to the " AC Power distribution box " not working or not normal. It has been noticed that when the AC switch in the front panel of the tester is switch on there was an arching like sound generated in the AC Power distribution box. Attached are the photos of the AC power distribution box and the AC switch for reference. We will go over to Venture to understand the issue tomorrow and see if we can replace the equipment with our spares. Let us know if you have any comments. | Wppl managed to recover the GTS tester with the spared AC Power distribution box on hand. It's a like for like kind exchange, so no qualification is required. | Desmond Ann | Richard Davis | 357
182493 | Test Engineering | 735002064 | 2018-04-09 | TE83013 | 1 | International | Desmond Ann | CLOSED |  | 2018-04-12 | PC of Tester TE83013-1 which is a tester of 186015042 cannot start window successfully. Performed hard reset and tried for a few times, still the same problem and cannot start window normally. We also used function key F8 and tried different booting modes, but still cannot start window. Please refers to attached picture for reference. | Hard disk moved to another identical PC, like for like replacement requiring no verification. | Matt Brule | Desmond Ann | 357
182501 | Test Engineering | N/A | 2018-04-12 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-04-12 | From: "Joel Mamalateo Deduyo" <joelmamalateo.deduyo@flex.com> Date: 2 April 2018 at 7:04:18 PM SGT To: "'Desmond Ann'" <Desmond_Ann@waters.com>, "yeehwa_chan@waters.com" <yeehwa_chan@waters.com> Cc: "Edwin Koh" <Edwin_Koh@waters.com>, "Nallajanula Phani Bhushan " <phanibhushan.nallajanula@flex.com> Subject: OESR Active Desktop Recovery Hi Desmond, Pls. refer to attached OESR to Restore. my Active Desktop. Autoloader-2 Active Desktop Recovery - display not able to return to default Waters display. Autoloader-2 IP Address : 10.249.26.34 1. The station restart due to Power Trip Off last 29 Mar 2018. 2. Upon restart, it does not go to default Active Desktop Recovery. 3. Pressing Restore my Active Desktop show as missing desktop.htt file. 4. Flex does not have admin right to “Restore my Active Desktop” | Active desktop should not be enabled. Logged in through VNC and disabled, changed the background to the default Waters/TE background. No explanation for how it got turned on. OK to close, nothing to qualify here. | Bill Bates |  | 357
182513 | Test Engineering | 279003844 | 2018-04-20 | TE95515 | N/A | International | Desmond Ann | CLOSED |  | 2018-04-10 | PDA Chip Alignment Station (Starlite) : 1. New PDA Chip Alignment Station used in 279003844 was down-v "Switch to Split Screen Mode" is not available on the display. 2. Enable "Remote Assistant, Remote Desktop and Remote Desktop - Remote X" @ Windows Firewall \ Allowed Program. Computer Description : 042398E IP Address : 10.249.26.59 DVD\CD Drive : Measure X CD System Properties ; Allow connection from computer running any version of Remote Desktop (less secure) 3. The QVI dongle is plug-in the usb port. I plug-out and plug-in again and restart the PC. Pls. take a look again tonight and advise what can we do here from our side as PDA and eLamda shipment is affected. | Repair performed by Starlite in Singapore | Eric Turner |  | 357
182514 | Test Engineering | 271129 | 2018-04-20 | N/A | N/A | International | Luo HuaZhen | CLOSED |  | 2018-06-29 | In order to support production capacity due to yield loss with carryover, we are setting up the 6th Alliance tester. Could you help support us on this effort to try get a tester up by end April? We already have the CAPEX approved to purchase the 6th tester. | Tester all set up and running. IQ performed and ready for use. | Mark Jussaume | Michael Cotter | 357
182518 | Manufacturing Engineering | N/A | 2018-04-23 | 999119657 | N/A | International | Desmond Ann | CLOSED |  | 2018-04-25 | Venture feedback that the shuttle press machine 999119657 cannot be powered on. Open casing found Y3 alarm on. Perform check on all wire connection no loose connection found. Use multi-meter to measure the 24 VDC Output connector found no output. Suspect power supply is faulty or other issue cause Y3 alarm on. Power supply FA-24PS is available in US but not available in Singapore. Will need Milford help to purchase 3 units of FA-24PS for our replacement and spares. | Parts were shipped 23-Apr on TRF SWRT AY4LC5. I was only able to get my hand on two power supplies. SW 23-Apr-2018 Apr 25, 2018: Steve to provide Tracking # | Stephen Wright |  | 357
182535 | Test Engineering | 735002056 | 2018-04-30 | TE82984 | 1 | International | Desmond Ann | CLOSED |  | 2018-05-02 | GES feedback that when UUT completed testing at TE82984-1 Column Manager test station, an error message will appear on the screen. It show “Error 7 occurred at New File”. (see attached) After click continue icon, a Test Record will print. But we are not able to continue next “Fluidic Test” at TE82987-1 Empower test station due to instrument serial number not found. The IP address of the TE82984-1 is 10.249.40.21 and TE82987-1 is 10.249.40.22. Please help to check and advise. | That tester writes geneology data to the server, which was recently replaced. Added the necessary directory to the server and no error message. Not sure about the serial number thing, the serial number on the unit connected to TE82984 looks OK. Emailed Desmond to have a unit connected. No issue found, per Desmond tester is up and operational. No qualification required as no changes were made tester was reconnected to server so data can be written. | Matt Brule |  | 357
182541 | Test Engineering | N/A | 2018-05-03 | N/A | N/A | International | Luo HuaZhen | CLOSED |  | 2018-07-10 | Recently a GTS tester down in Venture Singapore due to no power, the problem was fixed by replacing AC power distribution unit from WPPL spare parts. We tried to contact Eaton Singapore/distributors for repairing of this faulty AC power distribution unit, but getting no response regardless hotline or email. So we need Milford help to advise possibility of repairing this faulty AC power distribution unit in Milford or acquiring another spare AC power distribution box for future use. | Emailed Luo and asked him to return the power supply to Milford for repair or replace on 3-May-2018... They sent the power distribution box to Milford for review I found the plug melted replaced the 5 lug plug with a new like for like replacement and tested the power supply after confirming the repair corrected the problem I shipped the unit back to WPPL | Paul Deschenes |  | 357
182555 | Test Engineering | 735001613 | 2018-05-15 | TE82882 | N/A | International | Joel Mamalateo Deduyo | CLOSED |  | 2018-09-13 | SPARE Optic Bench PN 279003847 ASSY, OPTICS BENCH, ACQUITY PDA is testing is not inside 735001616 Rev U and to the following Test Station : TE82969-1, TE82882-8, TE82882-9, TE82882-10, TE82882-11, TE82882-13 TE82882-14 Pls. include inside the TSP or issue TMD to test the UUT | We do not test 279 level product | Eric Turner |  | 357
182556 | Test Engineering | 735002984 | 2018-05-15 | TE83064 | all | International | Richard Davis | CLOSED |  | 2018-05-23 | Investigate Sample Syringe Decay test failing LSL (-5 psi). This started during validation of test software 1.42 and new instrument firmware. It is occurring on all part numbers. Try to get a screen capture of the pressure trace. We would also like to know what they do when it fails like this. If we can get an instrument that fails the LSL consistently leave it on the test stations so Milford engineers to look at it. | The requested data and screen capture were sent to Richard. Richard agreed to close this ticket on 23 May 2018. No qualification required for data collection activities. | Luo Hua Zhen | Desmond Ann | 357
182566 | Test Engineering | 430005420 | 2018-05-18 | N/A | N | International | Cheah Su Fern | CLOSED |  | 2018-05-18 | Would like to Request Bill Irons to add 430005420 Rev B (Assy, Tube, Needle Wash) to critical clean Auto Bagger part number database. If the part number is already in the database, could you please forward a copy of database for us due to Key Point system will crash when connected to server? | emailed a fresh ( 17-May-2018 ) copy of "labels.dbf" to Cheah Su Fern | Bill Irons |  | 357
182580 | Test Engineering | WAT271129 | 2018-05-29 | TE82736 | N/A | International | Mike Cotter | CLOSED |  | 2018-06-28 | For the Luna project, there is some interest in previous Alliance product data. Specifically, the team may be interested in: Display test data Degasser test data Heater/Cooler test data The testing is manual, so it may require mining old records for data, but by mid-week (by May 31st), please let me know how long it would take to provide two years' worth of data for the above tests. Thanks, Mike | [28Jun2018] The requested data has been sent to Mike [26Jul2018] Pending for Mike to confirm if this case can be closed. This was a request for data only, with no changes made, no qualification required. | Desmond Ann | Luo Hua Zhen | 357
182587 | Test Engineering | WAT025373 | 2018-06-04 | TE82475 | N/A | International | Cheah Su Fern | CLOSED |  | 2018-06-29 | TE82475 - Pressure Decay Tester is not able to hold off pressure for position 1. Action Taken : Verified System Transducer during calibration to 1000 PSI - OK Verified Position 1 : during calibration to 1000 PSI - OK Verified Position 2 : during calibration 10 1000 PSI - OK Perform testing as per WAT025373 - to each position and result failed. Position 1 Pneumatic Valve 1 exhibit abnormal temperature during operation compare to other valve . Suspect to have issue with Valve Number 71216SN1BL00N0L111Q3 | Tester is working fine no action needed and no need to order any parts | Paul Deschenes |  | 357
182598 | Test Engineering | WAT043815 | 2018-06-14 | TE82503 | 05 | International | Cheah Su Fern | CLOSED |  | 2018-06-14 | TE82503-05 - Injector Seal Pack Tester failed on decay test.Tester able to ramp up to 5000PSI. However, there is heavy leakage and unable to meet specs i.e 660 PSI. - Replaced valve 1 with new valve. No improvement. - Replaced valve 2 with same new valve. No improvement. Replace back original valves. Replacing tubing from filter to valve 1 & Valve 2. In-progress. | Per Edwin, leak has been identified and repaired. | Michael Cotter | Mark Jussaume | 357
182599 | Test Engineering | N/A | 2018-06-14 | TE82503 | 05 | International | Cheah Su Fern | CLOSED |  | 2018-06-14 | Request schematic diagram of TE82503-05 for trouble shooting purpose. | Copied a copy of TE82503 rev 1.vsd to the Test Engineering folder of the Flextronics server. Also sent a copy via email. No qualification required, simply a request for information. | Michael Cotter | Mark Jussaume | 357
182632 | Test Engineering | 735003612 | 2018-06-29 | TE82943 | 1 | International | Luo HuaZhen | CLOSED |  | 2018-09-01 | Venture Singapore Production is seeing a spike of 0.3ml/min flow rate accuracy failures (lower spec limit – 0.297ml/min) on QSM (186017000) since Apr 2018 . Please kindly refer to the "Brainstorming for 0 3ml flowrate test failure-29 Jun.pptx" for detail study and advise if further action is required | This was closed previously and I'm closing again. The results in the report were not found in Milford and further results showed no spike in failures after the report date. | Edward Jackson |  | 357
182634 | Test Engineering | WAT210700 | 2018-06-29 | TE82526 | 2 | International | LUO HUAZHEN | CLOSED |  | 2018-08-02 | TE82526-2 Pump Tray Tester is keep on failing on Ripple Test \ Flow Test. Action taken to verify the failure: 1. Pin Plug FV2 during Ripple Test - Test Result PASS 2. Pin Plug FV2 during Flow Test - Test Resut PASS 3. Removed the Pin Plug during Ripple Test - Test Result FAIL and observed small amount of leak solvent during test. 4. Removed the Pin Plug during Flow Test - Test Result FAIL and observed small amount of leak solvent during test. 5. Replaced ASFVO Flow Valve, swapped Degasser. 6. The problem persists. Suspect valve as shown in attached. But we don't have spare to verify. Would you please check and advise if Milford is able to source this vale requested by Flex? Thanks Update from Flex on 02 Jul 2018, Flex found vent valve P/n:7716-053 (FV3) is having leakage in between the head and the body of tester So they need this valve to fix the ripple test issue. Please help to check if you're able to source this valve in Milford. Thanks | Replaced with new valve of same type so no qualification required. | Mark Jussaume | Paul Deschenes | 357
182641 | Test Engineering | 735001613 | 2018-07-09 | TE82882 | 9 | International | Mandy Stewart | CLOSED |  | 2018-07-30 | TE82882-9 and TE82882-14 Footer Mismatch with Serial Number. Action Taken: 1. Stop using TE82882-9 and 14 2. Retest UUT affected to other station 3. Refer to attached print-sample The serial number at the top of the printed test document is different from the generated serial number at the bottom of the screen. | It's a duplicate and similar to 182647, so it's closed. | Eric Turner |  | 357
182646 | Test Engineering | 735001613 | 2018-07-10 | TE82882 | 9 | International | Mandy Stewart | CLOSED |  | 2018-07-30 | TE82882-9 and TE82882-14 Footer Mismatch with Serial Number. Action Taken: 1. Stop using TE82882-8, 9, 10 amd 14 2. Retest UUT affected to other station 3. Refer to attached print-sample | It's a duplicate and similar to 182647, so it's closed. | Eric Turner |  | 357
182647 | Test Engineering | 735001613 | 2018-07-10 | TE82882 | 9 | International | Mandy Stewart | CLOSED |  | 2018-09-10 | TE82882-9 and TE82882-14 Footer Mismatch with Serial Number. Action Taken: 1. Stop using TE82882-8, 9, 10 amd 14 2. Retest UUT affected to other station 3. Refer to attached print-sample Joel mentioned that it is the same issue as they had last December and has been working well until this week where the first two testers mentioned stopped working correctly, then the other two testers stopped working last night. | New test software released and validated. ECR 9224731 | Eric Turner |  | 357
182648 | Test Engineering | NA | 2018-07-10 | TE82503 | 5 | International | Mandy Stewart | CLOSED |  | 2018-07-11 | Request for drawings for this tester to assist with failure analysis. Failures at ramp and decay test. | Documents supplied. Request for documentation only so no qualification required. | Mark Jussaume |  | 357
182649 | Test Engineering | NA | 2018-07-10 | TE82475 | 2 | International | Mandy Stewart | CLOSED |  | 2018-07-11 | Request for drawings to help assist in the failure analysis of tester. | Documents supplied. Request for documentation only so no qualification required. | Mark Jussaume |  | 357
182664 | Test Engineering | N/A | 2018-07-16 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-08-07 | Will like to request for Flex team to have the account privileges to view Test Records on Sigmasure 7 for testers reside in Flex only in order for them to review FPY data. Below are the users that will need to have the account privileges to view Test Records on Sigmasure7. Flex: - Account to view ONLY data at Flex JoelMamalateo.Deduyo@flex.com yambeegary.lai@flex.com phanibhushan.nallajanula@flex.com tan.cheekhim@flex.com | Gary and Joel have existing accounts, added the other 2 new accounts. Looks like they will only have access to PDA, TUV, and Pumps because of the way some other test equipment was input. They aren't shareable as they would also see other CMs' data, and we do not allow that (Sample Managers for instance are under 'Waters' then 'Flex', not just Flex. No qualification required for adding users to Sigmasure. | Bill Bates |  | 357
182670 | Test Engineering | NA | 2018-07-17 | TE82988 | 3 | International | Luo HuaZhen | CLOSED |  | 2018-07-20 | Tester was failing test " Serial Communication " and have constant beep sound. Please see attached Video. Suspect is the CPU faulty. Could you please send a CPU with FW V1.46.61 for this repair? Or can we use our production CPU with loaded FW V1.69.154 (Instrument No.186015043) to replace suspect CPU in GTS 210000478 test fixture and contiue production for PCB 210000478? If we're not suppose to use latest FW, please send us the old FW V1.46.61. Thank you | Two CPU2k s loaded with the correct firmware for test have been sent to WPPL. Also the firmware file has been sent to Lou incase he a future need. No qualification required...like for like replacements. | Mark Jussaume | Tony Fontes | 357
182674 | Test Engineering | 735003657 | 2018-07-18 | N/A | N/A | International | Chan Yee Hwa | CLOSED |  | 2018-07-18 | Waters Corporation Instrument Certificate Printer WATERSCOFP1 HP Laser Jet Enterprise M604 unable to print. The PC is able to access to server. Error Message (Microsoft Access): Run-time'76': Path not found | The Error Message (Microsoft Access): Run-time'76': Path not found was caused because the Certificate Station lost priviledges to connect to the server shared directory "mds_app". Waters BT (David Moss) had reconfigured FLEXTRONICSMFGnew server for backing up to Microsoft OneDrive for Business. When he moved the directories into the OneDrive target directory, the shared "Users" preferences were lost. I added the appropriate users to the server shares and successfully tested connectivity. | Bill Irons |  | 357
182685 | Test Engineering | N/A | 2018-07-23 | N/A | N/A | International | Chan Yee Hwa | CLOSED |  | 2018-07-24 | Instrument Serial Number printer " WATERSPNL1" DATAMAX PRINTER at Flex Singapore unable to print out with below error message: Problem with Shortcut "The Item 'Singapore_iclass_A_LV6_REV11.lbl' that is shortcut refer to has been change or move, so this shortcut will no longer work properly. Steps Followed to solve it internally: 1.Try to click the shortcut icon to print shows " Problem with shortcut ". see the attachment for the actual prompting msg error. 2.Try to re-start system ----Still unable to print and shows same problem as mentioned above. Please refer email for photo. | Some server files had been moved. Files have been restored to correct location and labeling successfully tested. | Bill Irons | Bill Bates | 357
182688 | Test Engineering | N/A | 2018-07-24 | N/A | N/A | International | Chan Yee Hwa | CLOSED |  | 2018-07-24 | Certificate Printer WATERSCOFP1 HP Laser Jet Enterprise M604 was unable to print with Run-time error '68' Device unavailable. Cannot link to Server. | Some server files had been moved. Files have been restored to correct location and labeling successfully tested. | Bill Irons | Bill Bates | 357
182695 | Test Engineering | N/A | 2018-07-26 | N/A | N/A | International | Chan Yee Hwa | CLOSED |  | 2018-07-26 | Flex Singapore insert Tap Media into Server for backup purpose. Error occur (please refer email for screenshot). This has cause Label Printers and Certificate Printer Down as well. | While inserting a back-up tape into the server, there was an unexpected server restart, which caused the temporary loss of connectivity to the shared folders. The server is now back on-line and I have successfully tested connectivity for both; labelers and certificate systems. | Bill Irons |  | 357
182696 | Test Engineering | N/A | 2018-07-26 | N/A | N/A | International | Yeehwa Chan B.Irons | CLOSED |  | 2018-07-26 | Yeehwa Chan has created three (3) OESR's (182685, 182688, 182695) recently, and no email notification was sent to the OESR Coordinators. This OESR has been created by BILL IRONS to test her ability to use the "EMAIL BUTTON". | EMAIL received successfully: International Test Engineering 182696 OESR ADDED. OESR System to: Ops_Eng_Support 07/26/2018 11:31 AM Cc: bill_irons | Bill Irons |  | 357
182701 | Test Engineering | N/A | 2018-07-27 | TE82943 | All | International | Luo HuaZhen | CLOSED |  | 2018-08-08 | WPPL QA did 1x BSM RMA inspection and noticed the PCBA P/N shown in HyperTerminal is PN: 210000308, but the PCBA physical P/N in the unit is 210000422. We checked a fresh built BSM in production and the same obsoleted P/N: 210000308 shown in HyperTerminal after loading firmware Version 1.71.387, 9Feb 2018 13:16:28. The questions is, 1) Would be a concern if an obsoleted P/N shown in HyperTerminal? 2) Is this P/N loaded into system during BSM firmware download? 3) Can this discrepancy be fixed? | The part number displayed in hyper terminal (210000308) is derived from a register in U16 on the personality board. However, RDE has informed us that the PLD U16 can not be modified at this time. Engineering is aware of this issue. They have concluded that this does not contribute to any other issues and that leaving this as is was the best solution. No Qualification Required as there is no change here. | Edward Jackson |  | 357
182706 | Manufacturing Engineering | 289001458 | 2018-08-01 | 999042549 | N/A | International | Desmond Ann | CLOSED |  | 2018-08-10 | During the assembly of Housing Check Valve (405004106) and Holder Cartridge (405004105) as per ASP 289001458 REV C, the swage fixture bushing break and unable to perform the press operation. This is affecting production line down for 186269505 & 186269506 models Will need Milford help to purchase and sent us 2 pcs of 999042549. | Two parts were sent to WPPL 01Aug waiting for TRF number to follow up. SW. Looking into creating a KANBAN system for consumable parts, Desmond is creating a preliminary list to evaluate. | Stephen Wright |  | 357
182734 | Test Engineering | NA | 2018-08-20 | TE82519 | 01 | International | Mandy Stewart | CLOSED |  | 2018-09-20 | Request for schematics and any other information to help with the repair and replacement of parts for this tester. Frequent down time of tester Frequent changing of valve tubings need to be changed every few days PCBA 056451 obsolete | Enclosed is the only information that we have that may be able to help you with the tester TE82519, 2690 Pump Head Tester. Frequent down time of tester: we do not have any data Frequent changing of valve: we do not have any data Tubing needing replace every few days: we have no data PCBA 056541 Obsoleted: That PCB Pump Driver Board was replace with the new one Part Number= 210000425 on 30, Apr., 2012. We have a system here that is used for testing Spares for the field, we are not having any issues with it so far. No qualification required - request for information only. | Tony Fontes | Mark Jussaume | 357
182748 | Test Engineering | N/A | 2018-08-27 | N/A | N/A | International | Desmond Ann | CLOSED |  | 2018-08-27 | Flex feedback that the Certificate of Functional Performance for UUT Serial Number H18TUV969A was printed with WRONG Model Description 5844441 instead of ACQUITY UPLC TUV DETECTOR. Please help to amend the COFP from 5844441 to ACQUITY UPLC TUV DETECTOR. (refer to scan sample in email) attachment on Pg.02 of this OESR. | Wrong description, entered by operator, on Certificate Of Functional Performance. I have inactivated the record, so that the certificate can now be reprinted with the correct description. | Bill Irons | Bill Bates | 357
182765 | Test Engineering | NA | 2018-09-03 | TE82720 | 2 | International | Luo Huazhen | CLOSED |  | 2018-09-19 | The tester PC always keep rebooting when PC is power on so not able to use for testing. Currently there are boards W1210000814 for testing. The following are the other PCBA list qualified to test using the tester ( W1210000423-R, W1210000399-R, W1279000592/Z, W1279001098, W1081318 ) Do you have a spare PC for replacement? If not, is there a specified requirement for this PC to be used on this tester? Please kindly advise if you have any concerns or troubleshooting tips for this repair. Thank you | Rick and I converted one of the GTS PCs to use as a replacement for this. Deleted the GTS components and Rick loaded the necessary UI for the front end tester. We shipped this PC out to WPPL on 5-Sep-2018, should be there by Friday. Will still need to be configured once on site. 19-Sep-2018 BB - Once PC was on site, the engineers there set up the PC with a little help from Rick with the configuration. No further actions needed at this time, OK to close. | Bill Bates | Richard Davis | 357
182766 | Test Engineering | 700010507 | 2018-09-03 | TE82903 | NA | International | Mandy Stewart | CLOSED |  | 2018-09-05 | SPARE PN : 700010507 SN 35J7V025 due to TE82903 Tester GUI 700010507 not able to accept the SPARE serial number 35J7V025. TE82903 Test Software 700010507 is showing INVALID SERIAL NUMBER (refer to screen shot 700010507) If using 700010507 with SN format for SN H18VTU025A, the GUI 700010507 is able to proceed on testing. (refer to screen shot SPARE 700010507). If not able to resolve, this will impact 1st shipment of 700010507 (1 unit) by 10 Sep 2018 | Per Mandy, It is a Flex old serial number format and this is why it isn't working. They (Flex) will change the format for testing. No qualification required since no actions required from Milford. | Bill Irons | Mark Jussaume | 357
182774 | Test Engineering | N/A | 2018-09-07 | N/A | N/A | International | Jen Chen | CLOSED |  | 2018-09-10 | Request access to SigmaSure server, for test records for instruments tested at CMOs/ Suppliers. Users requiring access: 1) Eugene Tan (SGETAN) 2) Loh Wan Jing (SGWJOH) 3) Shaynee Lim (SGSLIM) 4) Tan Kah Chin (SGKATAN) 5) Jen Chen (SGJCHE). Thanks a lot! | Added all 5 to Sigmasure 7, with the following priviledges: dashboard/Supplier Chain/Supplier Gen, and CEI/Flex/LC Pumps F/LC Pumps V/LC Sample Managers/Sigma6 Legacy/Venture. No qualification required for adding users to Sigmasure. | Bill Bates |  | 357
182805 | Test Engineering | NA | 2018-09-21 | TE82969 | 1 | International | Luo HuaZhen | CLOSED |  | 2018-09-21 | E82969-1 Not able to do the linearity Test for 186015032 IP Address : 10.249.26.75 Test Result as below: % RSD Measurement 1.#QNAN Outcome FAIL Mean Measurement 1.#INF Outcome PASS Std Dev Measurement 1.#QNAN Outcome PASS We will leave one unit for Eric for Eric configuration. | Older test version of linearity test software was installed. Updated to released version from finserv2. Ran model that was failing and it reported the correct result set. | Eric Turner |  | 357
182806 | Test Engineering | NA | 2018-09-21 | TE82882 | 13 | International | Luo HuaZhen | CLOSED |  | 2018-09-21 | TE82882-13 Test station test software version 4.1 completed test but not able to print report. IP ADDRESS : 10.249.26.84 Verify network connectivity and able to ping server. Verify test print using the computer and is able to print the report. Suspect is due to access by tester account is block. Flex to leave UUT for Eric to access tonight. Raise OESR for support | Found permissions error on system. Changed permissions on folder to allow tester to write to print output folder. System now prints so no qualification required. | Eric Turner |  | 357
182851 | Test Engineering | 735002155 | 2018-10-12 | TE82943 | 17 | International | Chan Yee Hwa | CLOSED |  | 2018-10-17 | BSM TSP: 735002010 Rev L, Software version: 2.2 QSM TSP: 735002155 Rev H, Software version: 1.7 User not able to sign in to test program at QSM & BSM station. Error message indicates valid ID must be included in Operators.txt file located on PC's Desktop. Verify Operators .txt file, user ID is in the file. 1. Mohd Redzuan; ID: 1074371. Unable to access Tester 17 2. Syed Shamsudin; ID 847838. Unable to access Tester 11 to Tester 18 Please refer Email for print screen. | 17-Oct-2018: I connected to TE82943-17 remotely. Opened Operator.txt file. Found Mohd Redzuan and Syed Shamsudin entries. Looking at the entry lines I discovered they had entered the wrong format. Entry should be: User_Name(TAB) User_SAP_Login_Name what I found was: User_Name(TAB) User_SAP_Login_Name(TAB) I removed the (TAB) after the User SAP Login Name or in this case Employee Number for CMs, and verified proper operation using QSM program currently being used on -17 tester. Chan Yee Hwa, please return back to testers 11 through 18 (excluding -17) and correct the Operators.txt file by removing the (TAB) that is "after" the Employee Number for employee Syed Shamsudin 847838. No Qualification is required. This was only an operator data entry issue. | Edward Jackson |  | 357
182865 | Test Engineering | 735002510 | 2018-10-19 | N/A | N/A | International | Luo HuaZhen | CLOSED |  | 2018-10-23 | CEI Singapore reported they have trouble to access all the TSP for CPSO and ASO model in Waters Portal. It show that 'No authority for this Material'. The following TSPs are affected. And user ID is V2917ATA. TSP WAT046538 GENERIC HI-POT & GND CONT Test PROC. 12 TSP 735001053 TEST PROCEDURE, ELEVATOR M TSP 735002510 CPSO FINAL TEST D TSP 735001199 Test Procedure ETS/ATS Shuttles B Please kindly help to check & fix this reported issue. Thank you | Correct user ID is V29177ATA, but still isn't able to access these above files. Recommended getting in touch with BT to get the correct access, ticket #B5S4RG was opened. OK to close this OESR, no qualification required because nothing was done. | Bill Bates |  | 357