## Sheet1
Log Number | Engineering Group | ASP TSP No | Date | TE Tool No | TE Tool No Instance | Request | Originator | STATUS | MPP Date | Date Completed | Problem Summary | Summary | Assignee 1 | Assignee 2 | TE Tool No Count
150165 | Test Engineering | N/A | 2015-06-08 | TE83026 | -1 | International | Desmond Ann | OFI |  |  | GES have reported that in the month of May 2015, SigmaSure data shows that “Failure Mode Not Indicate” has increase to 75% follow by “Ecord 3” with 13% for the Column Manager (186015043). Base on GES Manual Data Collection, Heater Cooler 2 test fail is the top contributor for “Failure Mode Not Indicate” (see attached pie chart). The failure symptom is the unit will immediately fail after Heater Cooler 1 is completed and this is not related to the unit as by using both Diagnostic and Acquity Console Program, we are able to verify that both Heater Cooler 1 & 2 is in working condition as there is no problem in Heating Up and Cooling Down the temperature. Most alarming is that when the unit fail Heater Cooler 2 test the test record show that it is a “PASS” and with missing “Left Valve” and “Right Valve” test step. Attach test record for your reference. Both GES test operators and QA Inspector have been trained and instructed not to release any unit with incomplete test record for shipment. Attach document is a completed test record. | Found a bug in program. Failures look the same as for returned PCBs. Actively working on change. Chia/Desmond will work with Matt on this issue on Thursday August 27, 2015. Matt on Sept 29, 2015: I believe this is the same issue as 150156, need to confirm and update code. This is a priority now that ISO audit is complete. Jan 5, 2016: Ongoing investigation. Working on update with new CMs code (150156, 150165), are both an issue with CMA test that needs to be resolved in Asia, this will came with the update to Sigmasure 7, the two OESRs could be combined. M. Brule May 12, 2016 OESR 150156 closed and combined with this request May 13, 2016. M. Cotter See OESR 150156 summary of problem(s) and summary of action taken for details. Update 25-SEP-2017: To be updated when test system is migrated to Camstar 7 and validated 20, AUG 2018 Same | Matt Brule |  | 19
150275 | Test Engineering | - | 2015-09-15 | TE82520 | X | International | Chau Nguyen | OFI |  |  | Please split drawing 999029399 into 3 separate drawings for 3 parts, and generate an additional ballooned assembly drawing to show how 3 parts assembled together. Drawing 999035654 to show 3 holes on the Carousel, which were not shown on the original drawing, please use it for 1 of the 3 separate drawings. | This is to help Model Shop divide the job clearer and easier to avoid duplicating unneeded parts. Jan 06, 2016: Ken approved to have Steve coordinate the work with other MEs. April 24, 2017: Consider moving to to 2017 List to be done when on-going Tool Drawing Project completed. June 28, 2017: OFI | Stephen Wright |  | 19
161127 | Test Engineering | 735002132 | 2016-10-04 | NA | NA | International | Desmond | OFI |  |  | Sigmasure is already implemented in Milford. WPPL is creating the access for Sigmasure 7 account. *** To update the document of 735002132 Genealogy that to use the manual document 730002048 as an alternative genealogy instruction in the event that the sigmasure server is down. | March 28, 2017: In discussion with Dan and team, we have de-prioritized the creation of a genealogy procedure. The activity has been put in the parking lot for now. May 30, 2017: Discuss with WPPL to switch to OFI. Jaaffar will discuss with Eric to put a plan together to migrate to Sigma 7. | Burt Bliss |  | 19
161329 | Test Engineering | 735002236 | 2016-12-20 | TE83026 | N/A | International | Wesley Wong | OFI |  |  | Problem Statement: 186015049 does not have Test Records Uploaded to SigmaSure6 Findings: i) Checked sigmasure6 web, no test records seen. ii) Checked on local testers TE83026-1 & -2, xml files are seen residing in sendbuffer folder iii) TSP 735002236 rev F mentions the test record data will be uploaded to sigmaquest collection system iv) Product 186015043 that runs on the same testers have test records uploaded | I reviewed the test systems. It appears that part number 186015049 is not properly apply the station GUID to the xml record. This appears to have been the case for some time. This will be fixed upon migration to the new Camstar server in 2017. Moved to OFI, to be migrated to Sigmasure 7 after updates to TE83013, slated to start later in 2017. 20, AUG 2018 Same | Matt Brule |  | 19
171474 | Test Engineering | WAT069481 | 2017-02-15 | TE82519 | N/A | International | Wesley Wong | OFI |  |  | Proposed to change below in TSP, WAT069481 rev4 i) On steps 3.3: Change "CRT" to "Monitor" to make it generic. | Feb 22, 2017: Edwin will discuss with Jaaffar to reach resolution. (WPPL has reviewed TSP thoroughly with operators at CM to detect as many changes as possible to correct when submit OESR to Milford.) Apr 24, 2017: Paul to submit ECR ? Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI | Paul Deschenes |  | 19
171475 | Test Engineering | 735000287 | 2017-02-15 | TE82625 | N/A | International | Wesley Wong | OFI |  |  | Proposed to change below in TSP's coversheet, 735000287_Rev B I) To remove "TEST EQUIPMENT PLANT LOCATION: Milford" and to utilize new document template coversheet. | May 30, 2017: OFI July 26: This is still an OFI. Updating this procedure requires specification alignment which will be performed and ECR'd if more significant changes are needed for TSP 735000287. | Tyler Modelski |  | 19
171483 | Test Engineering | WAT210700 | 2017-02-15 | TE82526 | N/A | International | Wesley Wong | OFI |  |  | Proposed to change below section in TSP, WAT210700_rev9 a)On step 1.13: Change "CRT " to "Monitor" to make it generic. b)On step 1.13: - Change "password=tester" to "password=PUMPTRAY-X", where X is the pc number, to allow current practice. OR - Remove the login credentials information | Feb 22, 2017: Edwin will discuss with Jaaffar to reach resolution. (WPPL has reviewed TSP thoroughly with operators at CM to detect as many changes as possible to correct when submit OESR to Milford.) Apr 24, 2017: Paul to submit ECR ? Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI | Paul Deschenes |  | 19
171504 | Test Engineering | 735000300 | 2017-02-21 | TE82656 | N/A | International | Wesley Wong | OFI |  |  | Proposed to change below in TSP's coversheet, 735000300_Rev C. i) To remove "TEST EQUIPMENT PLANT LOCATION: Milford" and to utilize new document template coversheet. | Feb 22, 2017: Edwin will discuss with Jaaffar to reach resolution. (WPPL has reviewed TSP thoroughly with operators at CM to detect as many changes as possible to correct when submit OESR to Milford.) Apr 24, 2017: WPPL to submit ECR ? Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. | Wesley Wong | Mark Jussaume | 19
171546 | Test Engineering | 735000615 | 2017-03-03 | TE82760 | N/A | International | Wesley Wong | OFI |  |  | Proposed to change below section in TSP, 735000615_revP a)On Pg3, step1.3: - Change "password=tester" to "password=TE82760-X", where X is the pc number, to allow current practice. OR - Remove the login/password credentials information | Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI | Eric Turner |  | 19
171547 | Test Engineering | 735002056 | 2017-03-03 | TE82984 | N/A | International | Wesley Wong | OFI |  |  | Propose to change below in TSP, 735002056_Rev B On pg 2, Equipment Requirement, i) Change "TE 82948" to "TE 82984" ii) Add 800000204, TUV Detector Refurb iii) Add 800000240, BSM Refurb iv) Add 800000227, Sample Mgr Refurb v) Change "5.75 mg/L" to "0.56 g/L" - This is what Venture is using vi) Remove "0.2" - No evidence of revision control in empower project on empower PC vii)On Pg4, To power up and log on to the workstations: - Change "password=tester" to "password=TE82984-X", where X is the pc number, to allow current practice. OR - Remove the login/password credentials information viii)On Pg9, step 5.13, remove "9.10." ix) On Pg9, step 5.16, change "4 column" to "5 column" - In empower report, there are 5 columns (1-4 & BP) | These suggestions will be addressed when time allows. April 26, 2017: TMD should be raised for May Audit at WPPL or put into OFI if agreed by WPPL. May 30, 2017: Mike has responded and waiting for feedback from WPPL. Desmond to respond to Mike's comments. | Michael Cotter |  | 19
171551 | Test Engineering | 735002011 | 2017-03-06 | TE83032 | NA | International | Yeap Siew Yee | OFI |  |  | Proposed to change below in TSP, 735002011 Rev C, Page 3 of 4, section 2.0, step 2.2: 1) To change 'Password tester' to 'Password TE83032-#' ,# refer to PC number or to remove login password from TSP | Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI | Eric Turner |  | 19
171552 | Test Engineering | 735000629 | 2017-03-06 | TE83031 | NA | International | Yeap Siew Yee | OFI |  |  | Proposed to change below in TSP, 735000629 Rev E, Page 3 of 4, section 2.0, step 2.2: 1) To change 'Password tester' to 'Password TE83031-#', # refer to PC number or to remove login password from TSP | Apr 26, 2017: Jaaffar and Edwin will discuss to put these non-technical changes into parking lot (OFI) or have a plan for WPPL to submit ECR if necessary. May 30, 2017: OFI | Eric Turner |  | 19
171576 | Test Engineering | 735002011 | 2017-03-15 | TE83032 | 1 | International | Wesley Wong | OFI |  |  | AcQuity FLR Detector Optics Alignment Station GUI on TE83032-1 shows optic bench selection as "Altus A30 FLR Optics Bench - 279003512". The GUI should display selection as "ASSY, OPTICS BENCH, FLR, ALTUS A-30 - 279003293" See attachments in email | March 27, 2017: FLR Optics Bench Alignment System qualification has been executed and is in sign off. Apr 24, 2017: Will be done after spec alignment - will set status to OFI. | Eric Turner |  | 19
171586 | Test Engineering | 735001199 | 2017-03-20 | TE82869 | NA | International | Yeap Siew Yee | OFI |  |  | Propose to change the following in TSP 735001199 Rev B: i) Test step 1.3 (Page 3 of 4): To change '120-volt AC' to '240-volt AC' ii) Test step 3.4, 3.5, 3.7, 3.8 (Page 3 of 4): To change 'green LED' to 'LED'. Current LED on tester are red in color. | i) The tester will accept universal line voltage 90 to 240 VAC, 50/60 hz. This step will be changed. ii) The two testers have different color LED's. The functions of the LEDs are labeled so color is not important. These steps will change. Edits to be completed and ECR submitted as time allows. | Chau Nguyen | Richard Davis | 19
171608 | Test Engineering | 735002236 | 2017-03-28 | TE83026 | NA | International | Yeap Siew Yee | OFI |  |  | Propose to change the following in TSP# 735002236 Rev F: (1) Page 2 of 4, Equipment Required : To change TE83026-1 to TE83026 (2) Page 3 of 4, Test System Setup: To change Password: tester to Password: TE-83026-#, #refer to PC number; or to remove login password from TSP. | TE83026 is on the list to be updated to move to Sigmasure 7 and correct an ECord test issue, this procedure will be updated then. 20, AUG 2018 Next priority after wrapping up HTCH tester and Hi-Pot updates | Matt Brule |  | 19
171694 | Test Engineering | NA | 2017-04-27 | N/A | N/A | International | Yeap Siew Yee | OFI |  |  | Files pending in the 'sendbuffer' folders yet to be sent out. Affected testers as below: 1) TE83013 (HTCH tester) 2) TE83013-2 (HTCH tester) 3) TE83026-1 (CM tester) 4) TE83026-2 (CM tester) | June 27, 2017: Both of these to be fixed with test software upgrade for SIgmasure, TE83026 is first followed by TE83013. 27-JUL-2017: this will be corrected during up-commoing software revision for Camstar. Sept 26, 2017: Work has started on an update to TE83013, TE83026 is next. 30 JAN 2018 Working on TE83013, complete update (win 10, SIgmasure 7, fix OESRs etc) 27-FEB-2018 Actively working on this, Test running working on EDV and bugs as they are found 24-APR-2018 TE83013-EDV drafted, code complete, all hardware delivered, working on PC config, build, validation. TE83026 will be started soon 20, AUG 2018, Fixing Ecord test bug, one duplicate tester built. | Matt Brule |  | 19
171705 | Test Engineering | 735002229 | 2017-04-28 | TE83026 | N/A | International | Yeap Siew Yee | OFI |  |  | Proposed to change the following in TSP 735002229 Rev H: 1) Section 1.0 UUT Test Setup: To remove steps 1.5 & 1.7. Pre-heaters and I-button will only be installed during test, as per on-screen instructions. | 27-JUL-2017, this will be corrected during up-coming software revision for Camstar 20, AUG 2018 Same | Matt Brule |  | 19
171852 | Test Engineering | N/A | 2017-07-05 | TE82620 | N/A | International | Desmond Ann | OFI |  |  | The Keithley module in the tester is faulty and have EOL. Based on the last updates, Milford will be replacing this tester with a new tester and the current old tester have already been returned to Milford. See attached on the history. This OESR is created in place of OESR 185 created in 2014 to close out 2014 List. | (Prior updates available in prior monthly meetings' minutes.) Jan 24, 2017: On list behind current CM projects per Matt. Consider moving it to 2017 List in March/April 2017 when Wesley comes to Milford for products' training and discussion of projects' works. Feb 21, 2017: Currently in hold, project is behind column product updates and validations. April 24, 2017: On to-do list. June 27, 2017: Status unchanged. WPPL suggests to consider moving it onto 2017 List. Currently testing in instrument, new tester will be developed as priorities allow. 30-JAN-2018 Same | Tyler Modelski | Matt Brule | 19
172017 | Test Engineering | N/A | 2017-09-06 | TE82520 | N/A | International | Chau Nguyen | OFI |  |  | Transfer from OESR 150275: to close out 2015 To split drawing 999029399 into 3 separate drawings for 3 parts, and generate an additional ballooned assembly drawing to show how 3 parts assembled together. Drawing 999035654 to show 3 holes on the Carousel, which were not shown on the original drawing, suggest to use it for 1 of the 3 separate drawings. | This is to help Model Shop divide the job clearer and easier to avoid duplicating unneeded parts. Jan 06, 2016: Ken approved to have Steve coordinate the work with other MEs. Feb 24, 2016: Plan to complete in Q2/Q3 of 2016 due to the plan to complete tool drawings for Flex in Q2/Q3 (30% completion so far). Sept 06, 2016: Low priority but moving along as planned. Oct 25, 2016: Chau might work on this to close it out after 2016 WPPL Trip. Jan 04, 2017: Chau will discuss with Steve (Frank) to start working on it when time permits. April 24, 2017: Move to 2017 List to be done when on-going Tool Drawing Project completed. June 28, 2017: OFI to be taken upon after Tool Drawing Project complete. | Stephen Wright | Chau Nguyen | 19