## Slide 1
Fallout Rate: 0.3ml/min flowrate
1* - Trim R1 & R2 tubes at tester 1 on 1st June 18 as per TSP (19/65 -29% failed at TE#1 until 31st Oct 18)
2* - Implemented monthly test tube change started from 26 Jul 18
3* - Trim R1 & R2 tubes at tester 24 on 5th Sep 18 as per TSP and monitor result
(2/21 (9.5%) failed at TE#24 until 31 Oct 18)
1*
2*
Chart: Total tested Qty Vs Percentage of failure
| 42995 | 43026 | 43057 | 43088 | 43119 | 43150 | 43181 | 43212 | 43243 | 43274 | 43305 | 43313 | 43361 | 43391
Test Qty | 119 | 129 | 116 | 23 | 87 | 90 | 110 | 101 | 94 | 77 | 109 | 110 | 96 | 55
% of failure | 0.02521 | 0.015504 | 0.051724 | 0.086957 | 0.18391 | 0.077778 | 0.090909 | 0.18812 | 0.26596 | 0.11688 | 0.21101 | 0.10909 | 0.0625 | 0.072727
3*

## Slide 2
Fallout Rate: Failure Mode Vs Test Station (Jun-Oct 18)
Chart: Jun-Oct 18 Failure by Teser ID
| Test Flow Rate Accuracy 0.3ml/min | Test Input D Prime with MeOH | Test Static Leak Decay VV to Mixer1 to VV | Test Flow Rate Accuracy 2.5ml/min | Test Seal Wash Pump Prime/Flow | Test Compositional Accuracy - AC95%AU | Test Input C Prime with MeOH | Test Input B Prime with MeOH | Test Stop Flow Input SHORT | Test Input A Prime with MeOH | Test Alarms at Bootup | Test Compositional Accuracy - BD5%AU | Test Static Leak Decay VV to Mixer2 to VV | Test Dynamic Leak - Accumulator Leak Rate | Test Gradient Start Input SHORT | Test Dynamic Leak - Primary Leak Rate | Test High Pressure | Test Degasser Pressure | Test Leak Detect Connected | Test Compositional Accuracy - AC5%AU | Test Compositional Accuracy - BD95%AU | Test LEDs | Test Reboot Firmware | Test Loader Download
TE82943-1 | 19 | 13 | 4 | 1 | 1 | 5 | 5 | 2 | 0 | 0 | 1 | 0 | 1 | 2 | 2 | 1 | 1 | 0 | 1 | 0 | 2 | 0 | 0 | 1
TE82943-19 | 8 | 15 | 2 | 2 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
TE82943-20 | 7 | 4 | 1 | 2 | 4 | 3 | 3 | 4 | 1 | 1 | 1 | 1 | 2 | 2 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0
TE82943-21 | 6 | 7 | 2 | 0 | 0 | 4 | 2 | 0 | 0 | 3 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
TE82943-22 | 9 | 3 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 2 | 0 | 1 | 0 | 0 | 0 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0
TE82943-24 | 13 | 3 | 1 | 6 | 5 | 5 | 3 | 2 | 2 | 2 | 2 | 2 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0

## Slide 3
FPY based on tester station in SigmaSure

## Slide 4
Failure rate by failure modes

## Slide 5
Can we replace these tunings to improve the FPY of 0.3ml/min flow test?
Do you have parts for us to do this experiment for flow rate yield improvement?