Alliance System
Separations Modules e2695 and e2695D

Installation Checklist

# General information

The installation checklist is completed by a Waters field service engineer (FSE) during the installation.

Note:	For systems containing a mass spectrometer: After installing the Alliance® system on the bench, complete the appropriate mass spectrometer Installation Checklist and its performance tests before completing this checklist.

Note:	If the customer purchased Qualification, you need only complete the Customer Training section of this checklist.

## Purpose

The installation checklist:

Ensures performance of all necessary steps in the installation

Provides proof of installation by a qualified Waters engineer

Demonstrates to the customer that the system functions according to installation specifications

Ensures that the customer is trained in basic operation of the system

## Safety advisory

There are various hazards associated with installing and operating this equipment. Follow the precautions detailed in the relevant system user manuals.

## Additional documents required

If the system includes a QDa, complete the appropriate installation checklist.

Appropriate training checklist. Attach training checklists for each component listed in System components training (page 20).

A full list of system installation documents are available on the Portal (FSEs only).

## Completing the checklist

If the system requires installation of additional components not included in the checklist, complete the appropriate checklist.

Note:	If this is a clinical diagnostic system, there may be different requirements.

If the customer has not purchased Qualification, the FSE must:

- Complete all checklist sections

Sign and date the Confirming the installation (page 24) and leave a signed copy of the checklist with the customer

Send or fax an electronic or printed copy of the checklist to the regional office

If the customer purchased Qualification, the FSE must:

- Complete only the checklist for Customer training (page 19)

Complete the Customer/system summary (page 23)

Sign and date the Confirming the installation and leave a signed copy of the checklist with the customer

Send or fax an electronic or printed copy of the checklist to the regional office

Note:	You do not need to fill out any other sections of the checklist.

# System setup

## Preparing the site

| The site is prepared as specified in the appropriate site preparation guide | Yes No N/A | Comments: |
| --- | --- | --- |

## Unpacking and positioning the system

Note:	Report any damage or discrepancies to Waters Global Services.

| Shipping containers and unpacked items are not damaged | Yes No N/A | Comments: |
| --- | --- | --- |
| All installation components are available | Yes No N/A | Comments: |
| System components are properly positioned in their final location | Yes No N/A | Comments: |

## Setting up the separations module

Note:	Do not power-up the module until all fluidic and signal connections are complete.

| Separations module type (select one of the following) | e2695 e2695D | Comments: |
| --- | --- | --- |
| Column heater or column heater/cooler (if applicable) is available | Yes No N/A | Comments: |
| AC power cord is connected | Yes No N/A | Comments: |
| Solvent bottle tray is available | Yes No N/A | Comments: |

## Making signal connections

| Ethernet cable (if applicable) is connected | Yes No N/A | Comments: |
| --- | --- | --- |
| IEEE-488 interface (if applicable) is connected | Yes No N/A | Comments: |
| Analog I/O connections required by the application (if applicable) are connected | Yes No N/A | Comments: |
| RS-232 connection (if applicable) is connected | Yes No N/A | Comments: |

## Making fluidic connections

| Solvent inlet supply lines are plumbed to the appropriate solvent reservoirs | Yes No N/A | Comments: |
| --- | --- | --- |
| Plunger seal-wash inlet line is plumbed to the plunger-seal wash reservoir | Yes No N/A | Comments: |
| Waste line is plumbed to a waste container | Yes No N/A | Comments: |
| Note: Do not use the same container or solution for plunger seal wash and needle wash. (See the Operator's Guide for choice of appropriate solutions.) Needle-wash inlet line is plumbed to the needle-wash reservoir | Yes No N/A | Comments: |
| Needle-wash outlet and waste lines are plumbed to a waste reservoir | Yes No N/A | Comments: |
| Spill tray waste and condensation tray waste lines are plumbed to the waste reservoir | Yes No N/A | Comments: |
| Degasser vent line is plumbed to a suitable fume hood (if applicable) | Yes No N/A | Comments: |
| Sample management outlet tubing is plumbed to the column inlet or column switching valve | Yes No N/A | Comments: |
| Note: Purge the column outlet line to an appropriate receptacle before making the fluidic connection between the column outlet and the detector inlet. Column is installed | Yes No N/A | Comments: |
| Separations module is powered-up | Yes No N/A | Comments: |
| System passed the initialization diagnostics | Yes No N/A | Comments: |

## Setting the separations module parameters

| Separations module is powered-up | Yes No N/A | Comments: |
| --- | --- | --- |
| System passed the initialization diagnostics | Yes No N/A | Comments: |
| Separations module parameters are set appropriately for this configuration | Yes No N/A | Comments: |
| IEEE address is set (if applicable) | Yes No N/A | Comments: |

## Setting up the LC detector

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Power cord is connected and voltage is set (if applicable) | Yes No N/A | Comments: |
| As applicable, the IEEE or Ethernet cable is connected between the separations module and the detector | Yes No N/A | Comments: |
| Flowcell inlet is plumbed to the column outlet | Yes No N/A | Comments: |
| Flowcell outlet is plumbed to a waste container | Yes No N/A | Comments: |
| Drip tray is plumbed to a waste container | Yes No N/A | Comments: |
| Note: Do not power-up the detector until detector fluidic and signal connections are complete. Detector is powered-up | Yes No N/A | Comments: |
| Detector passed the initial diagnostics | Yes No N/A | Comments: |
| IEEE address is set | Yes No N/A | Comments: |
| USB cable is connected between the detector and the system workstation (2432 only) | Yes No N/A | Comments: |

## Setting up the computer workstation

| Empower® or MassLynx™ Chromatography Manager workstation is installed | Yes No N/A | Comments: |
| --- | --- | --- |
| Empower or MassLynx software is installed on the workstation | Yes No N/A | Comments: |
| IEEE cable is connected | Yes No N/A | Comments: |
| Ethernet cable is connected (if applicable) and IP address is assigned | Yes No N/A | Comments: |
| Power is on and Empower software has found all system components | Yes No N/A | Comments: |

## Setting up the Waters Fraction Manager – Analytical (WFM-A)

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| The WFM-A is positioned in the system (Alliance System Site Preparation Guide, 715001178) | Yes No N/A | Comments: |
| Refer to Installing the Fraction Manager – Analytical (715003927) for guidance on the following: |  |  |
| WFM-A fluidic connections are made | Yes No N/A | Comments: |
| WFM-A power and signal connections are made | Yes No N/A | Comments: |
| WFM-A is calibrated, primed, and configured | Yes No N/A | Comments: |
| Appropriate delay coil is installed | Yes No N/A | Comments: |

# Preparing for the tests

For instructions on preparing for the tests, refer to the system installation manual or performance test.

| Highest-quality (HPLC-grade or better) water, solvents, buffers, and additives are available | Yes No N/A | Comments: |  |
| --- | --- | --- | --- |
| System diagnostic tests passed | Yes No N/A | Comments: |  |
| System is set up for performance tests as directed | Yes No N/A | Comments: |  |
| Solvent manager and sample manager are primed and flushed | Yes No N/A | Comments: |  |
| System is flushed to waste | Yes No N/A | Comments: |  |
| Test samples are available | Yes No N/A | Comments: |  |
| Lot number: |  |  |  |
| Part number: |  |  |  |
| Expiration date: |  |  |  |
| Solvent/water grade: |  |  |  |

# Running diagnostics

## Performing sample management system verification test runs

| Needle wash system is primed |  | Yes No N/A | Comments: |  |
| --- | --- | --- | --- | --- |
| e2695 or e2695D needle seal calibration is performed and the values are recorded |  | Yes No N/A |  | Comments: |
| The bottom count is: |  |  |  |  |
| The top count is: |  |  |  |  |
| The seal width is: |  |  |  |  |
| e2695 injector purge and compression test is performed |  | Yes No N/A | Comments: |  |

## Preparing the LC detector for verification test runs

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Note: Ensure that the flowcell is securely installed. |  |  |
| Flowcell is flushed for at least 15 minutes with mobile phase Note: For the 2489 and 2998 detectors use low UV absorbing (<220 nm) mobile phase | Yes No N/A | Comments: |
| 2489 detector is allowed to warm for at least 30 minutes | Yes No N/A | Comments: |
| Wavelength calibration for the 2489 detector is performed Calibration successful: Max deviation since last calibration: nm (must be less than 1.0 nm) | Yes No N/A Pass Fail | Comments: |

| 2998 detector is allowed to warm for at least for 60 minutes | Yes No N/A | Comments: |
| --- | --- | --- |
| Erbium calibration for the 2998 detector is performed Optimized exposure time: msec Graph Type: Erbium absorbance X Axis: Wavelength Wavelength calibration: λ 1: λ 512: - Calibration verification results: Maximum deviation: nm (must be less than 1.0 nm) | Yes No N/A Pass Fail | Comments: |
| 2414 refractometer is allowed to warm for at least 30 minutes (in some cases, the 2414 may require 24 hours to equilibrate) | Yes No N/A | Comments: |
| 2475 lamp is allowed to warm for at least 30 minutes | Yes No N/A | Comments: |

## Preparing the QDa detector for verification test runs

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Mobile phase and other solvents are prepared | Yes No N/A | Comments: |
| ESI performance, gradient test, and blank samples are prepared | Yes No N/A | Comments: |
| QDa detector is in Operate | Yes No N/A | Comments: |
| The two QDa detector status LEDs are green | Yes No N/A | Comments: |

| The following tasks are complete (Alliance System Performance Test – Empower-based with a QDa Detector, 715004380 or Alliance System Performance Test – MassLynx 4.1-based with a QDa Detector, 715004352): Solvent lines are inserted into their corresponding solvent reservoirs Each solvent line in the e2695 is dry- and wet-primed Seal wash and needle wash diagnostics are complete Purge injector function was run for 20 sample loop volumes System was flushed to waste (with no column attached) with 50:50 acetonitrile/water at 2.0 mL/min for 10 minutes Column was connected and installed into the column heater, and the adjust seals diagnostic was run Blank and sample vials were placed in the carousel | Yes No N/A | Comments: |
| --- | --- | --- |

Running the tests

For instructions on running the tests for each system component, refer to the appropriate installation manual or system performance test document (FSEs only).

| Performance tests for each configured system component were successfully completed | Yes No N/A | Comments: |
| --- | --- | --- |

# System performance test results

For instructions on test procedures, refer to the system installation manual or performance test.

| Performance test result checklists are complete for each system component, as necessary (see attached) | Yes No N/A | Comments: |
| --- | --- | --- |

## LC detector tests

Perform the following verification test during the customer training session. This test is used to show consistency in the retention times and peak areas. Where possible, use suggested column Xbridge 4.6 x 50 mm or similar. For each test, perform six replicate injections of 20 µL injection volume at a 1 mL/minute rate. You will need the following material:

Note:	You may need to reduce the 2475 gain to keep peaks on scale.

| This section applies to this configuration: Yes No |  |  |  |  |
| --- | --- | --- | --- | --- |
| Sample part number | Compound | Detector | Detector settings | Suggested mobile phase |
| 700001713 | Caffeine | 2414 | sens: 64 Run time: 3 min | 80% water, 20% methanol |
| 700002387 | Acetominophen | 2424 | Gas pressure: 30 psi Gain: 135 Flow rate (low): 1.0 mL/min Run time: 5 minutes | 90% water, 10% methanol |
| 700005906 (0.06 mg/mL ampule) | Caffeine | 2487, 2489, 2996, 2998 | 273 nm Run time: 2 min | 70% water, 30% methanol |
| WAT047685 | Anthracene | 2475 | ex: 249, em: 402 Gain: 10 Run time: 3 min | 20% water, 80% acetonitrile |

## LC detector repeatability tests

### UV and PDA detector test results

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Compound | Peak | Required %RSD (over 6 injections) |
| Caffeine |  | Area count 1.5 |
|  |  | RT 1.0 |

### RI detector test results

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Compound | Peak | Required %RSD (over 6 injections) |
| Caffeine |  | Area count 1.5 |
|  |  | RT 1.0 |

### 2424 ELS detector test results

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| Compound | Peak | Requirement (over 6 injections) |
| Acetominophen |  | S/N >1000 (for low) |
|  |  | S/N >4000 (for high) |
|  |  | Area count <5 %RSD (for high- and low-flow nebulizers) |
|  |  | RT 1.0 %RSD |

### 2475 detector test results

| This section applies to this configuration: Yes No |  |  |  |
| --- | --- | --- | --- |
| 2475 detector passes the start-up diagnostic routine running HPLC-grade water |  | Yes No N/A | Comments: |
| Peaks are visible, symmetrical, and integrated correctly |  | Yes No N/A | Comments: |
| Compound | Area* (over 6 injections) | Required %RSD | Test passed? |
| Anthracene |  | ≤1.5% | Yes No N/A |
|  | Peak | %RSD (over 6 injections) |  |
|  |  | Area count 1.5 %RSD |  |
|  |  | RT 1.0 %RSD |  |

*Print the six results using the Alliance report, Component Summary with Chromatogram.

### 2432 detector test results

| This section applies to this configuration: Yes No |  |  |  |  |
| --- | --- | --- | --- | --- |
| Compound | Retention time (over 6 replicates) | Standard deviation (SD) | Required SD | Test passed? |
| Chloride | s |  | ≤1.5 second (0.025 mins) | Yes No N/A |
| Bromide | s |  | ≤1.5 second (0.025 mins) | Yes No N/A |
| Compound | Area %RSD (over 6 replicates) |  | Requirement | Test passed? |
| Chloride |  |  | ≤2.0% | Yes No N/A |
| Bromide |  |  | ≤2.0% | Yes No N/A |

## WFM-A performance test
(for systems with TUV, PDA, ELSD, and QDa detectors only)

Refer to 715003927  for detailed instructions.

| This section applies to this configuration: Yes No |  |  |
| --- | --- | --- |
| With collection disabled, the system precision performance test sample was run and processed, and the results were saved | Yes No N/A | Comments: |
| Appropriate peak (caffeine for PDA, TUV, and ELSD detectors, or sulfadimethoxine for QDa detectors) was integrated, the start-time and end-time were determined, and the Fraction Manager instrument method was updated with the start and end times | Yes No N/A | Comments: |
| With collection enabled, the system precision performance test sample was rerun, processed, and collected in a sample vial | Yes No N/A | Comments: |
| With collection again disabled, the previously collected sample was injected, run, and processed | Yes No N/A | Comments: |
| Appropriate peak (caffeine for PDA, TUV, and ELSD detectors, or sulfadimethoxine for QDA detectors) is identified from the collected sample | Yes No N/A | Comments: |

## Completing the installation

| Instrument covers are in place | Yes No N/A | Comments: |
| --- | --- | --- |
| Installation projects were copied to a backup source | Yes No N/A | Comments: |
| You have logged out of service mode | Yes No N/A | Comments: |
| Service record date in the mass spectrometer console is set | Yes No N/A | Comments: |

# Customer training

The FSE trains the customer in the basic operation and maintenance of the system. If the customer desires more advanced training, they may contact their Waters representative for full details of Waters training courses.

For best training results, the user should be familiar with the system operator’s guide or guides.

## Safety training

Review the safety sections of the following documents.

| Safety sections of relevant user guides: - Alliance e2695 Separations Module Operator’s Guide (715003794) | Yes No N/A | Comments: |
| --- | --- | --- |

## Routine maintenance training

Review the maintenance chapters of the relevant operator’s guides.

| Maintenance chapters of relevant operator’s guides (715003794) | Yes No N/A | Comments: |
| --- | --- | --- |

## System components training

The Waters engineer will introduce the following topics while demonstrating the system performance specifications. Waters intends only to provide familiarization with the basic operation and maintenance of the system, and does not attempt to provide comprehensive training during installation. For full details of training courses available, please contact your Waters representative.

| Basic operation |  |  |
| --- | --- | --- |
| Primary user introduction to Empower or MassLynx software operation basics | Yes No N/A | Comments: |
| Configuring a chromatographic system | Yes No N/A | Comments: |
| Using Empower wizards (if applicable) | Yes No N/A | Comments: |
| New method set | Yes No N/A | Comments: |
| Processing method | Yes No N/A | Comments: |
| Sample set method | Yes No N/A | Comments: |
| Creating MassLynx Inlet methods (if applicable) | Yes No N/A | Comments: |
| Using the Run Samples window | Yes No N/A | Comments: |
| Interactively creating a processing method using Review | Yes No N/A | Comments: |
| Empower or MassLynx customer was shown how to review results | Yes No N/A | Comments: |
| Empower customer was shown how to generate a report | Yes No N/A | Comments: |
| Empower customer was shown how to back up and restore projects | Yes No N/A | Comments: |
| Review workflow on the ACQUITY QDa Detector Operator’s Quick Reference Card (715003957) | Yes No N/A | Comments: |

## Software

| Opening the software | Yes No N/A | Comments: |
| --- | --- | --- |
| ACQUITY QDa detector console page, sample list, inlet and method editors | Yes No N/A | Comments: |
| Switching into Operate and Standby | Yes No N/A | Comments: |
| Acquiring data | Yes No N/A | Comments: |
| Acquiring data from additional detectors when the QDa is in Standby mode (flow to the QDa must be disconnected) | Yes No N/A | Comments: |

## Waters’ resources and support training

| Location of operator’s overview and maintenance guides 1 | Yes No N/A | Comments: |
| --- | --- | --- |
| Accessing online Help and registering with the product online community | Yes No N/A | Comments: |
| Support information and how-to videos (www.waters.com) | Yes No N/A | Comments: |
| Education and training (www.waters.com) | Yes No N/A | Comments: |
| Obtaining spares (www.waters.com) | Yes No N/A | Comments: |
| Using Connections INSIGHT and submitting requests (if connected) | Yes No N/A | Comments: |
| Using iRequest™ to submit online service requests | Yes No N/A | Comments: |
| Benefits of Total Assurance Warranty (TAW) | Yes No N/A | Comments: |
| Benefits of Waters Analytical Standards and Reagents for supplying certified samples (www.waters.com) | Yes No N/A | Comments: |
| Obtaining technical, software, and spares support locally | Yes No N/A | Comments: |
| Waters Global Services | Yes No N/A | Comments: |

# Customer/system summary

Customer information

| Company: |  |  | Customer name: |  |  |
| --- | --- | --- | --- | --- | --- |
| Address 1: |  |  | Telephone: |  |  |
| Address 2: |  |  | Engineer name: |  |  |
| City: |  |  | SAP order: |  |  |
| State/prov/region: |  |  | Install date: |  |  |
| Postal code: |  |  |  |  |  |
| Country: |  |  |  |  |  |

System information

| Sample manager: |  |  | Solvent manager: |  |  |
| --- | --- | --- | --- | --- | --- |
| Sample organizer: |  |  | Autosampler: |  |  |
| Pump 1 model #: |  |  | Pump 1 serial #: |  |  |
| Pump 2 model #: |  |  | Pump 2 serial #: |  |  |
| Pump 3 model #: |  |  | Pump 3 serial #: |  |  |
| Detector 1 model #: |  |  | Detector 1 serial #: |  |  |
| Detector 2 model #: |  |  | Detector 2 serial #: |  |  |
| Detector 3 model #: |  |  | Detector 3 serial #: |  |  |
| Detector 4 model #: |  |  | Detector 4 serial #: |  |  |
| Mass spec model #: |  |  | Mass spec serial #: |  |  |
| Separations system: |  |  | Roughing/dry pump: |  |  |
| Column manager: |  |  | Column heater/cooler: |  |  |
| 30-cm column heater/cooler: |  |  | High-temp column heater: |  |  |
| Software model: |  |  | Software serial #: |  |  |
| Host PC: |  |  | Other 1: |  |  |
| Other 2: |  |  | Other 3: |  |  |
| Other 4: |  |  | Other 5: |  |  |

# Confirming the installation

Confirm that the installation was successfully completed.

| Installation and setup was performed and verified through the completion of this checklist | Yes No N/A | Comments: |
| --- | --- | --- |
| All tests pertinent to this installation were performed and have met the defined specification criteria | Yes No N/A | Comments: |
| The customer was trained in all aspects of instrument operation contained in this checklist | Yes No N/A | Comments: |

| I confirm on behalf of |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Company Name |  |  |  |  |  |
| that the instrument was installed in accordance with the procedures detailed in this installation checklist, and that instrument operation was demonstrated by conducting all tests relevant to this installation. The test specifications defined in this installation checklist were achieved. |  |  |  |  |  |  |  |
| Signed: |  |  |  | Date: |  |  |  |
| Customer signature |  |  |  |  |  |  |  |
| Printed name |  |  |  |  |  |  |  |
| I certify that the installation was successfully completed. |  |  |  |  |  |  |  |
| Signed: |  |  |  | Date: |  |  |  |
| Waters Engineer signature |  |  |  |  |  |  |  |
| Printed name |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |

# Reference documents

For details on installing and operating the system, see the following resources/documents for Alliance e2695/e2695D Systems:

Reference Information

Part Locator

Release Notes

Downloads

Technical Notes

User Manuals

Performance Maintenance (PM) Protocols

Attachments

Use the space below to add any necessary attachments (screen captures, chromatograms, spectra, and so on) to this installation checklist.

| Any necessary attachments are included below | Yes No N/A | Comments: |
| --- | --- | --- |

1 Shipped with the system and also available in the Waters Support Library.
2 Waters Global Services optimizes Waters products with service, support, upgrades, training, and Waters Quality Parts®.