## Carryover Test Summary
Summary
|  | Carryover % (Median) | Carryover % (Min) | Carryover % (Max) | Carryover % (SD)
Old Seal Pack | Old Final Test Template | 0.002 | 0.0007 | 0.0086 | 0.0018957
| New Final Test Template | 0.0018 | 0.0005 | 0.0039 | 0.00088482
New Seal Pack | New Final Test Template | 0.0013
Remark:
1 | All the result above collected using old Seal Pack (PN: 289001994) except for 1x GREEN data point which represent new seal pack.
2 | All instrument are loaded with Firmware 3.03.
3 | No Significant difference in carryover result between new and old Final Test Template
4 | No Significant difference in carryover result between new and old Seal Pack under Firmware 3.03

## CarryOver Test
Daily Alliance CarryOver Test
Detector S/N 1387E729A |  |  | Column S/N 01473126213809
Kit Lot Number & Expiration date : W01031720 ( 01 Mar 2020 )
Unit | Date | eAlliance S/N | Carryover Value w Old Template | Carryover Value w New Template | Current Spec<0.01 | Result Pass/Fail | Remarks
1 | 2017-12-19 | M17SM4911A | 0.0007 |  | 0.01 | Pass
2 | 2017-12-20 | M17SM4889A | 0.0015 |  | 0.01 | Pass
3 | 2017-12-21 | L16SM4368A | 0.0022 |  | 0.01 | Pass | RMA
4 | 2017-12-22 | L16SM4371A | 0.0018 |  | 0.01 | Pass | RMA | No High Blank
5 | 2017-12-26 | M17SM7568A | 0.0034 |  | 0.01 | Pass
6 | 2017-12-27 | M17SM4008A | 0.0012 |  | 0.01 | Pass
7 | 2017-12-28 | M17SM4031A | 0.0014 |  | 0.01 | Pass
8 | 2017-12-29 | M17SM4057A | 0.0013 |  | 0.01 | Pass
9 | 2017-12-30 | M17SM4071A | 0.004 |  | 0.01 | Pass
10 | 2018-01-02 | M17SM4075A | 0.0019 |  | 0.01 | Pass
11 | 2018-01-03 | M17SM4082A | 0.0072 |  | 0.01 | Pass
12 | 2018-01-04 | A18SM4103A | 0.004 |  | 0.01 | Pass
13 | 2018-01-05 | A18SM4116A | 0.0027 |  | 0.01 | Pass
14 | 2018-01-08 | A18SM4140A | 0.0008 |  | 0.01 | Pass
15 | 2018-01-09 | A18SM7616A | 0.0012 |  | 0.01 | Pass
16 | 2018-01-10 | A18SM7635A | 0.0017 |  | 0.01 | Pass
17 | 2018-01-11 | A18SM7652A | 0.0015 |  | 0.01 | Pass
18 | 2018-01-12 | A18SM4166A | 0.0018 |  | 0.01 | Pass
19 | 2018-01-15 | A18SM4196A | 0.0025 |  | 0.01 | Pass
20 | 2018-01-16 | A18SM4222A | 0.0086 |  | 0.01 | Pass
21 | 2018-01-17 | A18SM4234A | 0.001 |  | 0.01 | Pass
22 | 2018-01-18 | A18SM4254A | 0.0012 |  | 0.01 | Pass
23 | 2018-01-19 | A18SM4217A | 0.0076 |  | 0.01 | Pass
24 | 2018-01-22 | A18SM7675A | 0.0023 |  | 0.01 | Pass
25 | 2018-01-23 | A18SM7700A | 0.0037 |  | 0.01 | Pass
26 | 2018-01-24 | A18SM4290A | 0.0025 |  | 0.01 | Pass
27 | 2018-01-25 | A18SM4306A | 0.0013 |  | 0.01 | Pass
28 | 2018-01-26 | A18SM4317A | 0.0032 |  | 0.01 | Pass
29 | 2018-01-29 | A18SM4360A | 0.0017 |  | 0.01 | Pass
30 | 2018-02-05 | B18SM4396A | 0.0028 |  | 0.01 | Pass
31 | 2018-02-06 | B18SM7720A | 0.0019 |  | 0.01 | Pass
32 | 2018-02-07 | B18SM7745A | 0.0013 |  | 0.01 | Pass
33 | 2018-02-08 | B18SM4417A | 0.002 |  | 0.01 | Pass
34 | 2018-02-09 | B18SM4437A | 0.0062 |  | 0.01 | Pass
35 | 2018-02-12 | B18SM4463A | 0.0031 |  | 0.01 | Pass
36 | 2018-02-13 | B18SM4489A | 0.0017 |  | 0.01 | Pass
37 | 2018-02-14 | B18SM4482A | 0.0029 |  | 0.01 | Pass
38 | 2018-02-20 | B18SM7759A | 0.0029 |  | 0.01 | Pass
39 | 2018-02-21 | B18SM4538A | 0.005 |  | 0.01 | Pass
40 | 2018-02-27 | B18SM4574A |  | 0.0008 | 0.01 | Pass | complete test template 1.8.22 | No High Blank
41 | 2018-02-28 | B18SM4620A |  | 0.0025 | 0.01 | Pass | complete test template 1.8.22
42 | 2018-03-01 | B18SM4628A |  | 0.0018 | 0.01 | Pass | complete test template 1.8.22
43 | 2018-03-02 | B18SM4628A |  | 0.0019 | 0.01 | Pass | Complete test template 1.8.22
44 | 2018-03-06 | B18SM4628A |  | 0.0023 | 0.01 | Pass | Complete test template 1.8.23
45 | 2018-03-07 | B18SM4708A |  | 0.0017 | 0.01 | Pass | Complete test template 1.8.23
46 | 2018-03-08 | C18SM7815A |  | 0.0039 | 0.01 | Pass | Complete test template 1.8.23
47 | 2018-03-09 | C18SM7822A |  | 0.002 | 0.01 | Pass | Complete test template 1.8.23
48 | 2018-03-19 | C18SM4797A |  | 0.0005 | 0.01 | Pass | Final Test 1
49 | 2018-03-20 | C18SM7893A |  | 0.0017 | 0.01 | Pass | Final Test 1
50 | 2018-03-21 | C18SM4816A |  | 0.0017 | 0.01 | Pass | Final Test 4
51 | 2018-03-13 | C18SM7857A |  | 0.0013 | 0.01 | Pass | Complete test template 1.8.23 (new seal pack)

## Injector Tester Status
Injector Tester Status
Injector Tester | IQ Passed | Tested With Old Seal Pack | Tested With New Seal Pack | Remark
TE82503-4 | √ | √ | √ | Total 65 new Seal Packs were tested, all passed and shipped to Milford
TE82503-5 | √ | √ | √

## Alliance Final Test Station
Alliance Test Station Upgrade Status
Tester Station | Port 1 | Port 2 | Port 3 | Port 4 | Target Completion
1 | √ | √ | √ | √ | Done
2 |  |  |  |  | 2018-03-23
3 | √ | √ | √ | √ | Done
4 | √ | √ | √ | √ | Done
5 | √ | √ | √ | √ | Done