## for MTR
| Stage / Activity | Unit 1 | Unit 2 | Unit 3 | Unit 4
| MODEL | 186269506 (Alliance) | 186269506 (Alliance) | 186269506 (Alliance) | 186269506 (Alliance)
| Serial Number | Unit 1) SN L16SM4371A: | Unit 2) SN L16SM4368A: | Unit 3) SN L16SM4356A | Unit 4) SN C17SM4376A.
| FAILURE | Carryover Test Failures on 22-Feb'17 and 16-Mar'17, Sample Temp not attaining 5 degree on 22-Feb'17, and Vial missing on 12-Jun'17. | missing vial and carry over on 20-Mar'17 and 9-Jun'17. | Vial Missing error on 2-May'17, Carryover failing in customer method on 3-May'17 and 11-May'17. | Gradient test failed on 26 Jun 2017. Black particles deposited )Pressure fluctuations) on 18-Jul-2017 and Drift and Noise test results not meeting on 07 Aug 2017.
1 | Carton Box - visual Inspection | Damaged / dented and Foreign stickers on caton box | Carton Box found damaged or torn and foreign labels | Carton Box found damaged or torn and foreign labels | A recycled carton box is used (fold and damaged) Serial Number Label not visible (covered up by shipping label) (Refer to tab "Item 1" for photos)
2 | Open box - inside content | Instrument Serial Number on Carton Box not tally with physical unit. | Instrument Serial Number on Carton Box not tally with physical unit. | Instrument Serial Number on Carton Box not tally with physical unit. | Missing Start up Kit and missing Info set. Empty space filled up with bubble wrap. Instrument wrapped with bubble wrap.
3 | Overall Inspection of Instrument | Cosmetic defects such as stains, scratches and foreign material (tape) found on instrument. | Cosmetic defects such as stains and scratches found on instrument. | Cosmetic defects such as stains, scratches, paint chip off and foreign material (plastic) found on instrument. | Stain marks, foreign matters and dents (chip-off) found on top and side panels. Missing I/O A & B connector and missing screws.
4 | Open Front Door For Inspection | Mixer found stains / discoloration. Missing item (tubing) found on instrument. | Mixer found stains / discoloration. Missing item (tubing) found on instrument. | Mixer found stains / discoloration. Missing item (tubing) found on instrument. | Inline filter not connected to holder. Pump Tray Sub Assy and Degasser Sub Assy - few Screws not tightened.
5 | Open the Top Cover for Inspection | Missing item (screw). Tubing not packed inside plastic bag. Dirt & dust found on instrument compartment. | Cosmetic defects such as stains & paint chip off. Dirt & dust found on instrument compartment. Screw assembled at wrong location. | Vial Sensorr Cable not connected. Cable tie wrap not clipped and wires loose. Screw assembled at wrong location. Green Connectors A & B not at correct position (switched) found on instrument. | Waste tubing cut off. Cable tie wrap not clipped. Missing screw
| Stage / Activity | Unit 1 | Unit 2 | Unit 3 | Unit 4
| Serial Number | Unit 1) SN L16SM4371A: | Unit 2) SN L16SM4368A: | Unit 3) SN L16SM4356A | Unit 4) SN C17SM4376A.
1 | Visual | Scratches on Chassis, wire connections removed, inner compartment dirty etc.
6a | Run Full Vial Test | Pass | Pass | FAIL (39 out of 120 position failed) | not done. Not required as Failure reported is not related to Carry over test failure
6b | Run Carry Over Test | Pass | Pass | Fail vial Sensor. Not able to proceed. | not done. Not required as Failure reported is not related to Carry over test failure
6c | Verify the Bracket bending angle with Go-NoGo Fixture | Pass | Pass | Fail - "GO" | not done. Not required as Failure reported is not related to Carry over test failure
7 | Rework the Bracket of the Chassis and check with Fixture | NA | NA | Pass | not done. Not required as Failure reported is not related to Carry over test failure
8 | Run Final Test | Pass | Pass | pend | Instrument failed for 'Staric Leak'. Found a brown flexi tubing not assembled by Flex. After replaced to a new tubing, still failing 'Static Leak'. Per observation (item 4), the screws securing Pump Tray sub assy were loose, indicating it might have been removed before by end customer. Further observation found that the tubes were bent from the original way. The Pump Tray was removed to be tested in the Pump tray Sub Assy testing. Result - Failed Leak. After replaced 2 x Tubings and 1x Check Valve, the Pump Tray Sub Assy testing passed and re-assembled back to the Instrument. Re-Run Final Test - PASSED
9 | Run Full Vial Test and Carry Over Test | Pass | Pass | pend | Passed
10 | SQT | Completed. | Completed. | pend | Pend
11 | Update on 08 Sep 2018
12 | change to new seal pak? | Yes - change to new Sealpak | NO - still with the original Sealpak | Yes - change to new Sealpak | Yes - change to new Sealpak
13 | Final Test and 120 vial position test result? | Final Test - pass 120 vial postion test - pass | NA | Final Test - pass 120 vial postion test - pass | Final Test - pass 120 vial postion test - pass
14 | where is the instrument? | At Flex, standby for Drop Test evaluation. | At Flex | At WPPL | At WPPL
15 | Remarks? | Closed | SQT report reviewed by WPPL. See attached | Closed | Closed