Acquity BSM Tester Functional Requirements

Current Rear Panel Test Functionality
1. Download Firmware
2. Set Serial Number

3. Test Rear Panel (I/O Events, Ethernet, RS232, Analog Outputs)
Current Fluidic Test Functionality
4. Prime UUT
5. Verify Inlet (SSV) Plumbing

6. Test Degasser
7. Test Pump A Leakrate (Current Test = Ramp & Decay)
8. Test Pump B Leakrate (Current Test = Ramp & Decay)
9. Test Flowrate → Measure Flow Through 3 Restrictors

10. Prime With IPA (May Not Be Necessary, With ACV)
11. Prime with Air

New Functionality
12. Test System Leakrate (Currently a Subjective, Visual Test)

13. Empty Solvent From Balance
Optional New Functionality
14. Prime Sealwash

15. Test Sealwash Flow (Currently Checked, Manually)
16. Test Transducer (May not be necessary; inaccurate transducer could only pass with a double-fault → transducer fault combined with the wrong solvent and/or restrictor)

Other?
