# Validation MASTER Plan

Waters Pacific Pte Ltd (WPPL)

REVISION HISTORY

| Version | Author | Date | Summary of change |
| --- | --- | --- | --- |
| 00 | Ash Jeffery & Wan Jing Loh | 09/01/2018 | New document created |

TABLE OF CONTENTS

1	INTRODUCTION	3

2	Scope	3

2.1	Review period	4

3	RESPONSIBILITIES	4

4	QUALITY RISK MANAGEMENT	5

4.1	Initial Assessment	5

5	VALIDATION STRATEGY	6

5.1	Identification of a CTQ	6

5.2	Equipment Qualification	6

5.3	Re-validation/Re-Qualification	7

5.4	Change Management	7

5.5	Calibration	7

5.6	Method Validation	7

6	Validation Categories	9

6.1.1	Prospective	9

6.1.2	Concurrent validation	9

6.1.3	Retrospective validation	9

7	DOCUMENTATION MANAGEMENT	10

7.1	Document Approvals	10

7.2	Document Storage and Retention	10

7.3	Supplier Sourced Validation Documentation	10

7.4	Supplier Based Documentation	10

8	GLOSSARY	11

9	REFERENCES	12

10	VALIDATION PROJECTS	12

# INTRODUCTION

This Validation Master Plan (VMP) details how validation is carried out within WPPL (Waters Pacific Pte Ltd). This document also outlines how calibration and change management is used to ensure the validated state of equipment, methods and processes used by WPPL or within the WPPL supply chain.

This document has been written in accordance with document 730000837 – Waters Global Validation Policy.

All validation work for WPPL is supported by the Waters Wilmslow site, which is responsible for validation documentation generation. WPPL validation documentation will be managed according to the Waters Wilmslow Site Validation Master Plan – 47082.

# Scope

The scope of the WPPL VMP includes the following:

- Equipment used in the in-process and final testing of Instruments (Products)

- Equipment used in the testing of spare parts or components

- Test Methods used within the Chemical Prep lab

- Equipment/methods/processes in use at suppliers managed by WPPL

- Use of calibration management

- Use of change management

The following is not within the scope of the WPPL VMP:

- Computer system software, which is covered in 730001135 – Global Software Validation Procedure.

- Design and Development validation which is covered under the WPDP process – 7300000985.

## Review period

At a minimum, this document shall be reviewed 12-monthly.

The document shall be revised at each review by updating the revision number and history. The revision history shall be updated to state ‘Periodic Review Complete’ and all applicable changes detailed.

# Terms & definitions

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| CTQ | Critical to Quality |
| MS | Mass Spectrometer |
| PAE | Product Assurance Engineering |
| QA | Quality Assurance |
| QE | Quality Engineer |
| VE | Validation Engineer |
| VMP | Validation Master Plan |

# RESPONSIBILITIES

| Role | Responsibilities |
| --- | --- |
| Wilmslow Quality | Maintenance of site VMP Approval of validation documentation Produce plans, protocols, reports for validation activities, where required Maintenance of validation documentation Provide support and guidance on validation requirements |
| WPPL Quality | Maintenance of site VMP Notify Wilmslow Quality of: - All new equipment - Changes to validated or calibrated items - Unplanned maintenance to validated or calibrated items |
| Wilmslow PAE | Produce plans, protocols, reports for validation activities, where required Determination of classification, validation status, priority and Timelines/resources Performing validation activities Provide support and guidance on validation activities |
| WPPL Operations | Reviewing validation documentation Performing validation activities Identifying and notifying WPPL Quality of: - All new equipment - Changes to validated or calibrated items - Unplanned maintenance to validated or calibrated items |

# QUALITY RISK MANAGEMENT

## Initial Assessment

An assessment is performed for all new equipment entering the manufacturing process within the WPPL site or supply chain. The assessment is performed by Waters Wilmslow in line with the Wilmslow Site VMP - 47082. WPPL is responsible for notifying Waters Wilmslow of any new equipment to be installed on site or in the WPPL supply chain.

# VALIDATION STRATEGY

## Identification of a CTQ

Waters Wilmslow is responsible for identifying critical to quality (CTQ) attributes prior to carrying out qualification or validation work, this shall be defined within the Qualification or validation plan.

## Equipment Qualification

As defined within the Global Validation Policy - 730000837, equipment used in the manufacture, test and release of a product requires a level of assessment. The purpose of EQ is to ensure equipment is installed correctly, operates to manufacturer’s guidelines and performs against Waters’ requirements through predetermined acceptance criteria and is typically performed prior to any equipment being used to manufacture, test or release products. See Appendix 1.

## Re-validation/Re-Qualification

Revalidation or requalification may be performed on the process/method/equipment under certain conditions.  Examples include but are not restricted to:

- Change(s) in process/method/equipment have occurred that could affect product quality

- Transfer/moving of the process/method/equipment

- Observed negative trends (e.g. CAPAs or NCRs) related to process/method/equipment

- Change(s) in product design which affects the process/method/equipment

- Periodic requalification requirements defined within a report

## Change Management

Change management is required to control changes to equipment, method or process to ensure the change has not affected performance and that it remains fit for use.

The WPPL Change Management 730005503 shall be followed.

## Calibration

Calibration is a function to ensure equipment remains fit for use by periodic check to ensure it remains within operational tolerances. Equipment that has been assessed using the EQAF form – 7300003704 may need a level of calibration if the outcome was verification or validation.

The calibration process also includes routine planned maintenance. This is managed by WPPL-SP-15.

For all unplanned maintenance, the WPPL Change Management 730005503   shall be followed (Section 6.4)

## Method Validation

As defined within the Global Validation Policy, methods that are used for testing products are deemed high risk as they are used in both in-process and final release testing. All analytical methods used to test solutions are initially developed and validated by Waters Wilmslow and transferred to WPPL.

To ensure the method is accurate a minimum of 6 batches of solution are manufactured and tested at Waters Wilmslow and then a minimum of 3 batches are then manufactured at WPPL. .

Once methods have been suitably developed, validated and transferred, the manufacturing instruction is finalised and routine production can begin.

The overall process is managed and maintained by Waters Wilmslow.

# DOCUMENTATION MANAGEMENT

## Document Approvals

Electronic document approval should be performed following the SAP Control of Documents Procedure – 730003080.

Due to protocols requiring a two stage review process, the initial review and approval will be performed electronically in SAP. The second review step is performed once the protocol has been completed and the use of hand written signatures will be used to review and post approve the protocol. Upon post –approval completion, the QA representative will scan the completed document and associated attachments, in .PDF format, to the SAP database as an attachment within the document record.

## Document Storage and Retention

Wilmslow Quality is responsible for maintaining all electronic versions of the validation documents. If required, WPPL Quality shall request documents from Waters Wilmslow Quality.

# REFERENCES

This document uses the following references:

| Document Number | Document title |
| --- | --- |
| 730000837 | Waters Global Validation policy |
| 730003080 | SAP Control of Documents Procedure |
| 48044 | Production Sample Assessment Form |
| 730005503 | WPPL Change Management |
| 47940 | Control of Change Form |
| WPPL-SP-15 | Control & Monitoring of IMTE |

APPENDIX 1 – EQUIPMENT QUALIFICATION

| Serial Number | Equipment ID | Description | Location | Qualification Report Number |
| --- | --- | --- | --- | --- |
| KAA0270 | WPPL0039 | Acquity QDa Detector Instrument | WPPL | VAL001132 |
| KAD2550 | WPPL0012 | Acquity QDa Detector Instrument |  |  |
| N/A | WPPL0017 | Anachem Technologies Fume Cupboard |  |  |
| N/A | TBC | Milli-Q Advantage A10 UP Water System |  |  |
| 840303798 | WPPL0005 | Liebher LKUexv1610 Fridge |  |  |
| 840442763 | WPPL0004 | LGUex1500 Freezer |  |  |
| 74400208 | WPPL0025 | Miele Glassware Washer |  |  |
| 35701494 | WPPL0003 | Sartorius Balance |  |  |
| G42127G | WPPL0001 | Eppendorf P1000 Pipette |  |  |
| G42285G | WPPL0002 |  |  |  |
| I22994G | WPPL0036 | Eppendorf P200 Pipette |  |  |
| I22911G | WPPL0037 |  |  |  |
| TB685D | TBC | RGA Rig TB685D |  | VAL001171 |
| S4BA7494 | TBC | QDa Test Floor Test Bay | GES | VAL001193 |
| S4BK7626 | TBC | QDa Test Floor Test Bay |  | VAL001193 |
| S4BK9317 | TBC | QDa Test Floor Test Bay |  | VAL001193 |
| S4BA7494 | TBC | QDa Test Floor Test Bay |  | VAL001193 |
| MM7251 | TBC | AMITS Storage System (PC) |  | VAL001193 |
| 345-0747 | W1-17GC0122 | Pat & Flash Tester |  | N/A |
| 90001301209 | W1-17GC0095 | Helium Leak Detector |  | N/A |
| 100180047 | TBC | Steam Cleaner | Richport | 50803 |
| N/A | TBC | DI System |  | 51165 |
| N/A | TBC | Ultrasonic Cleaning Equipment |  | 48446 |

APPENDIX 2 – METHOD QUALIFICATION

| Reference Number | Description | Location | Qualification Report Number |
| --- | --- | --- | --- |
| 730005264 | Acquity QDa Setup Solution | WPPL | VAL001085 |