| IQ Test name: Associated Documentation Verification |  |  | Test Script No.: IQ-01 |  |  |
| --- | --- | --- | --- | --- | --- |
| Equipment No./ID: TE82988-12 |  | Equipment Name: Global Test System |  |  |  |
| Installation Site: Venture: WPPL, Singapore |  |  |  |  |  |
| Purpose: To verify that supporting documentation is available prior to the completion of qualification. Test execution: List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header and then complete the table below. If applicable documents are unavailable state why in the comments section. All applicable documents must be in either Ready, Pre-Rel-Prod, or Release General state in SAP. Enter N/A for any parameter or cell that is not used or not applicable. Objective evidence: Completed Test Script Data Sheet SAP Screen shot of below available document(s) |  |  |  |  |  |
| Document Type | Title; Doc. No.; Revision/Version No. |  |  | Status in SAP | Initial & Date |
| TSP | 50593 Ver. 03 |  |  |  |  |
| CAL | 735004006 Rev A |  |  |  |  |
| Comments: |  |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |  |

Copy this Test Script as needed

| IQ Test name: Site Verification |  |  |  | Test Script No.: IQ-02 |  |
| --- | --- | --- | --- | --- | --- |
| Equipment No./ID: TE82988-12 |  |  | Equipment Name: Global Test System |  |  |
| Installation Site: Venture: WPPL, Singapore |  |  |  |  |  |
| Purpose: To verify the location in which the test equipment is to be installed is suitable for use. Test execution: List any additional applicable environmental parameters in the Parameters column. Using the applicable equipment documentation, define the specifications in the Acceptance Criteria column, as appropriate. Verify and/or measure the applicable parameters and complete the table below. Do not delete any of the listed defined parameters. Enter N/A for any parameter or cell that is not used or applicable. Objective evidence: Completed Test Script Data Sheet. Photo(s) of the work area, if applicable. Copies of calibration certificates for gauges or meters used for measurement, if applicable. |  |  |  |  |  |
| Parameters | Acceptance criteria | As found |  | Pass/Fail | Initial & Date |
| Ambient Temperature | 59-86º F (15-30º C) |  |  |  |  |
| Relative Humidity | 20%-80%, Non-Condensing |  |  |  |  |
| Workspace | Adequate space for testing and service |  |  |  |  |
| Test Station Bench | Large and strong enough to support all equipment |  |  |  |  |
| Electrical Supply | 50-60 HZ, 85-132 or 180-350 VAC |  |  |  |  |
| Electrical Outlets | >= 1 |  |  |  |  |
| Network Available | >= 1 |  |  |  |  |
| Vacuum | 40 cfm (1.13 cubic meters/min) minimum vacuum flow and a minimum vacuum pressure of 23 inches Hg (78 kPa). |  |  |  |  |
| Comments: |  |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |  |

Copy this Test Script as needed

| IQ Test name: Instrumentation Calibration Verification |  |  |  |  |  | Test Script No.: IQ-03 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment No./ID: TE82988-12 |  |  | Equipment Name: Global Test System |  |  |  |  |
| Installation Site: Venture: WPPL, Singapore |  |  |  |  |  |  |  |
| Purpose: To verify that all critical instruments are in the calibration system and calibrated or a preventive maintenance has been performed, as appropriate. Test execution: List all critical instruments/components and/or sub-systems (System, TUV, Sample Manager, BSM, Column Manager, Test Solutions, Meters, Gauges, Scales, etc.) in the Description column and complete the table below. Verify that all critical instruments have had a preventive maintenance and/or calibration performed, in accordance with the PM and/or Calibration procedures, and have been included in the preventive maintenance/calibration schedule. For test solutions note the expiration date. Enter N/A for any parameter or cell that is not used or applicable. Objective evidence: Completed Test Script Data Sheet Copies of calibration and/or maintenance certificates. |  |  |  |  |  |  |  |
| Description | Waters P/N and Asset No. | Serial Number |  | Last Cal/PM Performed Date | PM Cal/Due / Expiration Date |  | Initial & Date |
| AGILENT 34411A | ID#97025 |  |  |  |  |  |  |
| AGILENT 33220A | ID#97026 |  |  |  |  |  |  |
| TEKTRONIX TDS 3054C | ID#97024 |  |  |  |  |  |  |
| GENH30-25 | ID#97019 |  |  |  |  |  |  |
| GENH30-25 | ID#97020 |  |  |  |  |  |  |
| GENH60-12.5 | ID#97015 |  |  |  |  |  |  |
| GENH60-12.5 | ID#97016 |  |  |  |  |  |  |
| GENH60-12.5 | ID#97017 |  |  |  |  |  |  |
| GENH60-12.5 | ID#97018 |  |  |  |  |  |  |
| CHROMA 63103A | ID#97010 |  |  |  |  |  |  |
| CHROMA 63103A | ID#97011 |  |  |  |  |  |  |
| CHROMA 63103A | ID#97012 |  |  |  |  |  |  |
| CHROMA 63103A | ID#97013 |  |  |  |  |  |  |
| National Instrument/ NI PXI-6229 | ID#97021 |  |  |  |  |  |  |
| National Instrument/ NI PXI-6229 | ID#97022 |  |  |  |  |  |  |
| National Instrument/ NI PXI-4130 | ID#97023 |  |  |  |  |  |  |
| Marvin Test Solution / GX1164 | ID#97014 |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |  |  |  |

Copy this Test Script as needed

| IQ Test name: Test Software and Instrument Firmware Version |  |  |  | Test Script No.: IQ-04 |  |
| --- | --- | --- | --- | --- | --- |
| Equipment No./ID: |  |  | Equipment Name: |  |  |
| Installation Site: |  |  |  |  |  |
| Purpose: To verify that the installed firmware/software is correct per the applicable design/operational documents (TSP, ASP, MPR, BOM, and/or other Design documents). Test execution: List all critical instruments/components and/or sub-systems that require software and/or firmware in the Description column. List the associated required software/firmware version, as defined in the applicable design/operational documents, in the Acceptance criteria column. Verify that the software/firmware version on the Test Equipment match the acceptance criteria and are identical. Enter N/A for any parameter or cell that is not used or applicable. Objective evidence: Completed Test Script Data Sheet Identify the applicable design/operational document that defined the required software/firmware version(s) in the associated Installation Qualification Report. |  |  |  |  |  |
| Description | Acceptance criteria | As found |  | Pass/Fail | Initial & Date |
| 210001115HVC | Version: C | Version: |  |  |  |
| 210001115 | Version: C | Version: |  |  |  |
| NI TestStand | Version: 4.2.1 | Version: |  |  |  |
|  | Version: | Version: |  |  |  |
|  | Version: | Version: |  |  |  |
|  | Version: | Version: |  |  |  |
|  | Version: | Version: |  |  |  |
|  | Version: | Version: |  |  |  |
| Comments: |  |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |  |

| IQ Test name: General Safety and Alarms |  |  | Test Script No.: IQ-05 |  |
| --- | --- | --- | --- | --- |
| Equipment No./ID: TE82988-12 |  | Equipment Name: Global Test System |  |  |
| Installation Site: Venture: WPPL, Singapore |  |  |  |  |
| Purpose: To verify that the test equipment can operate in a safe manner, within its designated location. Test execution: List any additional safety related conditions or alarms in the Safety Inspection column. List the associated safety requirement/specification in the Acceptance Criteria column. Perform an inspection/test of the equipment to assess its safety related conditions. Enter N/A for any parameter or cell that is not used or applicable. Objective evidence: Completed Test Script Data Sheet Photo of the secured glass bottles and/or tanks, if applicable |  |  |  |  |
| Safety Inspection | Acceptance criteria |  | Pass/Fail | Initial & Date |
| Mechanical Safety/ Pinch Points/Sharp Edges | None Found |  |  |  |
| Electrical Safety Hi Pot/Ground Continuity | Passed per test specification WAT046538 |  |  |  |
| Comments: |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |

Copy this Test Script as needed

| IQ Test name: Functional Verification |  |  | Test Script No.: IQ-06 |  |
| --- | --- | --- | --- | --- |
| Equipment No./ID: TE82988-12 |  | Equipment Name: Global Test System |  |  |
| Installation Site: Venture: WPPL, Singapore |  |  |  |  |
| Purpose: To verify that system transfer and/or qualification activities have not adversely affected system functionality. Test execution: One UUT shall be tested to verify that the test equipment functions as expected. If the system creates and/or tests more than one part number, the UUT to be tested per this Script should be a “worst case” product, and defined within the associated Installation Qualification Report. Enter N/A for any parameter or cell that is not used or applicable. Objective evidence: Completed Test Script Data Sheet Copy of the system Test Record for the UUT, if applicable. |  |  |  |  |
| UUT Part Number/Serial Number | Acceptance criteria |  | Pass/Fail | Initial & Date |
| GTS Self-Test Fixture Part Number: TE83004 | All required CTQ Specification measurements result in a Pass. |  |  |  |
| Test Fixture Part Number for Water UK PCBA: 279002953 UUT Part Number: 210001115 UUT Serial Number: | All required CTQ Specification measurements result in a Pass. |  |  |  |
| Comments: |  |  |  |  |
| Overall Test Conclusion: Pass Fail |  |  |  |  |

Copy this Test Script as needed