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IQ Test name: Associated Documentation Verification
	Test Script No.: IQ-01

	Equipment No./ID: TE82988-12
	Equipment Name: Global Test System 

	Installation Site: Venture: WPPL, Singapore

	Purpose:
To verify that supporting documentation is available prior to the completion of qualification.
Test execution:
List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header and then complete the table below. If applicable documents are unavailable state why in the comments section. All applicable documents must be in either Ready, Pre-Rel-Prod, or Release General state in SAP. Enter N/A for any parameter or cell that is not used or not applicable.
Objective evidence:
Completed Test Script Data Sheet
SAP Screen shot of below available document(s) 

	Document Type
	Title; Doc. No.; Revision/Version No.
	Status in SAP
	Initial & Date

	TSP
	50593 Ver. 03
	
	

	CAL
	735004006 Rev A
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	Comments:
[bookmark: _GoBack]






	Overall Test Conclusion:	 Pass			 Fail	


Copy this Test Script as needed
	
IQ Test name: Site Verification
	Test Script No.: IQ-02

	Equipment No./ID: TE82988-12
	Equipment Name: Global Test System

	Installation Site: Venture: WPPL, Singapore

	Purpose:
To verify the location in which the test equipment is to be installed is suitable for use.
Test execution:
List any additional applicable environmental parameters in the Parameters column. Using the applicable equipment documentation, define the specifications in the Acceptance Criteria column, as appropriate. Verify and/or measure the applicable parameters and complete the table below. Do not delete any of the listed defined parameters. Enter N/A for any parameter or cell that is not used or applicable. 
Objective evidence:
Completed Test Script Data Sheet.
Photo(s) of the work area, if applicable. 
Copies of calibration certificates for gauges or meters used for measurement, if applicable.

	Parameters
	Acceptance criteria
	As found
	Pass/Fail
	Initial & Date

	Ambient Temperature
	59-86º F (15-30º C)

	
	
	

	Relative Humidity
	20%-80%, Non-Condensing
	
	
	

	Workspace
	Adequate space for testing and service
	
	
	

	Test Station Bench
	Large and strong enough to support all equipment
	
	
	

	Electrical Supply
	50-60 HZ,
85-132 or 180-350 VAC
	
	
	

	Electrical Outlets
	>= 1
	
	
	

	Network Available
	>= 1
	
	
	

	Vacuum
	40 cfm (1.13 cubic meters/min) minimum vacuum flow and a minimum vacuum pressure of 23 inches Hg (78 kPa).
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	Comments:



	Overall Test Conclusion:	 Pass			 Fail	


Copy this Test Script as needed
	[bookmark: _Toc266368353]

IQ Test name: Instrumentation Calibration Verification
	Test Script No.: IQ-03

	Equipment No./ID: TE82988-12
	Equipment Name: Global Test System

	Installation Site: Venture: WPPL, Singapore

	Purpose:
To verify that all critical instruments are in the calibration system and calibrated or a preventive maintenance has been performed, as appropriate.
Test execution:
List all critical instruments/components and/or sub-systems (System, TUV, Sample Manager, BSM, Column Manager, Test Solutions, Meters, Gauges, Scales, etc.) in the Description column and complete the table below. Verify that all critical instruments have had a preventive maintenance and/or calibration performed, in accordance with the PM and/or Calibration procedures, and have been included in the preventive maintenance/calibration schedule. For test solutions note the expiration date. Enter N/A for any parameter or cell that is not used or applicable.
Objective evidence:
Completed Test Script Data Sheet
Copies of calibration and/or maintenance certificates.

	Description
	Waters P/N and Asset No.
	Serial Number
	Last Cal/PM 
Performed Date
	PM Cal/Due / Expiration Date
	Initial & Date

	AGILENT 34411A
	ID#97025
	
	
	
	

	AGILENT 33220A
	ID#97026
	
	
	
	

	TEKTRONIX TDS 3054C
	ID#97024

	
	
	
	

	GENH30-25
	ID#97019
	
	
	
	

	GENH30-25
	ID#97020
	
	
	
	

	GENH60-12.5
	ID#97015
	
	
	
	

	GENH60-12.5
	ID#97016
	
	
	
	

	GENH60-12.5
	ID#97017
	
	
	
	

	GENH60-12.5
	ID#97018
	
	
	
	

	CHROMA 63103A
	ID#97010
	
	
	
	

	CHROMA 63103A
	ID#97011
	
	
	
	

	CHROMA 63103A
	ID#97012
	
	
	
	

	CHROMA 63103A
	ID#97013
	
	
	
	

	National Instrument/ NI PXI-6229
	ID#97021
	
	
	
	

	National Instrument/ NI PXI-6229
	ID#97022
	
	
	
	

	National Instrument/ NI PXI-4130
	ID#97023
	
	
	
	

	Marvin Test Solution / GX1164
	ID#97014
	
	
	
	

	Comments:






	Overall Test Conclusion:	 Pass			 Fail	


Copy this Test Script as needed

	[bookmark: _Toc266368355]IQ Test name: Test Software and Instrument Firmware Version
	Test Script No.: IQ-04

	Equipment No./ID:
	Equipment Name: 

	Installation Site: 

	Purpose:
To verify that the installed firmware/software is correct per the applicable design/operational documents (TSP, ASP, MPR, BOM, and/or other Design documents).
Test execution:
List all critical instruments/components and/or sub-systems that require software and/or firmware in the Description column. List the associated required software/firmware version, as defined in the applicable design/operational documents, in the Acceptance criteria column. Verify that the software/firmware version on the Test Equipment match the acceptance criteria and are identical. Enter N/A for any parameter or cell that is not used or applicable.
Objective evidence:
Completed Test Script Data Sheet
Identify the applicable design/operational document that defined the required software/firmware version(s) in the associated Installation Qualification Report.

	Description
	Acceptance criteria
	As found
	Pass/Fail
	Initial & Date

	210001115HVC
	Version: C
	Version: 
	
	

	210001115
	Version: C
	Version:
	
	

	NI TestStand
	Version: 4.2.1
	Version:
	
	

	
	Version: 
	Version:
	
	

	
	Version:
	Version:
	
	

	
	Version:
	Version:
	
	

	
	Version:
	Version:
	
	

	
	Version:
	Version:
	
	

	Comments:






	Overall Test Conclusion:	 Pass			 Fail	




	


IQ Test name: General Safety and Alarms
	Test Script No.: IQ-05

	Equipment No./ID: TE82988-12
	Equipment Name: Global Test System

	Installation Site: Venture: WPPL, Singapore

	Purpose:
To verify that the test equipment can operate in a safe manner, within its designated location.
Test execution:
List any additional safety related conditions or alarms in the Safety Inspection column. List the associated safety requirement/specification in the Acceptance Criteria column. Perform an inspection/test of the equipment to assess its safety related conditions. Enter N/A for any parameter or cell that is not used or applicable.
Objective evidence:
Completed Test Script Data Sheet 
Photo of the secured glass bottles and/or tanks, if applicable

	Safety Inspection
	Acceptance criteria
	Pass/Fail
	Initial & Date

	Mechanical Safety/
Pinch Points/Sharp Edges
	None Found
	
	

	Electrical Safety 
Hi Pot/Ground Continuity
	Passed per test specification WAT046538
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	Comments:  







	Overall Test Conclusion:	 Pass			 Fail	


Copy this Test Script as needed
	


IQ Test name: Functional Verification
	Test Script No.: IQ-06

	Equipment No./ID: TE82988-12
	Equipment Name: Global Test System

	Installation Site: Venture: WPPL, Singapore

	Purpose:
To verify that system transfer and/or qualification activities have not adversely affected system functionality.
Test execution:
One UUT shall be tested to verify that the test equipment functions as expected. If the system creates and/or tests more than one part number, the UUT to be tested per this Script should be a “worst case” product, and defined within the associated Installation Qualification Report. Enter N/A for any parameter or cell that is not used or applicable.
Objective evidence:
Completed Test Script Data Sheet 
Copy of the system Test Record for the UUT, if applicable.

	UUT Part Number/Serial Number
	Acceptance criteria
	Pass/Fail
	Initial & Date

	GTS Self-Test Fixture Part Number:  TE83004

	All required CTQ Specification measurements result in a Pass.
	
	

	Test Fixture Part Number for Water UK PCBA:  279002953

UUT Part Number: 210001115
UUT Serial Number:
	All required CTQ Specification measurements result in a Pass.
	
	

	Comments:  




	

	Overall Test Conclusion:	 Pass			 Fail	


Copy this Test Script as needed
Waters Confidential		                                 Template 730000483 V06
It is the responsibility of the user to verify they are using the current released version.
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