Waters - Wilmslow

Equipment Qualification Report

Venture ACQUITY QDa Test Floor

For Internal Use Only

Document reviewed and approved electronically

Author:  Edwina Beanland

Date: 05 March 2018

# Purpose

The purpose of this report is to describe the work that was carried out, the results obtained and the recommendations for the ACQUITY QDa Test Bays which underwent equipment qualification activities to ensure fitness for purpose at Venture, Singapore.

The validation report outlines each of the tests performed on the ACQUITY QDa Test Bays and all related outcomes. This report will draw on the conclusion that the equipment is fit for its intended purpose, or requires further testing or work carried out.

The report will also provide recommendations for maintaining the validated state of the system.

# Scope

This report is restricted to the Equipment Qualification aspect of four of the AQUITY QDa Test Bays located in Venture, Singapore.  The test bays covered are:

Waters S4BA7494

Waters S4BK7626

Waters S4BK9317

Waters S4BA7494

Test Bays were identified by their Host PC names

The AMITS Storage System PC was also qualified:

MM7251

Each test bay was verified using a Goldbrick QDa Instrument:

Model/Type:  MKI

Serial Number:  KAA0270

Asset No. UK05 100 2274

# EQ Activities

Equipment Qualification was conducted by testing the equipment against the User Requirements Specifications, Design Specification or manufacturers’ operational specification.  A Qualification Test Matrix was created within the EQ plan which detailed the testing requirements during execution of the protocol.

| EQ No. | Details | PASS / Pass (with comment) / FAIL |
| --- | --- | --- |
| Test Bay S4BA7494 |  |  |
| 1 | Equipment Details | Pass |
| 2 | Training Requirements | Pass |
| 3 | Documentation Review | Pass |
| 4 | Calibration Check | Pass |
| 5 | Health and Safety Testing | Pass |
| 6 | Installation checks of 13amp Mains Wall Sockets | Pass |
| 7 | Installation Checks of Nitrogen Supply | Pass |
| 8 | Installation Checks of Rotary Pump Exhaust System | Pass |
| 9 | Installation Checks of the Source Exhaust System Comments: Only had access to a single QDa, so unable to test system with 4 x 1200 l/hr entering the source exhaust system. | Pass with comments |
| 10 | Installation Checks of the Test Bay Computer Hardware | Pass |
| 11 | Installation Checks of the Test Bay Computer Software | Pass |
| 12 | Installation Checks of the Goldbrick Acquity QDa | Pass |
| 13 | Operational Checks of 515 Pump | Pass with Deviation (no.1) |
| 14 | Operational Checks of Test Bay Software and Network | Pass |
| 15 | Operational Checks of the Sample Divert Valve | Pass with Deviation (no.2) |
| 16 | Operational Checks of the Test Bay System, including the Goldbrick Acquity QDa | Pass with deviation (no.2) |
| 17 | Shutdown of Test System | Pass with Deviation (no.2 and 3) |
| Test Bay S4BK7626 |  |  |
| 1 | Equipment Details | Pass |
| 2 | Training Requirements | Pass |
| 3 | Documentation Review | Pass |
| 4 | Equipment Calibration | Pass |
| 5 | Health and Safety Testing | Pass |
| 6 | Installation Checks of 13Amp mains wall sockets | Pass |
| 7 | Installation check of Nitrogen Supply | Pass |
| 8 | Installation Checks of Rotary Pump Exhaust System | Pass |
| 9 | Installation Checks of the Source Exhaust System Comments: Only had access to a single QDa so unable to test system with 4 x 1200l/hr entering the source exhaust system. | Pass with comments |
| 10 | Installation Checks of the Test Bay Computer Hardware | Pass |
| 11 | Installation Checks of the Test Bay Computer Software | Pass |
| 12 | Installation Checks of the Gold Brick Acquity QDa | Pass |
| 13 | Operational Checks of 515 Pump | Pass with deviation (No.1) |
| 14 | Operational Checks of Test By Software and Network | Pass |
| 15 | Operational Checks of the Divert Valve | Pass with deviation (No.2) |
| 16 | Operational Checks of the Test Bay System including the Gold Brick Acquity QDa | Pass with deviation (No.2) |
| 17 | Shutdown of Test System | Pass with deviation (No.2 and No.3) |
| Test Bay S4BK9317 |  |  |
| 1 | Equipment Details | Pass |
| 2 | Training Requirements | Pass |
| 3 | Documentation Review | Pass |
| 4 | Equipment Calibration | Pass |
| 5 | Health and Safety Testing | Pass |
| 6 | Installation Checks of 13Amp mains wall sockets | Pass |
| 7 | Installation check of Nitrogen Supply | Pass |
| 8 | Installation Checks of Rotary Pump Exhaust System | Pass |
| 9 | Installation Checks of the Source Exhaust System Comment: Only had access to a single QDa so unable to test system with 4 x 1200l/hr entering the source exhaust system. | Pass with Comment |
| 10 | Installation Checks of the Test Bay Computer Hardware | Pass |
| 11 | Installation Checks of the Test Bay Computer Software | Pass |
| 12 | Installation Checks of the Gold Brick Acquity QDa | Pass |
| 13 | Operational Checks of 515 Pump | Pass with Deviation (No.1) |
| 14 | Operational Checks of Test By Software and Network | Pass |
| 15 | Operational Checks of the Divert Valve | Pass with Deviation (No.2) |
| 16 | Operational Checks of the Test Bay System including the Gold Brick Acquity QDa | Pass with Deviation (No.2) |
| 17 | Shutdown of Test System | Pass with Deviation (No.2 and No.3) |
| Test Bay S4BK7529 |  |  |
| 1 | Equipment Details | Pass |
| 2 | Training Requirements | Pass |
| 3 | Documentation Review | Pass |
| 4 | Equipment Calibration | Pass |
| 5 | Health and Safety Testing | Pass |
| 6 | Installation Checks of 13Amp mains wall sockets | Pass |
| 7 | Installation check of Nitrogen Supply | Pass |
| 8 | Installation Checks of Rotary Pump Exhaust System | Pass |
| 9 | Installation Checks of the Source Exhaust System Comment: Only had access to a single QDa so unable to test system with 4 x 1200l/hr entering the source exhaust system. | Pass with Comment |
| 10 | Installation Checks of the Test Bay Computer Hardware | Pass |
| 11 | Installation Checks of the Test Bay Computer Software | Pass |
| 12 | Installation Checks of the Gold Brick Acquity QDa | Pass |
| 13 | Operational Checks of 515 Pump | Pass with Deviation (No.1) |
| 14 | Operational Checks of Test By Software and Network | Pass |
| 15 | Operational Checks of the Divert Valve | Pass with Deviation (No.2) |
| 16 | Operational Checks of the Test Bay System including the Gold Brick Acquity QDa | Pass with Deviation (No.2) |
| 17 | Shutdown of Test System | Pass with Deviation (No.2 and No.3) |
| MM7521 – AMITS Storage System |  |  |
| 1 | Equipment Details | Pass |
| 2 | Training Requirements | Pass |
| 3 | Documentation Review | N/A |
| 4 | Equipment Calibration | N/A |
| 5 | Utility Installation | N/A |
| 6 | Health and Safety Testing | N/A |
| 7 | Computer Hardware | Pass |
| 8 | Computer Software | Pass |
| 9 | Operational Checks of Test Bay Software and Network Comment: Delay in obtaining results due to Host PC account not allowing AMITS to reset DHCP service. Fault in GES-Venture account has been fixed. | Pass with comment |

# Discussion of Results

During the execution of the equipment protocols for the 4 test bays a number of deviations and ‘pass with comments’ were raised.  All the deviations were closed and there are no further actions required.  The details of the deviations can be found in section 5 of this report.

# Deviations

The following deviations were raised during execution of the EQ protocol and include information and resolution:

| Deviation No. | Details | Open/Closed |
| --- | --- | --- |
| 1 | Deviation Details: 515 pump failed 5ml/min flow test. It did pass 0.5ml/min and 1.0ml flow tests. Corrective Actions: None. The flow rate used by the system test is 0.35ml/min Outcome: No change required. The failure will not affect the system. | Closed |
| 2 | Deviation Details: Default position of divert vale was “2” not “1”. This meant there were a lot of errors in the documentation and the divert valve had to be plumbed differently to perform as expected. Corrective Actions: Peek tube between 515 pump and Divert valve = port 1 Peek tube between Divert valve and sample Duran = port 2 Peek tube between Divert valve and QDa = port 6 Outcome: By performing the above change the divert valve performed correctly. | Closed |
| 3 | Deviation Details: GES-Venture account created for the host PC system did not have the correct rights to allow AMITS to function fully. The failure was caused by the User Account Control notification setting. When changed to ‘Never Notify’ the software worked correctly. The AMITS software before this point used accounts with full administrative rights, which allowed users to see the Waters network. Corrective Actions: UAC reset to ‘Never Notify’. Outcome: AMITS functioned correctly. | Closed |

# Maintaining The Validated State

In order to maintain the validated state, test bays  S4BA7494, S4BK7626, S4BK9317, S4BK7529 and AMITS Storage System MM7521 will undergo an annual planned maintenance following the procedure on doc no. 48550 with all details recorded on the planned maintenance test record document no. 48552.

# Conclusion

In conclusion the equipment as detailed in section 2 for GES Venture, Singapore is fit for its intentional use.