## Sheet1
Tender No.: HOMHQ0ETT16000310
Appendix 6 to Section III
(6) REQUIREMENT SPECIFICATION FOR LIQUID CHROMATOGRAPHY MASS
SPECTROMETRY/MASS SPECTROMETER (LCMS/MS)
(A) SPECIFICATIONS | Compliance to Specifications (Y/N) | Remarks
1.1 General Description
a. The system shall be an ultra high performance liquid chromatography coupled to a triple quadrupole mass spectrometer. | Y
b. The LCMS/MS shall be a bench-top version. | Y
c. The dimensions of the system shall not be more than 150 cm (L) by 70 cm (W) by 70 cm (H). | Y
d. The weight of the system shall not be more than 150 kg. | Y
e. The system shall have equivalent if not better capability than the specifications stated from 1.2 to 1.15. | Y
f. Each system shall include item 1.2 to 1.15. | Y
1.2 Ultra High Pressure Binary Solvent Delivery System
a. The pump shall be a high pressure binary gradient pump | Y
b. Each pump shall be installed with a solvent selector valve to automatically switch between two solvents. | Y
c. The flow rate shall be adjustable between 0.001 to 5 mL/min in 0.0001 mL/min increments. | N | flow rate 0.001 – 2ml/min
d. The flow rate precision shall be less than ±0.07% RSD. | Y
e. The composition accuracy shall be less than 0.1% (1 mL/min, 100MPa). | Y
f. The pump shall be able to withstand an operating pressure of at least 120 MPa for flow rate of up to 2 mL/min and 80 MPa for flow rate up to 5 mL/min | Y
g. The composition precision shall be less than 0.15 % RSD at constant room temperature. | Y
h. Wash kit shall be included. | Y
1.3 Degassing Unit
a. The degassing unit shall have at least four flow lines. | Y | 6 flow lines
b. The maximum operating flow rate shall be up to 10 mL/min per flow line. | N
c. The error status shall be transferred to the operating software. | Y
d. The unit shall include a leak sensor and a drain-pan. | Y
1.4 Auto Sampler
a. The auto sampler shall use a direct injection method as standard. However, the system shall allow flexibility to change the injection method to the loop injection with adjustable or fixed volume injection, when required. | Y
b. The auto sampler shall have a leak sensor, automatic rack and vial recognition. | Y
c. The injection system shall be of variable injection volume type with zero sample loss during injection. | Y
d. The auto sampler shall perform customised continuous analyses according to the conditions specified for each sample. Priority samples may be inserted in the queue during automated run. | Y
e. The system shall have the function of rinsing the flow line before and after sampling. | Y
f. The auto sampler shall have functions for maintenance and validation which are accessible by dedicated operating buttons. | Y
g. The auto sampler shall be able to accommodate at least 100 2 mL vials in 1 tray. | Y
h. The maximum operating pressure shall not be less than 80 MPa. | Y | 18,000 PSI
i. The sample injection volume shall be variable from 0.1 µL to 40 µL. | Y | 0.1 μL to 250.0 μL
j. The injection volume precision shall not be more than 0.25 % from 5 to 40 µL. | Y | <1% area RSD 0.2 to 1.9 μL injection <0.5% area RSD 2 to 10 μL injection
k. The injection volume accuracy shall be within +1% (at 10 µL injection, n = 10). | Y
l. The system shall allow temperatures to be set from 4 to 40 ºC. | Y
m. The carry over shall be below 0.004 % without rinsing. | Y
1.5 Column Oven
a. The column oven shall have a temperature limit device and temperature fuse. It must also have solvent and gas leak sensors as safety features. | Y
b. The column oven shall be capable of complex temperature programming in linear and step programs. The temperature range shall be ambient 10oC to 80oC. | Y
c. Complex temperature programming in linear and step programs shall be possible. | Y
1.6 Ion Source
a. The system must be equipped with the following: | Y
i. One set of Electro Spray Ionisation (ESI) assembly | Y
ii. One set of Atmospheric Pressure Chemical Ionisation (APCI) assembly | Y
b. The source temperature should be settable by user up to 750 oC. | N | 650◦C
c. A syringe and syringe pump shall be available for direct infusion of samples. | Y
d. The instrument must have an active source exhaust for the removal of gases within the ionization source. | N
1.7 Mass Analyzer
a. The mass analyzer shall consist of two, high standard, fast scanning quadrupoles and one efficient collision cell. The mass range is from at least 10 to 1000 m/z. | Y | 0-2048 amu
b. The instrument should be able to carry out positive and negative mode scans in one single experiment at a polarity switching time of 50 msec. | Y | Less than 15ms
c. The maximum scanning speed of the mass analyzer should be 20,000 amu/s. | Y
d. For sensitivity in positive ESI, the measured signal-to-noise ratio obtained from the MRM transition m/z 609 to 195 for an injection of 1 pg of Reserpine should be more than or equal to 90,000:1. For sensitivity in negative ESI, the measured signal-to-noise ratio obtained from the MRM transition m/z 321 to 152 for an injection of 1 pg of chloramphenicol should be more than or equal to 90,000:1. S/N measurements are calculated based on 1 standard deviation of at least 3 points of noise which produce the smallest standard deviation, after applying up to 3 Gaussian smooths. | Y | >500,000 :1 S/N
a. Mass spectrometry-mass spectrometry (MS-MS) capability shall be available in quadrupole 2 via collision induced dissociation. | Y
b. The collision cell shall have a minimum dwell time of 2 milliseconds. | Y | 1 millisecond
c. Auto-tuning shall be standard with a syringe pump that delivers accurate injection under full workstation control in both positive and negative modes. Manual tuning shall also be allowed. At least one set of standard sample for auto-tuning shall be supplied. | Y
d. The vacuum system shall be a differential type with triple inlet turbo-molecular pump and one rotary pump. | Y
e. There shall be a shutter to cut-off nitrogen when no mobile-phase flows or during stand-by mode. | Y
f. Stainless steel piping for carrier gas shall be provided. | Y
g. The detector shall capable of very fast switching between positive and negative ion detection. | Y
h. A unit of nitrogen generator shall be supplied for each LCMS/MS system. The nitrogen generator shall be compatible with the system. | Y
1.8 Gas Lines
a. Installation of gas lines if necessary (including regulator and semi-automatic switch over manifold), from the source of the gas to the system. All materials and labour shall be provided by tenderer | Y
b. Regulators shall be provided for the all gases required for analysis. | Y
1.9 Software
a. Software shall be the latest version and provide for a seamless control of all the modules of the LC and MS. Software upgrades, if necessary, shall be provided without any additional costs. | Y
b. MS/MS acquisition mode must have various modes apart from MRM, for instance full scan, selected ion monitoring (Q1 and Q3), product ion scan, precursor ion scan and neutral ion scan. | Y
c. Auto start-up and shutdown of the MS shall be possible using a single key operation. | Y
1.10 Computer and Printer
d. The computer shall have a Intel® Core™ i7-6700 Processor (Quad Core, 8MB, 8T, 3.4GHz) Intel® Integrated Graphics 8GB (1x8GB) 1600MHz DDR3L Memory (Max 4 DIMM slots; up 16GB)1TB 7200 RPM 3.5" SATA Hard Drive 8x DVD+/-RW, at least 22” LCD wide monitor, keyboard, , wireless mouse & mouse pad. | Y
e. The computer shall be pre-installed with the latest and compatible Windows operating system and Microsoft office software (Word, Excel® and PowerPoint) | Y
f. A colour laser printer with duplex printing function and a spare set of colour toner cartridges shall be provided. | Y
1.11 Uninterrupted Power Supply (UPS)
a. One set of UPS with output power supply shall be provided together with batteries that are sufficient to serve a backup time of at least 30 minutes. | Y
b. The package shall include the installation and commissioning of the UPS with the LCMS/MS supplied. | Y
1.12 Accessories
a. Each system provided shall come with the full suite of the following accessories, or their equivalent. The accessories are as follow: | Y
i. 2 C18 columns. The dimension of the column is 100 mm x 2.1 mm id x 1.7 um. | Y
ii. 2 guard columns compatible with the column in 1.11a i. | Y
iii. 2 guard column holders compatible with the guard column in 1.11a ii. | Y
iv. 1 piece of ESI probes assy. | Y
v. 1 piece of APCI probes assay | Y
vi. 2 piece of ESI electrode | Y
vii.1 piece of APCI electrode | Y
1.13 Others
a. Tenderer shall collect samples, provided by the Authority, to analyse at no additional costs. The chromatograms and analysis results shall be provided within 14 days upon receipt of the samples. All samples must be returned to the Authority after submission of report. The analysis results cannot be published in any form or revealed to a third party. | Y
b. All modules must be GLP/GMP compliant or a declaration of conformity certificate must be provided. | Y
c. The tenderer shall deliver the LCMS/MS to separate locations. Refer to para 4 of Section III for delivery schedule. Lab staff shall be notified of the exact delivery date at least 1 week before. | Y
d. The tenderer shall take note that there is no lift available at two of the sites of the delivery. The price quoted shall include hoisting charges of the instrument, if necessary. | Y
e. Prices quoted must include all administrative, delivery, installation and freight charges. | Y
1.14 Training
a. A briefing shall be provided on the operation of the system and software after installation of the system. | Y
b. Training shall be provided for at least 10 users for the operation, trouble shooting and general maintenance for each system. | Y
c. There shall be at least one training session from the supplier’s application specialist which can be customized to the users’ needs within a year of installation of the instrument. | Y
d. Training shall be provided for the users of the instruments at site, as well as at the tenderer’s application laboratory, after installation and commissioning. | Y
1.15 Warranty and Support
a. On-site installation, commissioning and training by factory-trained engineer is required. | Y
b. The tenderer shall have a local dedicated team, consisting of sales, engineers and application engineers from direct factory/manufacturer for optimum support of the system. | Y
c. The tenderer shall provide after-sales service efficiently and effectively. The tenderer shall have in his/her facility a similar system to that proposed in this tender for training purposes. | Y
d. Back up equipment support for testing must be available at the tenderer’s facility during equipment breakdown. | Y
e. Comprehensive support shall be provided for the equipment (inclusive of LCMS/MS, nitrogen generator and UPS) for a period of 12 months which will consist of the following at no extra cost: | Y
i. Unlimited service calls | Y
ii. Travel and labour expenses of the engineer | Y
iii. Service parts used for repairs | Y
f. The warranty shall commence only upon successful completion of the acceptance test or commissioning. The service engineer shall respond within 24 hours for service calls and within 48 hours for on-site service. | Y
(B) COMPREHENSIVE MAINTENANCE FOR LCMS/MS SYSTEM
2.1 The Contractor shall provide a seven (7) year Preventive Maintenance and Corrective Maintenance after the expiry of the warranty period. | Y
2.2 The scope of activities for the Comprehensive Maintenance package shall include the following: | Y
a. Preventive Maintenance (PM)
i. The Contractor shall perform PM twice per year on the LCMS/MS | Y
ii. The Contractor shall perform PM once per year for the nitrogen generator and UPS. | Y
iii. The scope of PM is as specified in Appendix 6.1 to Section III. | Y
b. Corrective Maintenance (CM)
i. The Contractor shall perform CM for LCMSMS, nitrogen generator and UPS (inclusive of the battery) within the stated turnaround time. | Y
ii. The scope of CM is as specified in Appendix 6.2 to Section III. | Y
c. Technical Support during office hours from 0800 to 1800hrs.
d. 2-Way Transportation of the equipment if maintenance is done off-site. | Y
PREVENTIVE MAINTENANCE (PM) FOR LCMS/MS SYSTEM
2.3 The Contractor shall provide the detailed proposed scope of works for the PM upon submission of tender. | Y
The general scope of work for the PM for LCMS/MS shall be as follows:
a) Verify the pulsation of flow from the LC pump and ensure that there is correct output flow. | Y
b) Inspect the system for leakage, cable problems, tubing problems, water issues and chemical residue. Repair and/or replace the parts if necessary. | Y
c) Verify the proper functioning of the temperature sensor inside the column oven. | Y
d) Verify the proper functioning of the mass spectrometer Q1 and Q3 in both positive and negative mode. | Y
e) Conduct tuning results check on the Q1 and Q3 scan mode. A tuning check on the MS/MS mode shall also be done. | Y
f) Inspect and ensure that the vacuum pump is operating well. | Y
g) Cleaning of the ionization probe/ assembly. Change the ionization probe if necessary. | Y
h) Inspect and ensure that the rotary pump is operating well. | Y
i) To provide and change the oil for the rotary pump. | Y
j) Perform necessary PM checks. | Y
k) A service report and PM report shall be submitted after each PM. | Y
The general scope of work for the PM for nitrogen generator shall be as follows:
a) Inspect and ensure that the nitrogen generator is operating well | Y
b) Perform the following test - leak checks, system logs, relief valve testing. | Y
c) Inspect the compressor. If the compressor is beyond the third refit, it shall be replaced during the PM without any cost. | Y
d) Supply and replace annual service kit for every PM | Y
e) Turn on the generator to build up set point pressure | Y
The general scope of work for the PM for UPS shall be as follows:
a) Verify the proper functioning of the UPS. | Y
b) Check the battery of the UPS and replace if necessary. | Y
c) A service report shall be submitted after each PM. | Y
50 of 83
Tender No.: HOMHQ0ETT16000310
Appendix 6.2 to Section III
CORRECTIVE MAINTENANCE (CM) FOR LCMS/MS SYSTEM
2.4 The Contractor shall provide the detailed proposed scope of works for the CM upon submission of tender. The general scope of work for the CM for the LCMS/MS (inclusive of pump overhaul), nitrogen generator and UPS (inclusive of battery) is as follows:
a) To be performed as per request by the laboratories for the LCMS/MS, nitrogen generator and UPS. The response time (RT) upon receipt of fault call is within next business day of fault call from Mondays to Fridays (0800 to 1800hrs).
b) To include all parts and labour used.
c) To include all the necessary system upgrade without additional cost.
d) The charges for the corrective maintenance are to be covered under the comprehensive maintenance plan.
e) To provide technical support during office hours (0800 to 1800hrs).
f) If corrective maintenance is done off-site, the price quoted shall include 2-way transportation of the equipment.
g) The Contractor shall make arrangements with the Authority to collect the defective system within 3 business days of fault call.
h) A service report shall be submitted after each corrective maintenance
i) On-Site Turn-Around-Time (i.e., the time period between the time the Contractor arrives at the Site to the time the repaired Equipment and/or Services are successfully completed and accepted on site by the Authority) required is 7 calendar days.
j) Workshop Repair Turn-Around-Time (i.e. the time period between the receipt of defective Equipment by the Contractor and the receipt and acceptance of the repaired Equipment by the Authority) required is 21 calendar days.
k) All repairs of parts and servicing works carried out shall carry a warranty period of 6 months. Parts that are replaced shall come with a minimum 1 year warranty.