Health Sciences Authority (HSA)

Tender Analysis Report

- for -

ACQUITY H-Class UPLC System coupled with ACQUITY Fluorescence Detector (FLD)

Date			: 22 November 2017

Presented by	: Waters Singapore Solution Centre

Prepared by	: Chantel Lee (Applications Consultant)

Table of Contents

Section 1 – Analysis Objectives	3

Section 2 – Analysis Information	3

Section 3 – Sample Information	3

Section 4 – Instrument Parameters	4

Section 5 – Results	5

Section 6 – Summary	7

# Section 1 – Analysis Objectives

The main objective is to develop an instrument method for Aflatoxins B & G analysis with the following requirements:-

- maximum run time of 4 minutes;

- resolution between closest eluting peaks must be at least 2

- S/N of all analyte peaks must be greater than 3; and

- no any form of derivatization should be used.

# Section 2 – Analysis Information

Name of Analyst		: Chantel Lee

Date of Sample Collection	: 10 November 2017

Date of Analysis		: 21 November 2017

# Section 3 – Sample Information

| Sample | Description |
| --- | --- |
| A | Direct analysis |
| B | Diluted by 40 times with given diluents (10% MeOH in 1% Acetic Acid) |

# Section 4 – Instrument Parameters

LC Parameters

| Parameters | Description |
| --- | --- |
| LC Instrument | ACQUITY H-Class UPLC (50uL Sample Loop) |
| Column | ACQUITY BEH C18, 2.1 x 100 mm, 1.7um |
| Run Time | 5 min |
| Mobile Phase | A = Water B = Methanol C = Acetonitrile Isocratic Elution 64:18:18 A:B:C |
| Flow Rate | 0.4 mL/min |
| Wash Solvent | 50/50 Water/ACN |
| Purge Solvent | 90/10 Water/Methanol |
| Column Temperature (ºC) | 30 |
| Sample Temperature (ºC) | 4 |
| Injection Volume (µL) | 20 |

## Detector Parameters

| Parameters | Description |
| --- | --- |
| Detector | ACQUITY Fluorescence Detector (FLD) (Large Volume Flow Cell) |
| Excitation Channel | 365 nm |
| Emission Channel | 0.00 min – 456 nm 2.65 min – 429 nm |
| Sampling Rate | 20 pts/sec |
| PMT Gain | 10 |
| Filter Time Constant | Normal |

# Section 5 – Results

Sample A

| Aflatoxins | Retention Time (min) | Peak Area | Peak Height | Peak Width at Half Height (min) | S/N (USP)* | Resolution** |
| --- | --- | --- | --- | --- | --- | --- |
| G2 | 2.035 | 1,461,039 | 593,540 | 0.037 | 8,454.30 | - |
| G1 | 2.515 | 349,696 | 115,146 | 0.046 | 1,408.27 | 6.84 |
| B2 | 2.767 | 1,320,105 | 386,537 | 0.051 | 1,297.01 | 3.04 |
| B1 | 3.476 | 244,740 | 62,427 | 0.062 | 303.04 | 7.46 |

*S/N(USP) is auto-calculated by Empower using the formula: SN=2 ×( Height-0.5 × NoiseScalingNoiseScaling)

** Resolution is calculated using the formula: Rs= t2- t1×1.18( w0.5h, 1+w0.5h, 2 )

Sample B

| Aflatoxins | Retention Time (min) | Peak Area | Peak Height | Peak Width at Half Height (min) | S/N (USP)* | Resolution** |
| --- | --- | --- | --- | --- | --- | --- |
| G2 | 2.037 | 41,763 | 16,678 | 0.038 | 235.80 | - |
| G1 | 2.518 | 9,912 | 3,251 | 0.046 | 39.28 | 6.76 |
| B2 | 2.768 | 39,040 | 11,038 | 0.053 | 36.10 | 2.98 |
| B1 | 3.482 | 6.714 | 1,846 | 0.059 | 8.21 | 7.52 |

*S/N(USP) is auto-calculated by Empower using the formula: SN=2 ×( Height-0.5 × NoiseScalingNoiseScaling)

** Resolution is calculated using the formula: Rs= t2- t1×1.18( w0.5h, 1+w0.5h, 2 )

# Section 6 – Summary

- A UPLC-FLR method under 4 minutes was used to analyze 4 aflatoxins: B1, B2, G1 and G2.

- Baseline separation was achieved for all the 4 aflatoxins (resolution > 2.98).

- All the analyte peaks have S/N larger than 8.6.

- The ACQUITY H-Class UPLC system is capable of auto-blending up to four-solvents, avoiding the needs of solvent pre-mixing which can be prone to human error.

- The ACQUITY H-Class UPLC coupled with ACQUITY FLD using the large volume flow cell, provides the needed analysis sensitivity for the analysis of aflatoxins B and G without derivatization.