Requirement Specifications for Supply, Delivery, Installation and Commissioning of one (1) unit of Fraction Collector for lab use

| Agri-Food & Veterinary Authority (AVA) is inviting interested suppliers to quote for Supply, Delivery, Installation and Commissioning of one (1) unit of Fraction Collector for lab use |  |  |  |
| --- | --- | --- | --- |
| Prices quoted for each line item in GeBIZ shall be deemed final, all inclusive (including shipping, freight, handling and any other costs) and firm to fulfil the requirement stated in the Requirement Specifications. Any additional claims on grounds of omission shall not be entertained. |  |  |  |
| Item | Quantity | Product Description | Unit Price S$ [Exclude GST] |
| 1 | 1 unit | Fraction Collector # Able to operate at 220/240 volts, 50/60Hz # Sample collection completely controlled by software - Fully compatible with our existing Postnova AF2000 Asymmetric Flow Field Flow Fractionation (AF4) system software # Must not require the use of an additional computer to control the fraction collector (due to space constraint) # Allow specific collection of targeted peaks and peak fractions # Methods of collection: # time - drop count - drop volume - signal - Metal free system - Supplied sample rack with capacity for at least 40 numbers of 2ml vials - Dimensions: width max 320mm, Height max 600mm, Depth max 400mm - Weight: Max 25kg - To supply at least 1 set of vials or tubes for use with the fraction collector - Contractor must install and commission fraction collector for use with our existing Postnova AF2000 Asymmetric Flow Field Flow Fractionation (AF4) system and ensure that they can communicate with each other. - One-year warranty against defective parts from date of commissioning - Contractor shall supply handbooks, installation guides or operation manuals required for the installation, operation and maintenance of the system - Contractor should have a facility with an in house team of technicians to provide after sales services within 2 working days during warranty period | S$ (Price to be entered in GeBIZ) |
| OTHER INFORMATION: - Supplier to submit duly completed Compliance Table (Annex A) and product information (e.g. reports, certificates etc.) via GeBIZ for evaluation. - Prices quoted for each line item in GeBIZ shall be deemed final, all inclusive (including shipping, freight, handling and any other costs) and firm to fulfil the requirement stated in the Requirement Specifications. Any additional claims on grounds of omission shall not be entertained. - Delivery shall be made within 6 to 8 weeks upon order confirmation. - Equipment/operation manual to be provided during delivery of item. - Goods to be delivered to No. 10 Perahu Road, Singapore 718837. |  |  |  |
| EVALUATION CRITERIA Offers will be evaluated based on the lowest that is technically acceptable. |  |  |  |
| Person to contact for clarification: - Mdm Chua Sew Lay, Tel no. 6795 2814/892 or Email: chua_sew_lay@ava.gov.sg |  |  |  |

ANNEX A

Compliance Table

(TO BE COMPLETED AND SUBMITTED VIA GEBIZ)

- Compliance to Specifications. Supplier to indicate Yes or No.

- If there is any variation to specifications embodied in this Section of this ITQ document, the supplier must state the variations in the “Remarks” column (Annex A) clearly.

| Specifications | Supplier to indicate “YES” or “NO” |  |
| --- | --- | --- |
|  | Yes/No | Remarks |
| - Able to operate at 220/240 volts, 50/60Hz |  |  |
| - Sample collection completely controlled by software |  |  |
| - Fully compatible with our existing Postnova AF2000 Asymmetric Flow Field Flow Fractionation (AF4) system software - Must not require the use of an additional computer to control the fraction collector (due to space constraint) - Allow specific collection of targeted peaks and peak fractions - Methods of collection: - time - drop count - drop volume - signal |  |  |
| - Metal free system |  |  |
| - Supplied sample rack with capacity for at least 40 numbers of 2ml vials |  |  |
| - Dimensions: width max 320mm, Height max 600mm, Depth max 400mm |  |  |
| - Weight: Max 25kg |  |  |
| - To supply at least 1 set of vials or tubes for use with the fraction collector |  |  |
| - Contractor must install and commission fraction collector for use with our existing Postnova AF2000 Asymmetric Flow Field Flow Fractionation (AF4) system and ensure that they can communicate with each other. |  |  |

| - One-year warranty against defective parts from date of commissioning |  |  |
| --- | --- | --- |
| - Contractor shall supply handbooks, installation guides or operation manuals required for the installation, operation and maintenance of the system |  |  |
| - Contractor should have a facility with an in house team of technicians to provide after sales services within 2 working days during warranty period. |  |  |
| - Supplier to submit duly completed Compliance Table (Annex A) and product information (e.g. reports, certificates etc.) via GeBIZ for evaluation. |  |  |
| - Delivery shall be made within 6 to 8 weeks upon order confirmation. |  |  |

Submitted by: _________________________

Company: ____________________________

Email: _______________________________

Telephone contact: _____________________