| S/N | Description of item | Compliance with specifications |  | Information/Remarks |
| --- | --- | --- | --- | --- |
|  |  | Yes | No |  |
|  | Requirement specification |  |  |  |
| 1 | System Controller |  |  |  |
| 1.1. | Direction system controller from touch screen keypad operation with option of touch panel operation |  |  |  |
| 1.2. | Function as a communication bus module with data buffering capability |  |  |  |
| 1.3 | Able to acquire up to 24 hours for one analysis, at 500ms sampling rate |  |  |  |
| 1.4 | Able to store up to 20 analysis files with a total up to 400 steps of time programs |  |  |  |
| 1.5 | Must be controllable from a web-based interface via a network. It allows the system to be controlled, monitored and maintained via Internet Explorer Web browser |  |  |  |
| 1.6 | Must have an expert function in that if pressure falls below specified value, the expert function will automatically purge the mobile phase |  |  |  |
| 1.7 | The system should be easily upgradable to enable mobile phase reservoir monitoring; providing life update on reservoir volume, and usage through PC or mobile device |  |  |  |
| 2 | Solvent Delivery System |  |  |  |
| 2.1 | Should be capable to configure to run isocratic or quaternary low-pressure gradient modes |  |  |  |
| 2.2 | The pumping mechanism design should be parallel-type plunger with 10 uL/stroke |  |  |  |
| 2.3 | The flow rate should be set between 0.0001 to 10 ml/min in 0.1 ul/min step |  |  |  |
| 2.4 | Flow rate accuracy should be ±1% |  |  |  |
| 2.5 | Flow rate precision should be less than ±0.06% RSD or 0.02 min SD |  |  |  |
| 2.6 | The maximum pump pressure must be able to set for running UHPLC analysis method. The maximum pressure setting range should not be less than 70 MPa at <3 ml/min flowrate |  |  |  |
| 2.7 | The standard gradient mixer should have delay volume of 20 µl, with mixer recognition device |  |  |  |
| 2.8 | The unit should be capable to run pH range 1 to 14 |  |  |  |
| 2.9 | The composition accuracy should be ±0.5% |  |  |  |
| 2.10 | Should employ active check valves that allow stable delivery of even non-polar organic solvents such as hexane |  |  |  |
| 2.11 | Should include intelligent feature that capable of carrying out auto-diagnostic on system status to detect air bubble formation, and perform auto-purging to remove air bubble to restore system pressure |  |  |  |
| 2.12 | Should include intelligent feature to automatically halve or user-settable value of the setting flowrate to protect column |  |  |  |
| 2.13 | It must have a leak sensor as safety feature |  |  |  |
| 2.14 | Automated rinsing kit should be supplied as standard |  |  |  |
| 3 | Degassing Unit |  |  |  |
| 3.1 | Membrane degassing unit for 5 flow lines |  |  |  |
| 3.2 | Maximum operating flow rate up to 10ml/min per flow line |  |  |  |
| 3.3 | An internal capacity of 400 µl per flow line should be available for decreasing the time |  |  |  |
| 3.4 | Pressure resistance should be maintained at ±0.1 MPa |  |  |  |
| 3.5 | It must have a leak sensor as a safety feature |  |  |  |
| 3.6 | Error status should be transferrable to the operating software |  |  |  |
| 3.7 | It should include a drain-pan |  |  |  |
| 3.8 | Self-cleaning capability that extends vacuum pump life by drawing in air when the pump is running |  |  |  |
| 3.9 | The liquid contact surfaces of the degasser should employ special synthetic polymers designed for all solvents |  |  |  |
| 4 | Auto-Sample Injector with Accessories |  |  |  |
| 4.1 | The auto sampler design should be a flow through design with variable total volume injection, with loop injection should be available as option. |  |  |  |
| 4.2 | Sample injection volume should be variable between 0.1 µl to 50 µl, with optional sample loop up to 2000 uL should be available as option |  |  |  |
| 4.3 | Injection volume step 0.1 uL |  |  |  |
| 4.4 | It must be capable of ultra fast injection time of <7 sec/sample, with higher injection speed is preferred for high throughput analysis capability |  |  |  |
| 4.5 | The maximum operating pressure must be at least 70 MPa |  |  |  |
| 4.6 | Flow line rinse capability both before and after sampling should be possible |  |  |  |
| 4.7 | It should be able to use at least two different solvents for rinsing, with more options available is preferred for effective rinsing. |  |  |  |
| 4.8 | System injection volume accuracy should be ≤±1% |  |  |  |
| 4.9 | Injection linearity ≥0.9999 |  |  |  |
| 4.10 | High sample processing volume capability should be available; 252x1 mL vials, 162x1.5 mL vials, 84x4 mL vials, 36x10 mL vials. The system should be designed to handle 96-well microtiter plate or 384-well microtiter plate |  |  |  |
| 4.11 | Temperature setting range should be from 4 to 45ºC, with smart air circulation cooling design. Temperature accuracy achievable ±2˚C |  |  |  |
| 4.12 | Dip rinsing outside injection needle and injection port rinsing should come as standard, with internal rinsing |  |  |  |
| 4.13 | The injection needle mechanism should incorporate zero dead volume between injection needle and injection loop for lowest carryover |  |  |  |
| 4.14 | The Carry over must be below 0.0015 % without rinse and below 0.0003% with rinse |  |  |  |
| 4.15 | Operating pH range from 1 to 14 |  |  |  |
| 4.16 | It should be capable of standalone operation with digital touch-screen operating keys |  |  |  |
| 4.17 | The autosampler should be capable to do on-site upgrade to add another injection pot, to perform individual dual injection mode for 2 individual analysis under one system |  |  |  |
| 4.18 | It should have safety features like leak sensor and automatic rack and vial recognition. Maintenance kit should be available |  |  |  |
| 4.19 | It should have functions for maintenance and validation which are accessible by a dedicated operation button |  |  |  |
| 5 | Column Oven |  |  |  |
| 5.1 | It should be forced-air-circulation type for uniform temperature distribution with a quick feedback mechanism to maintain constant temperature level even when power source voltage fluctuates |  |  |  |
| 5.2 | The temperature range should be ambient -10˚C to 85˚C |  |  |  |
| 5.3 | Temperature control precision should be ±0.1˚C |  |  |  |
| 5.4 | The column oven compartment should be able to handle at least 3x30 cm L. column |  |  |  |
| 5.5 | The oven should have temperature limit device and temperature fuse and a solvent leak sensor |  |  |  |
| 5.6 | It should have built-in slots for valve control functions |  |  |  |
| 5.7 | Capable of complex temperature programming in linear and step programs |  |  |  |
| 6 | Photodiode Array Detector (PDA) |  |  |  |
| 6.1 | The detector should be designed with light source from D2 Lamp and Tungsten Lamp |  |  |  |
| 6.2 | UV Cut-off filter should be made available to prevent photo-degradation on photo-sensitive chemicals |  |  |  |
| 6.3 | The wavelength range should be 190 nm - 800 nm |  |  |  |
| 6.4 | The photo-diode array detector should have at least 1024 elements, with an element resolution of 0.6 nm/element |  |  |  |
| 6.5 | The PDA Detector should include Advanced TC Optics design to allow 3-way temperature control on Lamp, Flow Cell and Spectrometer for maximized baseline stabilisation |  |  |  |
| 6.6 | The detector must have 2 modes of operation using a variable slit: High Resolution mode at a slit width of 1 nm and a High Sensitivity mode at a slit width of 8nm |  |  |  |
| 6.7 | A Standard flow cell [1.0 μL volume, 10 mm cell path length, 8 MPa pressure max.] with temperature control should be available as standard |  |  |  |
| 6.8 | Wavelength reproducibility should be ±0.1nm |  |  |  |
| 6.9 | Wavelength accuracy should be ±1 nm |  |  |  |
| 6.10 | The Drift must be equal or smaller than ≤0.4x10-3 AU/H |  |  |  |
| 6.11 | The Noise Level must be equal or smaller than ≤4.5x10-6 AU |  |  |  |
| 6.12 | Linearity should be equal or more than 2.5 AU |  |  |  |
| 6.13 | Light sources and cell should be accessible from the front for easy maintenance |  |  |  |
| 6.14 | Operating pH 1 to 13 |  |  |  |
| 6.15 | It must have data buffering function to collect up to 20 min of data for all wavelength regions in case of PC malfunctions |  |  |  |
| 6.16 | The detector should be designed with Flow Cell ID Chip to record flow cell ID and Lamp ID Chip to record Lamp ID to ensure data performance |  |  |  |
| 7 | Fluorescence Detector |  |  |  |
| 7.1 | The light source should be a Xenon lamp and a low-pressure mercury lamp should be available as an option for checking wavelength accuracy |  |  |  |
| 7.2 | The wavelength range should be 200nm to 650nm. An option should be available to achieve the range of 200 to 750 nm |  |  |  |
| 7.3 | The Spectral bandwidth should be 20nm |  |  |  |
| 7.4 | Wavelength accuracy should be ± 2nm |  |  |  |
| 7.5 | Wavelength reproducibility should be ± 0.2nm |  |  |  |
| 7.6 | The Signal to noise ratio for Water Raman peak should be minimum 1200 |  |  |  |
| 7.7 | Cell temperature control from 4º c to 40 º c in 1 º c step should be available as an option |  |  |  |
| 7.8 | The detector must be capable of monitoring any 2 wavelengths between 200 and 650nm simultaneously |  |  |  |
| 7.9 | For ease of maintenance, the xenon lamp and flow cell must be accessible from the front panel with no tools required and alignment free |  |  |  |
| 8 | Fraction Detector |  |  |  |
| 8.1 | Drive system : Arm movement (X-Y) system |  |  |  |
| 8.2 | Collection method : Solenoid |  |  |  |
| 8.3 | Fraction collection modes: Basic programmable mode, Time Program mode, Overlapping Peak Collection mode, Off-Scale Detector Signals mode, Ratio Chromatogram mode, etc |  |  |  |
| 8.4 | It should be capable of collecting at least 144 fractions |  |  |  |
| 8.5 | It should be able to use over a wide range of flow rates, covering small and large scale preparative work. The maximum flow rate settable is 150ml/min |  |  |  |
| 8.6 | It should be flexibly applied for various fractionation both manual and continuous automatic fractionation |  |  |  |
| 8.7 | The delay time from the detector to the fractionation head should be calculated automatically every time the flow rate is changed |  |  |  |
| 8.8 | When exceeding the volume of the vial, the collector must automatically move to the next vial before it is overflow to prevent loss of precious fractions |  |  |  |
| 9 | Accessories |  |  |  |
| 9.1 | Must include a manual injector with 20 uL and 2mL stainless steel loop used for semi-preparative system. 1 mL syringe must be included as standard |  |  |  |
| 9.2 | Columns must be provided as follows: 9.2.1. 2 sets of C18 column 3µ, 75 x 4.6mm column 9.2.2. 2 sets of C18 column 5µ, 250 x 4.6mm column |  |  |  |
| 9.3 | Caffeine solution must be provided as follows: 9.3.1. 2 bottles of 250mg/L 9.3.2. 2 bottles of 20mg/L |  |  |  |
| 9.4 | Hand tool kit |  |  |  |
| 9.5 | Minimum 2 packs of 3.5 mL test tubes (glass) with OD 10mm x L 75mm vial for rack no.1 |  |  |  |
| 9.6 | Minimum 1 pack of rack no.1 |  |  |  |
| 9.7 | Minimum 8 reservoir bottles with cap |  |  |  |
| 9.8 | Must include mobile phase monitor unit with bottle holder and adapter |  |  |  |
| 9.9 | Installation kit should be included |  |  |  |
| 10 | Hardware |  |  |  |
| 10.1 | Intel Core 2 Duo processor or higher |  |  |  |
| 10.2 | At least 8GB RAM on board |  |  |  |
| 10.3 | At least 250 GB hard disk |  |  |  |
| 10.4 | DVD-RW drive |  |  |  |
| 10.5 | At least 21 inches monitor |  |  |  |
| 10.6 | Key board and mouse |  |  |  |
| 10.7 | Pre-installed Windows 10 or higher |  |  |  |
| 10.8 | Laser printer |  |  |  |
| 11 | Software |  |  |  |
| 11.1 | Software must be able to link with Windows Users or Active Directory Users |  |  |  |
| 11.2 | It should cover full one-point digital instrument control, qualitative and quantitative processing, report creation and self-diagnosis |  |  |  |
| 11.3 | Sample schedule wizard function should be standard |  |  |  |
| 11.4 | There should be an on-line help function context sensitive |  |  |  |
| 11.5 | The reporting format should be flexible and easy to use in any desired format |  |  |  |
| 11.6 | The data can be converted to other (AIA, ASCII) formats. Spread Sheet software and word-processing software can be readily employed to provide data in tables or graphs through industry standard protocols |  |  |  |
| 11.7 | The software should allow automatic execution of system checks, auto-purge and baseline checks |  |  |  |
| 11.8 | Software must register all events (log files) audit trails for Data, Method, Batch, Report, System Policy and User Administration |  |  |  |
| 11.9 | Software must display the online status of instruments (Name, Type, Analysis, Status, User Running, Queued Count, Estimated End time) |  |  |  |
| 11.10 | In-built pdf generator feature |  |  |  |
| 11.11 | Functions to check PC Information, Software Program Files Check, User List, User Groups, Group Rights, Security Policy, Instrument Connection information from software and printable in pdf format |  |  |  |
| 11.12 | An audio-visual multi-media CD-ROM for Maintenance and Troubleshooting must be provided |  |  |  |
| 11.13 | System suitability, System security as well as System check functions must be provided which comply with Good Laboratory Practice (GLP) and Regulatory Conformity |  |  |  |
|  | Other requirement specifications |  |  |  |
| 1 | Delivery |  |  |  |
| 1.1 | Delivery to offshore SJINML, 51, Saint John’s Island, Singapore 098634 at pre-agreed date and timing |  |  |  |
| 1.2 | Delivery must be made within 8 weeks from the issue of PO, unless otherwise stated by NUS |  |  |  |
| 1.3 | Vendors shall have a technical engineer to oversee the loading of equipment onto the designated boat/ferry and off-loading of equipment from the designated boat/ferry to SJINML |  |  |  |
| 2 | Installation, commissioning and training |  |  |  |
| 2.1 | Installation and commissioning are required at the offshore SJINML, 51, Saint John’s Island, Singapore 098634 |  |  |  |
| 2.2 | On-site installation, commissioning and training shall be conducted by a qualified factory-trained engineer |  |  |  |
| 2.3 | Vendor shall complete testing and commissioning of the system within 2 weeks from the date of delivery, unless otherwise stated by NUS |  |  |  |
| 2.4 | Training must be provided and conducted on offshore SJINML, 51, Saint John’s Island, Singapore 098634 by vendor on the operation and maintenance of the equipment (at least 2 users) |  |  |  |
| 2.5 | Vendor to provide both on-site and operator training for users on the system start-up, usage, maintenance, quality control, trouble shooting, etc. including comprehensive classroom training |  |  |  |
| 3 | Warranty |  |  |  |
| 3.1 | The warranty shall commence only upon successful completion of the Acceptance Test or commissioning |  |  |  |
| 3.2 | At least 36 months warranty, complete support shall include the following at no extra cost: 3.2.1 Service parts used for repairs 3.2.2 Travel and labour expenses of service engineer 3.2.3. Unlimited service call |  |  |  |
| 3.3 | Vendor to provide service guarantee: should the system require service during the warranty period, vendor must guarantee or replacement of instrument for free |  |  |  |
| 3.4 | Vendor to have logistic support to ensure that over at least 95% of the service parts are readily available and upkeep delivery within 24 hours |  |  |  |
| 3.5 | Vendor shall be authorized distributors or agents of equipment |  |  |  |
| 3.6 | The local service team shall provide on-site support within 48 hours upon contact |  |  |  |
| 3.7 | All services and repairs shall be done at the offshore SJINML, 51, Saint John’s Island, Singapore 098634, whenever applicable |  |  |  |
| 3.8 | Minimum one time of preventive maintenance should be inclusive for second and third year |  |  |  |
| 3.9 | At least 1 set of PM kit should be inclusive |  |  |  |
| 4 | Other requirements |  |  |  |
| 4.1 | The system should be accompanied with Conformity Certificate |  |  |  |
| 4.2 | A demo unit shall be physically available in Singapore. Site visit for the demo may be necessary upon request by the buyer |  |  |  |
| 4.3 | The vendor shall have proven track record of selling the unit model both locally and internationally for the past 2 years |  |  |  |
| 4.4 | The working language of the manual must be available in English |  |  |  |
| 4.5 | Vendor must demonstrate that it has a proven appropriate set-up and capability to provide after-sales service efficiently and effectively. The supplier should have in his facility a similar system to that proposed in this tender for training purpose |  |  |  |
| 4.6 | A declaration of Conformity certificate must be provided |  |  |  |
| 4.7 | A declaration of System Validation certificate must be provided |  |  |  |
| 4.8 | All modules must be GLP compliant |  |  |  |