# Tandem Quadrupole Mass Spectrometry

# Sample Analysis Report

Customer:  Singapore Food Agency

Samples: 	DR-STD

DR-BLK A, DR-BLK B

DR-SPK A, DR-SPK  B

Instruments: 	ACQUITY UPLC I Class

Xevo TQ-XS

Prepared by: Lee Jun Xiang (Applications Consultant

Waters Solution Centre

Table of Contents

1	Executive Summary	1

2	Instrument Methods	3

2.1	UPLC Method	3

2.2	MS Method	4

2.2.1	Setting up a screening method for potential drugs substance in sample	4

3	Results	6

3.1	Screening of potential compounds in DR-STD (0.1 ppm)	6

3.2	Setting up a screening and quantitation method for samples (DR-BLK A, DR-BLK B, DR-SPK A, DR-SPK B)	17

3.3	Screening and Quantitation result for DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B using established MRM method and calibration curves	28

3.3.1	Results for DR-BLK A, DR-BLK B	28

3.3.2	Results for DR-SPK A, DR-SPK B	30

4	Conclusions	37

5	List of Figures	38

6	List of Tables	38

# Executive Summary

Screening of standards DR-STD

- Standard DR-STD was screened using an MRM method that was established using Quanpedia. Out of the 250 potential compounds, 84 compounds were identified. Table 1 contains the summary of the list of compounds being identified in DR-STD (0.1 ppm).

- Calibration curves were established for all the detected analytes to further screen and quantitate for samples DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B. All calibration curves achieve good linearity of >0.995.

Screening and quantitation of samples DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B

- Triplicate analysis of the samples was performed. Total number of compounds identified in the samples are summarized as follow:

  - DR-SPK A: 32 drugs identified

  - DR-SPK B: 30 drugs identified

  - DR-BLK A: 5 drugs identified

  - DR-BLK B: 4 drugs identified

- Quantitation result and identities of all the drugs detected in the samples are shown in Table 2.

| DR-STD (0.1)ppm |  |  |  |
| --- | --- | --- | --- |
| S/N | Drugs identified | S/N | Drugs identified |
| 1 | 16-beta-hydroxy stanozolol | 43 | Lasalocid |
| 2 | 17-alpha-methyltestosterone | 44 | Mabuterol |
| 3 | 4-Epichlortetracycline | 45 | Maduramycin |
| 4 | 4-Epioxytetracycline | 46 | Mapenterol |
| 5 | 4-Epitetracycline | 47 | Marbofloxacin |
| 6 | 5-Hydroxyflunixin | 48 | Metronidazole |
| 7 | 6-alpha-Methylprednisolone | 49 | Metronidazole-hydroxy |
| 8 | Albendazole | 50 | Monensin |
| 9 | Altrenogest | 51 | Monepantel sulfone |
| 10 | Amoxicillin | 52 | Nafcillin |
| 11 | Ampicillin | 53 | Nandrolone |
| 12 | Brombuterol | 54 | Narasin |
| 13 | Buqinolate | 55 | Norfloxactin |
| 14 | Carbmazepine | 56 | Novobiocin |
| 15 | Cefalexin | 57 | Oxacillin |
| 16 | Cefazolin | 58 | Oxibendazole |
| 17 | Cefoperazone | 59 | Penicillin G |
| 18 | Ceftiofur | 60 | Penicillin V |
| 19 | Cephalothin | 61 | Prednisolone |
| 20 | Cephradine | 62 | Ractopamine |
| 21 | Chlortetracycline | 63 | Robenidin |
| 22 | Cimbuterol | 64 | Ronidazole |
| 23 | Clenbuterol | 65 | Salbutamol |
| 24 | Clenpenterol | 66 | Sarafloxacin |
| 25 | Clopidol | 67 | Stanozolol |
| 26 | Cloxacillin | 68 | Sulfadimethoxine/Sulfadoxine |
| 27 | Decoquinate | 69 | Sulfanilamide |
| 28 | Desacetylcefapirin | 70 | Terbutaline |
| 29 | Dexamethasone | 71 | Thiabendazole |
| 30 | Dicloxacillin | 72 | Tildipirosin |
| 31 | Difloxacin | 73 | Vancomycin |
| 32 | Dimetridazole | 74 | Virginiamycin |
| 33 | Doramectin | 75 | Zilpaterol |
| 34 | Doxycycline/Tetracycline | 76 | 17-alpha-ethylnylestradiol |
| 35 | Enrofloxacin | 77 | betamethasone |
| 36 | Epi-testosterone/Testosterone | 78 | Diclazuril |
| 37 | Eprinomectin | 79 | Diethylstilbesterol |
| 38 | Flumequine | 80 | Estradiol |
| 39 | Furazolidone | 81 | Estriol |
| 40 | Hydoxy dimetridazole | 82 | Estrone |
| 41 | Ipronidazole | 83 | Ethyl estradiol |
| 42 | Ipronidazole hydroxy | 84 | Nicarbazin |

Table 1: List of drugs identified in DR-STD (0.1 ppm)

| S/N | DR-SPK A |  | S/N | DR-SPK B |  | S/N | DR-SPK A |  | S/N | DR-SPK B |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Drugs identified | Conc (ppb) |  | Drugs identified | Conc (ppb) |  | Drugs identified | Conc (ppb) |  | Drugs identified | Conc (ppb) |
| 1 | 16-beta-hydroxy stanozolol | 3.02 | 1 | 16-beta-hydroxy stanozolol | 2.7 | 24 | Salbutamol | 0.12 | 24 | Stanozolol | 1.7 |
| 2 | 4-epioxytetracycline | 2.07 | 2 | 4-epioxytetracycline | 3.0 | 25 | Stanozolol | 2.72 | 25 | Sulfanilamide | 67.9 |
| 3 | Albendazole | 309.19 | 3 | Albendazole | 58.2 | 26 | Sulfanilamide | 84.70 | 26 | Zilpaterol | 0.1 |
| 4 | Altrenogest | 0.08 | 4 | Amoxicillin | 1.0 | 27 | Zilpaterol | 0.12 | 27 | betamethasone | 0.7 |
| 5 | Amoxicillin | 0.51 | 5 | Ampicillin | 1.6 | 28 | betamethasone | 2.72 | 28 | Diethylstilbesterol | 0.3 |
| 6 | Ampicillin | 0.80 | 6 | Chlortetracycline | 39.4 | 29 | Diethylstilbesterol | 1.06 | 29 | Ethyl estradiol | 0.2 |
| 7 | Chlortetracycline | 21.63 | 7 | Clenbuterol | 0.2 | 30 | Estrone | 0.06 | 30 | 17-alpha-ethylnylestradiol | 0.3 |
| 8 | Clenbuterol | 0.22 | 8 | Dexamethasone | 6.7 | 31 | Ethyl estradiol | 1.04 |  |  |  |
| 9 | Dexamethasone | 6.78 | 9 | Dimetridazole | 0.7 | 32 | 17-alpha-ethylnylestradiol | 1.08 |  |  |  |
| 10 | Dimetridazole | 0.92 | 10 | Epi-testosterone/Testosterone | 13.9 |  |  |  |  |  |  |
| 11 | Furazolidone | 752.13 | 11 | Furazolidone | 773.8 |  |  |  |  |  |  |
| 12 | Hydoxy dimetridazole | 0.35 | 12 | Hydoxy dimetridazole | 0.3 |  |  |  |  |  |  |
| 13 | Ipronidazole | 1.22 | 13 | Ipronidazole | 1.2 | S/N | DR-BLK A |  | S/N | DR-BLK B |  |
| 14 | Maduramycin | 2.51 | 14 | Maduramycin | 3.8 |  | Drugs identified | Conc (ppb) |  | Drugs identified | Conc (ppb) |
| 15 | Metronidazole | 0.34 | 15 | Metronidazole | 0.3 | 1 | Albendazole | 130.1 | 1 | Albendazole | 49.4 |
| 16 | Metronidazole-hydroxy | 0.28 | 16 | Metronidazole-hydroxy | 0.3 | 2 | Altrenogest | 0.1 | 2 | Furazolidone | 363.1 |
| 17 | Monensin | 4.89 | 17 | Monensin | 5.4 | 3 | Furazolidone | 428.1 | 3 | Sulfanilamide | 78.5 |
| 18 | Nandrolone | 2.28 | 18 | Nandrolone | 2.1 | 4 | Monensin | 0.1 | 4 | 17-alpha-ethylnylestradiol | 0.3 |
| 19 | Novobiocin | 8.65 | 19 | Novobiocin | 8.5 | 5 | Sulfanilamide | 79.1 |  |  |  |
| 20 | Oxibendazole | 23.56 | 20 | Penicillin V | 28.1 |  |  |  |  |  |  |
| 21 | Penicillin V | 25.94 | 21 | Prednisolone | 8.4 |  |  |  |  |  |  |
| 22 | Prednisolone | 6.48 | 22 | Ractopamine | 0.2 |  |  |  |  |  |  |
| 23 | Ractopamine | 0.14 | 23 | Salbutamol | 0.1 |  |  |  |  |  |  |

Table 2: List of drugs identified and the amount present in sample (DR-SPK A, DR-SPK B, DR-BLK A and DR-BLK B)

# Instrument Methods

## UPLC Method

This method was used for screening and quantitation of drug residues in the standard (DR-STD) and samples (DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B). The UPLC system coupled to the MS system was ACQUITY UPLC I-Class System (SM-FTN). The UPLC conditions are as follows:

Binary Solvent Manager

- Analytical column:  Acquity UPLC CSH C18, 1.7 µm, 100 × 2.1 mm column

- Injection volume: 10 µL

- Flow rate: 0.4 mL/min

- Mobile phase:

For positive ionization:	A – 0.1% formic acid in H2O

B –0.1% formic acid in ACN:MeOH (50:50)

For negative ionization:	A – H2O

B – ACN

- Mobile phase gradient: 1% B (0.0 min) → 20% B (4.0 min) → 50% B (5.0 min) → 99% B (7.0 to 9.0 min) → 1% B (9.1 to 12.0 min)

Sample Manager FTN

- Wash solvent: 50:30:40 water:acetonitrile: IPA

- Purge solvent: 10% ACN in water

- Sample temperature: 8 ºC

- Column temperature: 40 ºC

## MS Method

### Setting up a screening method for potential drugs substance in sample

The MRM method for the screening of drugs was established using information based on Quanpedia. Triplicate experiments of screening were performed on samples (DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B), Results obtained were then process using Targetlynx Software.

The MS system was Xevo TQ-XS. The instrument setting is as follows:

- Capillary voltage: 3.10 kV (ESI positive), 2.50 kV (ESI negative)

- Source temperature: 150 ºC

- Desolvation temperature: 400 ºC

- Cone gas flow: 150 L/h

- Nebuliser gas flow: 7 bar

- Desolvation gas flow: 1000 L/h

# Results

## Screening of potential compounds in DR-STD (0.1 ppm)

A list of 250 compounds containing various classes of drugs such as coccidiostatic agents, anti-biotics, beta-agonists, macrolides and glycopeptide antibiotic is provided. 2 MRM transition was used for screening of the potential compounds in the DR-STD. A total of 84 compounds, 75 in positive ionization mode and 9 in negative ionization mode, were identified in DR-STD. Table 3 show the compounds identified in the mixture along with the MRM transition used. Due to some suspected compounds being isomers and having the same fragmentation patterns, they are both being suspected due to the lack of retention time information. Figure 1 display the MRM chromatogram of the detected compounds at 10 ppb.

Table 3. List of detected compounds in DR-STD using MRM method with 2 MRM transitions

| No | Compound | RT (min) | Quantifier Trace | Reference Trace |  | No | Compound | RT (min) |  | Quantifier Trace | Reference Trace |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 16-beta-hydroxy stanozolol | 6.38 | 345.1 > 81 | 345.1 > 95 |  | 7 | 6-alpha-Methylprednisolone | 6.43 |  | 375.4 > 357.3 | 375.4 > 339.3 |  |
| 2 | 17-alpha-methyltestosterone | 7.22 | 303.1 > 97 | 303.1 > 97 |  | 8 | Albendazole | 6.14 |  | 265.9 > 234.1 | 265.9 > 191.2 |  |
| 3 | 4-Epichlortetracycline | 4.45 | 479.3 > 444.2 | 479.3 > 462.2 |  | 9 | Altrenogest | 7.04 |  | 311 > 269 | 311 > 227 |  |
| 4 | 4-Epioxytetracycline | 4.05 | 461.1 > 426.1 | 461.1 > 443.2 |  | 10 | Amoxicillin | 2.16 |  | 366.1 > 349 | 366.1 > 133.7 |  |
| 5 | 4-Epitetracycline | 3.21 | 445.1 > 410.1 | 445.1 > 427.3 |  | 11 | Ampicillin | 3.84 |  | 350.1 > 105.9 | 350.1 > 159.9 |  |
| 6 | 5-Hydroxyflunixin | 7.04 | 313 > 226 | 313 > 252 |  | 12 | Brombuterol | 4.75 |  | 367.1 > 214 | 367.1 > 293.1 |  |
| No | Compound | RT (min) | Quantifier Trace | Reference Trace |  | No | Compound | RT (min) |  | Quantifier Trace | Reference Trace |  |
| 13 | Buqinolate | 6.20 | 361.8 > 260.2 | 361.8 > 204 |  | 28 | Desacetylcefapirin | 1.95 |  | 382 > 226 | 382 > 292 |  |
| 14 | Carbmazepine | 6.43 | 237 > 193 | 237 > 194 |  | 29 | Dexamethasone | 6.43 |  | 393.3 > 355.2 | 393.3 > 373.2 |  |
| 15 | Cefalexin | 3.65 | 348.2 > 158 | 348.2 > 139.9 |  | 30 | Dicloxacillin | 6.76 |  | 470 > 211.9 | 470 > 254 |  |
| 16 | Cefazolin | 5.60 | 455.1 > 323 | 455.1 > 156 |  | 31 | Difloxacin | 4.63 |  | 400.3 > 356.2 | 400.3 > 382.2 |  |
| 17 | Cefoperazone | 5.79 | 646.1 > 142.8 | 646.1 > 530.1 |  | 32 | Dimetridazole | 2.90 |  | 142 > 95.8 | 142 > 80.9 |  |
| 18 | Ceftiofur | 5.97 | 524.2 > 241.2 | 524.2 > 285 |  | 33 | Doramectin | 7.80 |  | 916.9 > 145 | 916.9 > 331.3 |  |
| 19 | Cephalothin | 6.20 | 419 > 315 | 419 > 359 |  | 34 | Doxycycline/Tetracycline | 3.21 |  | 445.2 > 410.1 | 445.2 > 428.1 |  |
| 20 | Cephradine | 4.10 | 350.1 > 158 | 350.1 > 176 |  | 35 | Enrofloxacin | 4.22 |  | 360.3 > 316.3 | 360.3 > 342.3 |  |
| 21 | Chlortetracycline | 5.22 | 479.1 > 444.1 | 479.1 > 462.1 |  | 36 | Epi-testosterone/Testosterone | 6.82 |  | 289.1 > 97 | 289.1 > 109 |  |
| 22 | Cimbuterol | 2.22 | 234.2 > 143 | 234.2 > 160 |  | 37 | Eprinomectin | 7.80 |  | 914.8 > 154 | 914.8 > 186.1 |  |
| 23 | Clenbuterol | 4.08 | 277.1 > 168.1 | 277.1 > 140 |  | 38 | Flumequine | 1.76 |  | 262.1 > 244 | 262.1 > 202 |  |
| 24 | Clenpenterol | 5.05 | 291.1 > 203 | 293.1 > 205 |  | 39 | Furazolidone | 6.49 |  | 226.2 > 139.2 | 226.2 > 122 |  |
| 25 | Clopidol | 3.76 | 192 > 101 | 192 > 103 |  | 40 | Hydoxy dimetridazole | 2.51 |  | 158 > 140.1 | 158 > 112 |  |
| 26 | Cloxacillin | 6.67 | 436.2 > 159.9 | 436.2 > 276.9 |  | 41 | Ipronidazole | 5.53 |  | 170 > 124 | 170 > 109.1 |  |
| 27 | Decoquinate | 7.76 | 418.3 >204.1 | 418.3 > 372.3 |  | 42 | Ipronidazole hydroxy | 4.96 |  | 186.1 > 127.9 | 186.1 > 167.9 |  |
| No | Compound | RT (min) | Quantifier Trace | Reference Trace |  | No | Compound | RT (min) |  | Quantifier Trace | Reference Trace |  |
| 43 | Lasalocid | 8.15 | 613.4 > 377.4 | 613.4 > 577.4 |  | 58 | Oxibendazole | 5.45 |  | 250 > 176.1 | 250 > 218.1 |  |
| 44 | Mabuterol | 4.97 | 311.1 > 217.1 | 311.1 > 237.1 |  | 59 | Penicillin G | 6.31 |  | 335.1 > 160 | 335.1 > 176 |  |
| 45 | Maduramycin | 8.17 | 939.5 > 877.6 | 939.5 > 719.5 |  | 60 | Penicillin V | 6.49 |  | 351.1 > 160.1 | 351.1 > 114 |  |
| 46 | Mapenterol | 5.35 | 325.1 > 237.1 | 325.1 > 217.1 |  | 61 | Prednisolone | 6.20 |  | 361.2 > 147 | 361.2 > 343.2 |  |
| 47 | Marbofloxacin | 3.30 | 363.1 > 72 | 363.1 > 320 |  | 62 | Ractopamine | 3.88 |  | 302.2 > 164.1 | 302.2 > 284.2 |  |
| 48 | Metronidazole | 2.64 | 172 > 82 | 172 > 128 |  | 63 | Robenidin | 6.10 |  | 334.2 > 111.1 | 334.2 > 155.2 |  |
| 49 | Metronidazole-hydroxy | 2.16 | 188 > 126 | 188 > 143.9 |  | 64 | Ronidazole | 3.00 |  | 201 > 140 | 201 > 55 |  |
| 50 | Monensin | 8.08 | 693.5 > 461.4 | 693.5 > 501.4 |  | 65 | Salbutamol | 1.78 |  | 240.2 > 148.1 | 240.2 > 222.1 |  |
| 51 | Monepantel sulfone | 6.76 | 504 > 186 | 504 > 165.9 |  | 66 | Sarafloxacin | 4.64 |  | 386.2 > 342.1 | 386.2 > 299.1 |  |
| 52 | Nafcillin | 6.75 | 415.2 > 199.2 | 415.2 > 171.1 |  | 67 | Stanozolol | 6.92 |  | 329 > 81 | 329 > 121 |  |
| 53 | Nandrolone | 6.69 | 275 > 109 | 275 > 257 |  | 68 | Sulfadimethoxine/Sulfadoxine | 7.04 |  | 311 > 156 | 311 > 92 |  |
| 54 | Narasin | 8.29 | 787.6 > 431.3 | 787.6 > 531.3 |  | 69 | Sulfanilamide | 5.28 |  | 173 > 156 | 173 > 92 |  |
| 55 | Norfloxactin | 3.68 | 320.1 > 233 | 320.1 > 276.1 |  | 70 | Terbutaline | 1.66 |  | 226.1 > 107 | 226.1 > 125 |  |
| 56 | Novobiocin | 7.41 | 613 > 189 | 613 > 133 |  | 71 | Thiabendazole | 5.35 |  | 202 > 130.9 | 202 > 174.9 |  |
| 57 | Oxacillin | 6.56 | 402.2 > 160 | 402.2 > 243.1 |  | 72 | Tildipirosin | 8.18 |  | 734.6 > 174 | 734.6 > 132 |  |
| No | Compound | RT (min) | Quantifier Trace | Reference Trace |  | No | Compound | RT (min) |  | Quantifier Trace | Reference Trace |  |
| 73 | Vancomycin | 2.99 | 725.8 > 100 | 725.8 > 144.2 |  | 79a | Diethylstilbesterol |  | 6.52 | 267.1 > 237.1 | 267.1 > 251.1 |
| 74 | Virginiamycin | 6.56 | 526.4 > 508.3 | 526.4 > 355.1 |  | 80a | Estradiol |  | 6.24 | 271.2 > 145.1 | 271.1 > 183.1 |
| 75 | Zilpaterol | 1.76 | 262.2 > 185.1 | 262.2 > 202.1 |  | 81a | Estriol |  | 5.53 | 287.2 > 145 | 287.2 > 171 |
| 76a | 17-alpha-ethylnylestradiol | 6.37 | 295.2 > 145.1 | 295.2 > 158.9 |  | 82a | Estrone |  | 6.40 | 269.2 > 145.1 | 269.2 > 269 |
| 77a | betamethasone | 5.87 | 361.2 > 307.2 | 361.2 > 325.2 |  | 83a | Ethyl estradiol |  | 6.37 | 295 > 145 | 295 > 269 |
| 78a | Diclazuril | 6.80 | 405.1 > 334.1 | 407.1 > 336.1 |  | 84a | Nicarbazin |  | 6.47 | 301 > 107.1 | 301 > 137.2 |

Note: a  are compounds that are detected in negative mode

#### Figure 1(a) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 1 ~ 12)

#### Figure 1(b) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 13 ~ 24)

#### Figure 1 (c) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 25 ~ 36)

#### Figure 1(d) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 37 ~ 48)

#### Figure 1(e) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 49 ~ 60)

#### Figure 1 (f) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 61 ~ 72)

#### Figure 1(g) MRM chromatograms of detected compounds in DR-STD at 10ppb (Compound 73 ~ 84)

## Setting up a screening and quantitation method for samples (DR-BLK A, DR-BLK B, DR-SPK A, DR-SPK B)

The standard mixture, DR-STD, was diluted with 10% MeOH into various concentrations: 0.02 ppb, 0.1 ppb, 1 ppb, 10 ppb, 50 ppb and 100 ppb. A calibration curve using the quantitation trace was established for quantitation of compounds that was spiked into the samples. Due to the different response for the different classes of compounds in the mixture, each analyte has different calibration ranges. Table 4 shows the range and linearity of the 84 compounds. Figure 2 shows the calibration curves of the individual analytes in neat solution.

Table 4. Calibration range and linearity of detected compounds in DR-STD

| No | Compound | Linearity | Range (ppb) |  | No | Compound | Linearity | Range (ppb) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 16-beta-hydroxy stanozolol | 0.999 | 0.1 ~ 100 |  | 11 | Ampicillin | 1.000 | 0.02 ~ 100 |  |
| 2 | 17-alpha-methyltestosterone | 0.999 | 10 ~ 100 |  | 12 | Brombuterol | 1.000 | 1 ~ 100 |  |
| 3 | 4-Epichlortetracycline | 0.999 | 0.1 ~ 100 |  | 13 | Buqinolate | 0.999 | 10 ~ 100 |  |
| 4 | 4-Epioxytetracycline | 0.999 | 0.1 ~ 100 |  | 14 | Carbmazepine | 0.999 | 10 ~ 100 |  |
| 5 | 4-Epitetracycline | 0.999 | 10 ~ 100 |  | 15 | Cefalexin | 1.000 | 0.1 ~ 100 |  |
| 6 | 5-Hydroxyflunixin | 0.997 | 10 ~ 100 |  | 16 | Cefazolin | 1.000 | 0.1 ~ 100 |  |
| 7 | 6-alpha-Methylprednisolone | 1.000 | 10 ~ 100 |  | 17 | Cefoperazone | 1.000 | 0.02 ~ 100 |  |
| 8 | Albendazole | 1.000 | 1 ~ 100 |  | 18 | Ceftiofur | 1.000 | 0.02 ~ 100 |  |
| 9 | Altrenogest | 1.000 | 0.02 ~ 10 |  | 19 | Cephalothin | 1.000 | 1 ~ 100 |  |
| 10 | Amoxicillin | 1.000 | 0.02 ~ 100 |  | 20 | Cephradine | 1.000 | 0.02 ~ 100 |  |
| No | Compound | Linearity | Range (ppb) |  | No | Compound | Linearity | Range (ppb) |  |
| 21 | Chlortetracycline | 0.999 | 1 ~ 100 |  | 36 | Epi-testosterone/Testosterone | 1.000 | 10 ~ 100 |  |
| 22 | Cimbuterol | 0.999 | 0.02 ~ 50 |  | 37 B | Eprinomectin | - | 50 ~ 100 |  |
| 23 | Clenbuterol | 0.999 | 0.02 ~ 50 |  | 38 | Flumequine | 0.9978 | 1 ~ 100 |  |
| 24 | Clenpenterol | 0.998 | 0.02 ~ 100 |  | 39 | Furazolidone | 1.000 | 10 ~ 100 |  |
| 25 | Clopidol | 1.000 | 0.1 ~ 10 |  | 40 | Hydoxy dimetridazole | 1.000 | 0.02 ~ 100 |  |
| 26 | Cloxacillin | 0.999 | 0.1 ~ 100 |  | 41 | Ipronidazole | 1.000 | 0.1 ~ 10 |  |
| 27 | Decoquinate | 0.997 | 10 ~ 100 |  | 42 | Ipronidazole hydroxy | 1.000 | 0.02 ~ 100 |  |
| 28 | Desacetylcefapirin | 0.999 | 10 ~ 100 |  | 43 | Lasalocid | 1.000 | 0.02 ~ 10 |  |
| 29 | Dexamethasone | 1.000 | 0.1 ~ 10 |  | 44 | Mabuterol | 1.000 | 1 ~ 100 |  |
| 30 | Dicloxacillin | 1.000 | 1 ~ 100 |  | 45 | Maduramycin | 0.997 | 0.1 ~ 10 |  |
| 31 | Difloxacin | 0.992 | 10 ~ 100 |  | 46 | Mapenterol | 0.998 | 0.02 ~ 50 |  |
| 32 | Dimetridazole | 0.9994 | 0.02 ~ 50 |  | 47 B | Marbofloxacin | - | 50 ~ 100 |  |
| 33 B | Doramectin | - | 50 ~ 100 |  | 48 | Metronidazole | 0.998 | 0.02 ~ 50 |  |
| 34 | Doxycycline/Tetracycline | 0.998 | 1 ~ 100 |  | 49 | Metronidazole-hydroxy | 1.000 | 0.02 ~ 100 |  |
| 35 B | Enrofloxacin | - | 50 ~ 100 |  | 50 | Monensin | 0.999 | 0.1 ~ 10 |  |
| No | Compound | Linearity | Range (ppb) |  | No | Compound | Linearity | Range (ppb) |  |
| 51 | Monepantel sulfone | 1.000 | 0.1 ~ 10 |  | 66 | Sarafloxacin | 0.997 | 10 ~ 100 |  |
| 52 | Nafcillin | 1.000 | 0.02 ~ 100 |  | 67 | Stanozolol | 0.998 | 1 ~ 100 |  |
| 53 | Nandrolone | 0.999 | 0.1 ~ 100 |  | 68 | Sulfadimethoxine/Sulfadoxine | 0.997 | 10 ~ 100 |  |
| 54 | Narasin | 0.995 | 0.02 ~ 10 |  | 69 | Sulfanilamide | 1.000 | 1 ~ 100 |  |
| 55 B | Norfloxactin | - | 50 ~ 100 |  | 70 | Terbutaline | 1.000 | 0.02 ~ 10 |  |
| 56 | Novobiocin | 0.998 | 0.1 ~ 50 |  | 71 | Thiabendazole | 1.000 | 10 ~ 100 |  |
| 57 | Oxacillin | 1.000 | 0.1 ~ 100 |  | 72 | Tildipirosin | 0.996 | 10 ~ 100 |  |
| 58 | Oxibendazole | 1.000 | 10 ~ 100 |  | 73 | Vancomycin | 1.000 | 1 ~ 100 |  |
| 59 | Penicillin G | 1.000 | 0.1 ~ 100 |  | 74 | Virginiamycin | 0.999 | 0.1 ~ 100 |  |
| 60 | Penicillin V | 1.000 | 0.1 ~ 100 |  | 75 | Zilpaterol | 1.000 | 0.1 ~ 10 |  |
| 61 | Prednisolone | 1.000 | 0.02 ~ 100 |  | 76 | 17-alpha-ethylnylestradiol | 0.999 | 0.1 ~ 50 |  |
| 62 | Ractopamine | 0.999 | 0.02 ~ 50 |  | 77 | betamethasone | 1.000 | 0.02 ~ 10 |  |
| 63 | Robenidin | 0.996 | 10 ~ 100 |  | 78 | Diclazuril | 0.999 | 0.1 ~ 10 |  |
| 64 | Ronidazole | 0.999 | 0.02 ~ 50 |  | 79 | Diethylstilbesterol | 1.000 | 0.1 ~ 10 |  |
| 65 | Salbutamol | 1.000 | 0.02 ~ 10 |  | 80 | Estradiol | 0.997 | 10 ~ 100 |  |
| No | Compound | Linearity | Range (ppb) |  | No | Compound | Linearity | Range (ppb) |  |
| 81 | Estriol | 1.000 | 0.1 ~ 10 |  | 83 | Ethyl estradiol | 0.999 | 0.1 ~ 50 |  |
| 82 | Estrone | 1.000 | 0.02 ~ 50 |  | 84 | Nicarbazin | 0.998 | 1 ~ 100 |  |

Note: B refers to compound having only 2 calibration points, thus linearity is not provided.

#### Figure 2(a) Calibration curve of 84 compounds found in DR-STD (Compound 1 ~ 12)

103

743

443

1

113

543

1203

943

843

643

33

2

#### Figure 2(b) Calibration curve of 84 compounds found in DR-STD (Compound 13 ~ 24)

223

1943

1643

13

17

233

203

14

243

213

183

15

#### Figure 2(c) Calibration curve of 84 compounds found in DR-STD (Compound 25 ~ 36)

3143

283

25

343

29

353

323

26

3633

303

333

27

#### Figure 2(d) Calibration curve of 84 compounds found in DR-STD (Compound 37 ~ 48)

4343

403

37

463

41

473

443

38

4833

423

453

39

#### Figure 2(e) Calibration curve of 84 compounds found in DR-STD (Compound 49 ~ 60)

5543

523

49

583

53

593

563

50

6033

543

573

51

#### Figure 2(f) Calibration curve of 84 compounds found in DR-STD (Compound 61 ~ 72)

6743

643

61

703

65

713

683

62

7233

663

693

63

#### Figure 2(g) Calibration curve of 84 compounds found in DR-STD (Compound 73 ~ 84)

7943

763

73

823

77

833

803

74

8433

783

813

75

## Screening and Quantitation result for DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B using established MRM method and calibration curves

Sample DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B were diluted with water (1:1, vol:vol) before LCMSMS injection. Triplicate analysis was performed on each sample to obtain reliable results. For each detected analyte in the sample, they were quantitated based on the calibration curve established with the standards.

### Results for DR-BLK A, DR-BLK B

DR-BLK A and DR-BLK B are blank matrices provided. There are traces of drug residue detected in both blank matrices. 5 analytes were detected and quantitated in DR-BLK A while 4 analytes were detected and quantitated in DR-BLK B. Table 5 shows the drug residue detected and the amount present in the blank matrices.  Figure 3 display the MRM chromatogram of detected analytes in DR-BLK A and Figure 4 display the MRM chromatogram of detected analytes in DR-BLK B

Table 5. List of detected compounds in DR-BLK A and DR-BLK B

| No | Compounds | Conc (ppb) |  |
| --- | --- | --- | --- |
|  |  | DR-BLK A | DR-BLK B |
| 1 | Albendazole | 130.1 | 49.4 |
| 2 | Altrenogest | 0.1 | ND |
| 3 | Furazolidone | 428.1 | 363.1 |
| 4 | Monensin | 0.1 |  |
| 5 | Sulfanilamide | 79.1 | 78.5 |
| 6 | 17-alpha-ethylnylestradiol |  | 0.3 |

#### Figure 3 MRM chromatograms of detected compounds in DR-BLK A

#### Figure 4 MRM chromatograms of detected compounds in DR-BLK B

### Results for DR-SPK A, DR-SPK B

DR-SPK A and DR-SPK B are spiked samples that was provided. A total of 32 analytes were detected in DR-SPK A and 30 analytes were detected in DR-SPK B. Due to the interference from the matrices, fronting and tailing are observed for some of the analytes. Table 6 list down the drug residues detected and the amount present in the spiked matrices. Figure 5 display the MRM chromatogram of detected analytes in DR-SPK A and Figure 6 display the MRM chromatogram of detected analytes in DR-SPK B

Table 6. List of detected compounds and its amount in DR-SPK A and DR-SPK B

| No | Compounds | Conc (ppb) |  |
| --- | --- | --- | --- |
|  |  | DR-SPK A | DR-SPK B |
| 1 | 16-beta-hydroxy stanozolol | 3.02 | 2.67 |
| 2 | 4-epioxytetracycline | 2.07 | 3.00 |
| 3 | Albendazole | 309.19 | 58.22 |
| 4 | Altrenogest | 0.08 | ND |
| 5 | Amoxicillin | 0.51 | 1.03 |
| 6 | Ampicillin | 0.80 | 1.62 |
| 7 | Chlortetracycline | 21.63 | 39.45 |
| 8 | Clenbuterol | 0.22 | 0.22 |
| 9 | Dexamethasone | 6.78 | 6.74 |
| 10 | Dimetridazole | 0.92 | 0.68 |
| 11 | Epi-testosterone/Testosterone | ND | 13.91 |
| 12 | Furazolidone | 752.13 | 773.79 |
| 13 | Hydoxy dimetridazole | 0.35 | 0.35 |
| 14 | Ipronidazole | 1.22 | 1.20 |
| 15 | Maduramycin | 2.51 | 3.75 |
| 16 | Metronidazole | 0.34 | 0.28 |
| 17 | Metronidazole-hydroxy | 0.28 | 0.28 |
| 18 | Monensin | 4.89 | 5.35 |
| 19 | Nandrolone | 2.28 | 2.10 |
| 20 | Novobiocin | 8.65 | 8.48 |
| 21 | Oxibendazole | 23.56 | ND |
| 22 | Penicillin V | 25.94 | 28.15 |
| 23 | Prednisolone | 6.48 | 8.42 |
| 24 | Ractopamine | 0.14 | 0.16 |
| 25 | Salbutamol | 0.12 | 0.12 |
| 26 | Stanozolol | 2.72 | 1.72 |
| 27 | Sulfanilamide | 84.70 | 67.87 |
| 28 | Zilpaterol | 0.12 | 0.0.10 |
| 29 | 17-alpha-ethylnylestradiol | 1.08 | 0.30 |
| 30 | betamethasone | 2.72 | 0.68 |
| 31 | Diethylstilbesterol | 1.06 | 0.30 |
| 32 | Estrone | 0.06 | ND |
| 33 | Ethyl estradiol | 1.04 | 0.24 |

#### Figure 5(a) MRM chromatograms of detected compounds in DR-SPK A

#### Figure 5(b) MRM chromatograms of detected compounds in DR-SPK A

#### Figure 5(c) MRM chromatograms of detected compounds in DR-SPK A

#### Figure 6(a) MRM chromatograms of detected compounds in DR-SPK B

#### Figure 6(b) MRM chromatograms of detected compounds in DR-SPK B

#### Figure 6(c) MRM chromatograms of detected compounds in DR-SPK B

# Conclusions

- Standard DR-STD was screened using an MRM method that was established using Quanpedia. Out of the 250 potential compounds, 84 compounds were identified

- Calibration curves were established for all the detected analytes to further screen and quantitate for samples DR-BLK A, DR-BLK B, DR-SPK A and DR-SPK B. All calibration curves achieve good linearity of >0.995.

- Triplicate analysis of the samples was performed. Total number of compounds identified in the samples are summarized as follow:

  - DR-SPK A: 32 drugs identified

  - DR-SPK B: 30 drugs identified

  - DR-BLK A: 5 drugs identified

  - DR-BLK B: 4 drugs identified

# List of Figures

Page

Figure 1 MRM chromatograms of detected compounds in DR-STD at 10ppb			10 ~ 16

Figure 2 Calibration curve of 84 compounds found in DR-STD				21 ~ 27

Figure 3 MRM chromatograms of detected compounds in DR-BLK A 			29

Figure 4 MRM chromatograms of detected compounds in DR-BLK B				29

Figure 5 MRM chromatograms of detected compounds in DR-SPK A				31 ~ 33

Figure 6 MRM chromatograms of detected compounds in DR-SPK B				34 ~ 36

# List of Tables

Page

Table 1. List of drugs identified in DR-STD (0.1 ppm) 						2

Table 2. List of drugs identified and the amount present in sample 				3		(DR-SPK A, DR-SPK B, DR-BLK A and DR-BLK B)

Table 3. List of detected compounds in DR-STD using MRM method 				6 ~ 9		with 2 MRM transitions

Table 4. Calibration range and linearity of detected compounds in DR-STD 			17 ~ 20

Table 5. List of detected compounds in DR-BLK A and DR-BLK B 				28

Table 6. List of detected compounds and its amount in DR-SPK A and DR-SPK B 		30