Evaluation of liquid chromatography high resolution mass spectrometry (LC-HRMS)

The evaluation will be performed using analysis results of three drug compounds specified by the Illicit Drugs Laboratory (IDL).

Vendors can either

- obtain samples from the IDL for analysis and submit the analysis results; or

- submit analysis results of similar drug compounds which had been acquired on the instrument with similar specifications as the proposed system.

Instructions:

The evaluation will take place from 5 January to 5 February 2021. The results of analyses can be submitted in hard or soft copies by 5 February 2021. For evaluation of the acquisition and data processing functions of the proposed system, the tenderer shall provide a ‘live’ demonstration upon request.

Please refer to Appendix A for further instructions.

Appendix A: List of samples for analysis

The standard solutions in the table below are prepared at a concentration of 0.01 mg/mL. S1 and S2 are prepared in methanol, while S3 and S4 are prepared in water.

| Sample | Name of Drug | Instructions |
| --- | --- | --- |
|  | 5-Chloro-APINACA | - Analyse each of the sample and acquire the total ion chromatograms of these compounds. - Inject each sample 10x consecutively and obtain the accurate mass. - Tabulate retention times as well as the accurate mass and the corresponding mass error of each injection in the excel template which will be provided. - For one of the injection of each sample, acquire the MS2 spectrum. Note: Please indicate if internal and/or external mass calibration had been performed during acquisition. |
|  | 5-Fluoro-cumyl-PEGACLONE |  |
|  | N-Ethylhexedrone hydrochloride |  |
|  | Caffeine | - Analyse the sample 10x consecutively per day over three days and acquire the total ion chromatogram. - Tabulate the accurate mass and corresponding mass error in the excel template provided. Note: Please indicate if internal and/or external mass calibration had been performed during acquisition. |

The following shall be provided by the tenderers as part of the evaluation submission:

- LC and MS conditions used for the analyses.

- Excel sheet with the requested information.

- Relevant chromatograms and/or reports of all sample analyses.

- Using 5-fluoro-cumyl-PEGACLONE (S2), a demonstration of the utility of the proposed data processing software for the elucidation of unknown compounds shall be provided.

Appendix B: Information on compounds

| Sample | Name of Drug | Structure | Molecular formula |
| --- | --- | --- | --- |
|  | 5-Chloro-APINACA |  | C23H30ClN3O |
|  | 5-Fluoro-cumyl-PEGACLONE |  | C25H27FN2O |
|  | N-Ethylhexedrone |  | C14H21NO |
|  | Caffeine |  | C8H10N4O2 |