PART 3

PARTICULARS OF TENDERER

| 1. | General Information |  |  |
| --- | --- | --- | --- |
|  | State whether: Public Limited Co. / Private Limited Co./ Partnership / Sole Proprietorship |  |  |
|  | Registered Name of Company/Firm : |  |  |
|  | Registered Address : |  |  |
|  | Telephone Number : |  |  |
|  | Fax Number : |  |  |
|  | Correspondence Address (if different from registered address) : |  |  |
|  | Registration No.: |  |  |
|  | Date of Incorporation/Registration: |  |  |
|  | No. of years in the relevant industry: |  |  |
|  | GST status of Tenderer under the GST Act: | *Taxable/Non-taxable (delete as necessary) |  |
|  | For taxable Tenderer, state GST Registration No: |  |  |
|  | If not, please indicate whether you will be a taxable person under the GST Act : | Will be/Will not* be a taxable person |  |
|  | Interbank GIRO Account Number, if available: |  |
| 2. | Financial Information [To be completed by private companies only] |  |  |
|  | If body corporate, to state:- |  |  |
|  | Total issued capital: |  |  |
|  | Total shareholders’ equity: |  |  |
|  | If partnership or sole proprietorship, to state capital set aside for business of this nature:- |  |  |
|  | Partnership/Sole Proprietorship (to delete whichever is inapplicable) |  |  |
|  | Total owner’s equity: |  |  |
| 3. | Name and Address of Banker(s): |  |  |
|  | [To be completed by private companies only] |  |  |

| 4. | Shareholders/ Partners [To be completed by private companies only] The Tenderer shall provide full information on:- - the name and address of any person or company which owns, directly or indirectly, at least 50% of the total number of shares in the Tenderer; - the number, percentage and class of shares held by such person or company. Name; Position in Firm/Company; Address; % of Share. |
| --- | --- |
| 5. | Persons Legally Empowered To Act and Sign |  |
|  | Name; Position in Firm/Company; Address. |
| 6. | I am/We are* registered with the Expenditure and Procurement Policies Unit (EPPU)/ Building Construction Authority (BCA)* as follows: |  |
|  | Certificate of Registration No. |  |
|  | Date of Expiry : |  |
|  | Head; Financial Category. |
| 7. | List of Debarments (if any) |  |
|  | S/No.; Name of Authority; Reasons for Debarment; Period when Debarment was Effective. From; To. * delete as necessary |

| 8. | Particulars and Employment History of Professional/ Supervisory/ Technical Staff proposed or available for tendered Project (Foreman, Supervisory Staff, Drivers, etc.): |
| --- | --- |
|  | Name; Qualification; Institution; Year of Award; Relevant Working Experience (with Post Held & Responsibilities). ###### Professional Degree holder or equivalent. ###### Supervisory Diploma holder or equivalent. ###### Technical (e.g. Drivers) Trade Certificate holders. If the space provided above is insufficient, please continue on an extension page setting out the required data in a similar manner. |
| 9. | Projects executed for private organizations or other government agencies over the last 3 years to which reference may be made |
|  | S/No.; Project Title; Client & Address; Value of Work ($); Contract Period (DD/MM/YY); Officer-in-charge. From; To. |
|  | If the space provided above is insufficient, please continue on an extension page setting out the required data in a similar manner |

10.	I declare that the information contained is in the best of my knowledge and belief in all respects true and complete and should there be any false statement I understand that my Firm/Company may be debarred from tendering for any NUS work.

I hereby undertake to inform NUS of any changes of partnership/ directorship/ ownership of my Firm/ Company taking place during the term of the contract.

I confirm that the consent of all persons whose personal data are disclosed in this tender submission have been obtained for NUS’s use for purposes of this tender and consequent contract, if awarded.

Signed/Confirmed byb:-

_________________________		          	_______________________

Name:							Firm/Company Stamp

Designation:

(Please state designation of signatory

in Firm/Company i.e. Manager, Managing

Partner or Managing Director)

Date : ________________

Important Notes:

- All items in these forms must be filled. Any item not applicable should be so stated. Incomplete forms may render the tender liable to be rejected.

- These forms must be signed/confirmed by persons or person legally empowered to act and sign on behalf of the Firm/ Company.

FORM OF TENDER

| To : National University of Singapore 21 Lower Kent Ridge Road Singapore 119077 | Name(s) of Tenderer(s): Address : |
| --- | --- |

SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING, TESTING AND TRAINING OF LIQUID CHROMATOGRAPHY HIGH RESOLUTION MASS SPECTROMETER SYSTEM   TO DEPARTMENT OF FOOD SCIENCE & TECHNOLOGY, FACULTY OF SCIENCE, NATIONAL UNIVERSITY OF SINGAPORE

Tender Ref: FST-T-2020-04

- We hereby offer and undertake on your acceptance of this tender to supply the Goods and/or Services as specified in the Requirement Specifications under Part 2 of your Invitation to Tender and in accordance with the Instructions to Tenderer and Conditions of Contract in Part 1 of your Invitation to Tender for the following price/s in SINGAPORE DOLLARS (exclusive of GST):

- (a)   Contract Price:						 	(S$		);

- (b)   Option Price:  							(S$		);

- as detailed in the Schedule to Form of Tender.

- Our tender is fully consistent with and does not contradict or derogate from anything in Part 1 of your Invitation to Tender or downgrade anything in your Requirement Specifications. You are entitled to reject our tender if it is inconsistent with or contradict or derogate from anything in Part 1 of your Invitation to Tender or downgrade anything in Part 2 of your Invitation to Tender.

- We undertake that as and when required by you, we shall execute a formal agreement in the appropriate form incorporating the Conditions of Contract set out in this Invitation to Tender together with such further terms and conditions, if any, agreed upon between the National University of Singapore (herein referred to as “NUS”) and us. Until the said formal agreement is prepared and executed, this offer together with your written acceptance (either by way of a letter of acceptance or purchase order which may be sent by hand or post or electronically through EPS) thereof, shall constitute a binding agreement between us.

- OUR OFFER IS VALID FOR NINETY (90) DAYS FROM THE CLOSING DATE OF THIS TENDER.

- We agree that as and when requested by NUS, we shall extend the validity of this offer for one or more periods not exceeding in total ninety (90) days.

- We understand that you are not bound to accept the lowest, the whole or any tender you may receive and that you reserve the right to and we agree that you may accept our tender in whole or in part in accordance with Clause 9 of the Instructions to Tenderer.

- We further undertake to give you any further information which you may require.

- We warrant, represent and declare that we are duly authorised to submit, sign this tender, receive instruction, give any information, accept any contract and act for and on behalf of _________________  (Insert Name of firm, company or consortium).

Dated this ______________ day of _________________, 20_____

| Signed/confirmed by : |  |  |
| --- | --- | --- |
| Name of signatory |  | Signature & Company/Firm stamp |
| as |  | of |
| Designation of signatory |  | Name, Address & Tel No. of Company/Firm |
| In the presence of Signature of Witness] Name : Address : |  |  |

NOTICE: This Form duly completed MUST accompany every tender proposal. Any change to its wordings may render the tender liable to DISQUALIFICATION

SCHEDULE TO FORM OF TENDER

- PRICE SCHEDULE

| S/N | Item Description | Qty (a) | Price Per Unit (Incl freight, delivery & Installation) (b) | Price Per Item (a x b) = (c) | Remarks |
| --- | --- | --- | --- | --- | --- |
| 1. | Supply, Delivery, Installation, Commissioning, Testing and Training of Liquid Chromatography High Resolution Mass Spectrometer System | 1 |  |  |  |
| Total Price (Excluding GST) |  |  |  |  |  |

- OPTION PRICE SCHEDULE

| S/N | Item Description | Qty (a) | Price Per Unit (Incl freight, delivery & Installation) (b) | Price Per Item (a x b) = (c) | Remarks |
| --- | --- | --- | --- | --- | --- |
| 1. | 1st Year Comprehensive Preventive Maintenance | 1 Year |  |  |  |
| 2. | 2nd Year Comprehensive Preventive Maintenance |  |  |  |  |
| 3. | 3rd Year Comprehensive Preventive Maintenance | 1 Year |  |  |  |
| 4. | 4th Year Comprehensive Preventive Maintenance | 1 Year |  |  |  |
| 5. | 5th Year Comprehensive Preventive Maintenance | 1 Year |  |  |  |
| Total Price (Excluding GST) |  |  |  |  |  |

- SCHEDULE OF RATES

| IV. STATEMENT OF COMPLIANCE FOR PROPOSED GOODS/SERVICES |  |  |  |  |
| --- | --- | --- | --- | --- |
| S/N | Description of Item | Compliance with Specifications (tick accordingly) |  | Information/ Remarks |
|  |  | Yes | No |  |
|  | Requirement Specifications |  |  |  |
|  | Technical Specifications |  |  |  |
| 2.1 | Solvent Delivery System with Integrated Degasser |  |  |  |
|  | Instrument should be equipped with quaternary or binary mixing solvent delivery system with Integrated Degasser for high accuracy. |  |  |  |
|  | Must support and function in the pH range of 1.0 – 12.5. |  |  |  |
|  | Operating pressure range of the binary pump should be at least from 0- 15,000PSI. |  |  |  |
|  | Flow rate at least up to 2 mL/min with flow precision: ≤ 0.08% RSD or ≤ 0.02 min Standard Deviation (SD), whatever is greater. Flow accuracy: ± 1% or lower. |  |  |  |
|  | Must come with automatic seal wash. |  |  |  |
|  | Equipped with leak sensors for leak safe handling. |  |  |  |
|  | Gradient formation: at least 2 solvents. |  |  |  |
|  | Composition precision and accuracy as low as possible: <0.15% RSD or 0.02 min SD, whatever is greater. Composition accuracy: at least ± 0.35% |  |  |  |
|  | Low delay volumes of less than 375µL. |  |  |  |
| 2.2 | Autosampler with cooler |  |  |  |
|  | The auto sampler must come with flow through injection design with variable injection volume |  |  |  |
|  | The injection linearity is 0.9999 in the range of 0.1 – 100 µL. |  |  |  |
|  | The auto sampler should have an injection volume range of at least 0.1 – 20 µL in 0.1 µl increments without any hardware changes |  |  |  |
|  | Providing cooling and heating in the range at least from 4 ° C - 40 °C, temperature settable in 1 °C increments. |  |  |  |
|  | Lowest Carryover < 0.002% |  |  |  |
|  | Able to accommodate 2mL vial container |  |  |  |
| 2.3 | Column oven |  |  |  |
|  | The temperature range is at least from ambient to 90 ° C. |  |  |  |
|  | High temperature accuracy is ± 0.5° C or better |  |  |  |
|  | High temperature precision is 0.05° C or better |  |  |  |
|  | Able to accommodate at least 2 columns. |  |  |  |
|  | Able to tighten the column without using any tools. |  |  |  |
| 2.4 | High Resolution Quadrupole Time-Of-Flight Mass Spectrometer System |  |  |  |
|  | Minimum mass range: 50-10,000 and above |  |  |  |
|  | The system should preferably have built-in calibrant delivery system for automated introduction of reference masses and carry out real-time internal reference mass correction for MS and MS/MS operation for better mass accuracy. |  |  |  |
|  | The mass analyzer must consist of a quadrupole mass filter, a quadrupole collision cell. |  |  |  |
|  | The mass analyzer must be able to perform full scan MS and MS/MS, data independent acquisition and multiple reaction monitoring (MRM). |  |  |  |
|  | The mass measurement accuracy of the instrument should be better than 1 ppm RMS. |  |  |  |
|  | The sensitivity at positive MS mode with 1 pg injection of reserpine, while maintaining maximum mass resolution shall preferably be at least 500:1 RMS signal-to-noise ratio. |  |  |  |
|  | The sensitivity at positive MS/MS mode with 1 pg injection of reserpine, monitoring the most intense product ions while maintaining maximum mass resolution shall preferably be at least 1,500:1 RMS signal-o-noise ration. |  |  |  |
|  | The sensitivity at negative MS mode with 1 pg injection of chloramphenicol, while maintaining maximum mass resolution shall preferably be at least 500:1 RMS signal-to-noise ratio. |  |  |  |
|  | The sensitivity at negative MS/MS mode with 1 pg injection of chloramphenicol, monitoring the most intense product ions while maintaining maximum mass resolution shall preferably be at least 1,500:1 RMS signal-to-noise ratio. |  |  |  |
|  | Instrument should have low-noise analogue-to-digital (ADC) acquisition system for wide dynamic range accurate mass measurement. |  |  |  |
|  | The tuning solution should preferably be fully integrated with the built-in automated tuning system that enables optimization of ion optics and full mass axis calibration. |  |  |  |
|  | The mass resolution must be at least >30,000 FWHM |  |  |  |
|  | The dynamic range must be equal to or greater than 4 orders of magnitude without loss of mass resolution. |  |  |  |
|  | The spectral acquisition rate at MS mode shall be at least 40 spectra/second and at least 25 spectra/second acquisition rate at MS/MS mode shall be while maintaining a mass resolution unchanged. |  |  |  |
|  | Instrument should preferably have excellent Isotope fidelity with Relative Isotopic Abundance (RIA) error less than <5 % |  |  |  |
| 2.5 | Software and Workstation |  |  |  |
|  | The system must include a Windows®-based data acquisition and editing software package that incorporates a graphical user interface utilizing multi-pane windows for easy data acquisition and analysis |  |  |  |
|  | The software should preferably allow several users to access to process and evaluate their data at the same time. |  |  |  |
|  | One PC is for data acquisition and another for data processing. |  |  |  |
|  | Must be able to provide fully integrated control for both the (ultra) high-performance liquid chromatography (UHPLC) and the hi-res mass spectrometer system |  |  |  |
|  | Has on-line help for instrument control programs as well as user manual in PDF file format |  |  |  |
|  | The system must include a Windows®-based data acquisition and editing software package that incorporates a graphical user interface utilizing multi-pane windows for easy data acquisition and analysis |  |  |  |
|  | It should cover digital instrument control, qualitative and quantitative processing, report creation, self-diagnosis and auto- tune. |  |  |  |
|  | Auto tuning with sensitivity and resolution optimization for both positive ion and negative ion should be standard. |  |  |  |
|  | The proposed solution must be single point data system method capability with full control for HPLC system and HRMS and communication between the two such that if the HPLC or MS experiences a “not ready” condition, error, or leak, the system stops injecting the sample. |  |  |  |
|  | Data acquisition software should feature single point method capacity with full control for HPLC modules at all times during LC run. Software should be of single software program from single manufacturer. |  |  |  |
|  | Software must incorporate molecular feature extraction, molecular formula confirmation and molecular formula generation tools. |  |  |  |
|  | Software should have the capability to determine the molecular formula from accurate mass measurement and isotope fidelity (including both isotope spacing and isotope distribution) information. |  |  |  |
|  | Software should have the ability to predict the molecular formula of fragment ion from accurate mass information at MS/MS level. |  |  |  |
|  | Software should have “Personal Compound Database Library (PCDL) Manager” a user creatable & editable accurate mass retention time database with MS/MS library spectral information. Scoring should be based on monoisotopic mass, isotopic spacing and isotopic ratio distribution for MS search and option to choose reverse & forward score for MS/MS spectral match. |  |  |  |
|  | Software should have “Structure Correlator” capabilities, which can predict and match experimental fragments into proposed structure via a “novel systematic bond breaking approach. |  |  |  |
|  | Software should have capability to: Perform multivariate chemometric analysis and advanced statistical analysis such as post-hoc test for two or more sample conditions. |  |  |  |
|  | Support ICPMS, GC/MS and LC/MS data format or generic file on a single platform. |  |  |  |
|  | Support data visualization through heatmaps, correlation plots. |  |  |  |
|  | Perform pathway analysis using open-source or customizable pathway database on a single platform |  |  |  |
|  | Method file should be capable to save all analysis parameters, data processing parameters and report format. |  |  |  |
|  | The setting parameters of data acquisition and data analysis/ reporting for existing/remaining compounds should not be affected by the addition or deletion of compounds |  |  |  |
|  | Comes with a powerful chemometrics software or omics data processing software with many advanced statistical and visualization tools that allow getting biological insight from complex data faster and easier or omics data. |  |  |  |
| 2.6 | Nitrogen Generator (optional) |  |  |  |
|  | Nitrogen Flow Rate supply is at least 70L/Min |  |  |  |
|  | Operation Outlet Pressure at least 100 psi |  |  |  |
|  | Operates with Oil free Compressor based solution, such that generator can run on its own without the need for an external source of air |  |  |  |
|  | Contain at least 4 Oil Free Compressor and offer Compressor Redundancy. In the event if 1 of the compressor is faulty, the remaining compressor can still produce Nitrogen Gas at 70L/min. |  |  |  |
|  | Compressor must be installed in insulated chamber to minimize noise and vibration with Thermal Management to optimize internal air flow and system ventilation |  |  |  |
|  | Gas will be supplied on demand |  |  |  |
|  | Fitted Touch Screen Full colour user interface with on-boards service scheduling and system diagnostic |  |  |  |
|  | Ability to power down all compressor, reducing overall energy consumption |  |  |  |
|  | Ultra-fast Start up, not more than 30 minutes to operating purity from installation |  |  |  |
|  | Internal Status Check on start-up |  |  |  |
|  | Generator supplied must be tested & approved by LC/MS manufacturer |  |  |  |
|  | Generator must be fitted with wheels for easy mobility |  |  |  |
|  | Noise level <60 dB(A) |  |  |  |
|  | Generator must have declaration of conformity, which declare relates in conformity with the applicable European Commission (EC) Directives, harmonized standards, and other normative requirements |  |  |  |
|  | Generator must be fully compliant with all applicable certification standards and tested to comply with current Electromagnetic Compatibility requirement |  |  |  |
|  | Generator must be tested to delivered Nitrogen of suitable quality for the LCHRMS system. |  |  |  |
| 2.7 | Safety Features |  |  |  |
|  | The instrument should have undergone testing under the Consumer Protection Safety Requirements and bear the Safety Mark for use in Singapore |  |  |  |
| 2.8 | Options |  |  |  |
|  | Tenderers shall quote for the following options in the Schedule to Form of Tender. These options shall be exercisable by NUS within validity period if agreeable. |  |  |  |
|  | - Yearly comprehensive preventive maintenance for the next 5 years |  |  |  |
|  | - Further discount or provision can be offered if working on collaborative mode on Foodomics for food safety, sustainability, functional nutrition and Health |  |  |  |
| 3. | OTHER REQUIREMENTS |  |  |  |
| 3.1 | Utilities Requirement |  |  |  |
| 3.2 | Safety Requirements |  |  |  |
| 3.3 | Factory Acceptance Test (FAT) |  |  |  |
| 3.4 | Delivery |  |  |  |
| 3.5 | Price |  |  |  |
| 3.6 | Export and Import licenses |  |  |  |
| 3.7 | Site Preparation |  |  |  |
| 3.8 | Pre evaluation sample analysis |  |  |  |
| 3.9 | Installation, Commissioning and Training |  |  |  |
| 3.10 | Site Acceptance Test (SAT) |  |  |  |
| 3.11 | Documentation |  |  |  |
| 3.12 | Support and Maintenance |  |  |  |
| 3.13 | Warranty |  |  |  |
| 3.14 | Extended Warranty (Option) |  |  |  |
| 3.15 | Comprehensive Service Contract (Option) |  |  |  |
| 3.16 | Track Record |  |  |  |
| 3.17 | Delivery Timeline |  |  |  |

If the tender is submitted by a consortium, each member of the consortium shall be named.

If the tender is submitted by a consortium, each member of the consortium shall be listed.  For example, “… for and on behalf of (Name of 1st Member) , (Name of 2nd Member)  and  (etc – List out Names of remaining Members) …”

If the tender is submitted by a consortium, the Lead Member of the consortium shall submit the tender on behalf of the consortium members. Documentary proof must be provided that the Lead Member is authorised by all members of the consortium to submit, sign the tender, receive instruction, give any information, accept any contract and act for and on behalf of all the members of the consortium.