## HPLC
HPLC
Takeda requirements |  |  |  | Waters |  | Agilent |  | Thermofisher |  | ThermoFisher
Item | Description | Criteria for HPLC | Requirement | Response (yes/no) | Quotation ( SGD) | Response (yes/no) | Quotation ( SGD) | Response (yes/no) | Quotation ( SGD) | Response (yes/no) | Quotation ( SGD)
1 | Capability | Supply of HPLC
Hardware requirements
|  | Able to interface with plant network and use existing software | must have | yes
|  | Able to support up to 4 solvents per method |  | yes
|  | Capable of flowrates between 0.01 to 5.000 mL/min in 0.001 mL/min increments |  | yes
|  | Operate at pressures up to 600 bar |  | yes
|  | Withstand solvent pH between 1 - 12.3 |  | yes
|  | Solvent Composition range: 0 - 100 % |  | yes
|  | Injection volume range: 0.1 to 100 μL |  | yes
|  | Column heater or Column heater/cooler temperature accuracy test at 25°C, 30°C and 40°C |  | yes
|  | Sample heater/ cooler temperature accuracy test at 4°C |  | yes
Compatible with minimally 3 detectors - PDA/DAD, FLD, RID
PDA / DAD:
|  | Wavelength range: 190 - 950 nm or 190 - 800 nm |  | yes
|  | Wavelength accuracy: ±1 nm |  | yes
|  | Linearity range: Deviation at 2.0 AU ≤ 5% |  | yes
|  | Baseline noise: ± µAU, 230 nm, 2 Hz s, wavelength compensation 310 to 410, WET (2D channel) |  | yes
|  | Drift: ≤ 1.0 x 10^-3/AU/Hour/°C |  | yes
|  | Data acquisition / sampling rate: Up to 80Hz |  | yes
|  | Pressure limit: ≥1000 PSI |  | yes
|  | Operating temperature: 4 to 40°C |  | yes
|  | Operating Humidity: 20 to <95% RH |  | yes
FLD:
|  | Minimally meet requirements outlined in VAL-151669 |  | yes
RID:
|  | Minimally meet requirements outlined in VAL-151669 |  | yes
|  | Operation requirements (software and computerized system) |  | yes
|  | Complies with CFR 24 Part 11 |  | yes
|  | The system must be able to create and store electronic records |  | yes
|  | Operational system checks must exist to enforce sequencing of steps and events |  | yes
|  | Able to interface with plant network and use existing software |  | yes
|  | The system must produce records that uniquely identify the individual that created the record along with the date and time the record was created |  | yes
|  | The system must support electronic signatures |  | yes
|  | The system must have provision to select the sample name and type from a dropdown list of standard and should be entered for the sample name. |  | yes
|  | System must have ability to perform diagnostic checks to ensure that it is functioning correctly upon initialization. |  | yes
|  | There must be a system check in place before operation to determine if the instrument is operating properly |  | yes
|  | The system must have the ability to run a single determination or a series of samples |  | yes
|  | The system must be able to generate, display, and record error, warning, or abort messages when unexpected conditions are encountered. |  | yes
|  | The systems must analyze, and report result when only one of the detectors is used. |  | yes
|  | The Software shall have a minimum of two levels of access. |  | yes
|  | An access level for administrator shall be available with access to the software administrative functions (e.g. create users, reset password etc.) |  | yes
|  | There shall be a user access level for performing user activities/tasks such as measurements and report generation. |  | yes
|  | System must have data checks to determine invalid and altered records. |  | yes
|  | When exchanging data electronically with other systems, the system must include appropriate built-in checks for the correct and secure entry and processing of data. |  | yes
|  | Regular back-ups of all relevant data should be done. Integrity and accuracy of backup data and the ability to restore the data must be checked during validation and monitored periodically |  | yes
|  | Data shall be stored in a suitable location e.g. local hard disk or validated server which can be periodically backup for disaster recovery purpose |  | yes
|  | The system shall encompass data security, access control, data integrity and audit trail features |  | yes
|  | The audit trail shall minimally have the following: i. Be date-time stamped ii. Assigned to specific user iii. Describe activity type (for parameter changes, the audit trail entry shall reflect both the old and new values) iv. Audit trail is secure (unable to change even by an administrator level access) v. The audit trail shall be generated by the software |  | yes
|  | The software shall have reporting capabilities for calibration and measurements |  | yes
Testing and commissioning
|  | Verify system for material of construction, start up of equipment, testing and commissioning to user requirement Provide documents as folllowing: Calibration cert Testing report | must have | yes
| Familiarity with existing system | Knowledge of existing system
|  | Experience of providing similar or same equipment to Takeda, Woodland plant | Good to have | yes
|  | Experience working in Takeda, Woodland plant | Good to have | yes
|  | Experience in Singapore Pharma Projects | Good to have | yes
| Safety Requirement | EHS Requirement
|  | Construction and components used shall comply to local code | must have | yes
|  | Process equipment with a potential for operator interaction, advertent or inadvertent, shall be equipped with proper safety controls such as emergency stops, safety devices, guarding, or interlocks (as applicable to equipment) to prevent injury. Safety controls shall be installed according to approved design. | must have | yes
|  | Device must be designed for safe, ergonomic, and accessible operation | must have | yes
|  | System and all related equipment shall be designed and constructed in accordance with all applicable Shire and Singapore requirements, including Singapore Standards and Codes of Practices. | must have | yes
|  | Equipment noise level lower than 80DB |  | yes
|  | IEC 61010 Safety requirements for electrical equipment for measurement, control and laboratory use. | must have | yes
| Quality & Documentation | Quality & Documentation
|  | Documentation will be provided by the system vendor to support equipment commissioning and qualification. Documentation must be in English language | must have | yes
|  | Instruction manual, detailed manual for installation / startup, Maintenance and transportation | must have | yes
|  | Supplier certificate of quality | must have | yes
|  | Spare parts list | must have | yes
|  | Final test Report | must have | yes
| Schedule | Deliver and commission within 3 months of PO | must have | yes
| Pricing | Pricing
Scoring
|  | Capability | 35%
|  | Familiarity with existing system | 10%
|  | Safety requirement | 10%
|  | Quality & Documentation | 15%
|  | Competitiveness (price) | 30%
100%
Note on C59: Muthukrishnan, Vaidhyanathan: tuas?