## Sheet1
SNo | Current practice | Improvement Required
Material
1 | Approx. 3 -5 % Yield fallouts due to Check valve (289000144) | Redesign Check valve (289000144) due to approx. 3 -5 % Yield fallouts
2 | Leak issues at Pump tray and Final test and approx. 3 to 5 % fallout | Redesign Tubing Assembly (Tubing+ Compression screw + Ferrule) as approx. 3 to 5 % fallout
3 | Not much space to access to connect interface cables between 210000411 PCBA and 210000416 PCBA | Combine 210000411 PCBA and 210000416 PCBA into one and remove interface cables
4 | Current power module is bulky and occupy lot of space. | Redesign current power module into compact power module
Process
1 | Tubing insert and tightening to HPMV 1&2 bottom side is tedious. | Improve the Tubing insert location or change method of tube tightening
2 | Apply Epoxy to hold fiber optic vial sensors to vial brackets and use post it notes till it dries | alternative solution to replace epoxy process
Testing
1 | Current test steps are manual and labor intensive | Automate test steps so they can run one after the other and generate test report automatically after testing complete
2 | Current test cycle time is long 5~6 hours | Reduce test cycle time
3 | Daily verification of CCD carousel measurement check using smart scope is manual and labor intensive | Automate Daily verification of CCD carousel measurement check using smart scope & program
4 | Manually prepare HPLC UV test solutions (4x types) into respective vials as per defined frequency | Pre-order HPLC UV test solutions (4x types) in vial sizes and use them directly in testing
5 | prepare Pre-blank& post blank solution (70-30) | Pre-order Pre-blank& post blank solution (70% Water-30%Methanol) in vial sizes and use them directly in testing
Machine/Testers
1 | All eAlliance testers/Tooling/Machine were old and many of the tester/machine components were obsolete. | Refurbish all eAlliance testers/Tooling/Machine