## Scope
Scope
1. The intent of this evaluation is to compare results of carryover test between Waters SQT carryover test and Waters Alliance Mfg carryover test (WAT270119)
Conditions:
A. Please execute test per test group A and test group B
B. All units must be fresh units without injector change. Do not include units with injectors change in this evaluation
C. All chemistry to be retain in the original vials and not move swap.

## Test Group A
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM4690A | F18SM44699A | F18SM7084A | F18SM4632A | F18SM7125A | F18SM7140A
1 | Perform carryover using manufacturing final test | (Pass/Fail) | Pass, but high blank | Pass, but high blank | Pass, but high blank | Failed, high blank | Failed, high blank | Failed, high blank
2 | Perform carryover using manufacturing final test (2x or above) | (Pass/Fail) | N/A | N/A | N/A | Failed (4 times), high blank | Failed, high blank | Failed, high blank
3 | Rerun same units using SQT carryover test 1st run | (Pass/Fail) | Fail, high blank | Pass, but high blank | Missing, but high blank | Failed 0.002 spec, bad peak shape of post blank 1 | Failed 0.002 spec | Failed 0.002 spec
4 | Rerun same units using SQT carryover test 2nd run | (Pass/Fail) | Missing, but high blank | Pass, but high blank | Missing, but high blank | Passed at 0.002, but bad peak shape/incorrect integration | Failed 0.002 spec | Failed 0.002 spec
5 | Rerun same units using SQT carryover test 3rd run | (Pass/Fail) | Missing, but high blank | N/A | N/A | Failed 0.002 spec, bad peak shape of post blank 1 | Failed 0.002 spec | Failed 0.002 spec
6 | Multiple runs SQT after adjusting seal pack, compression test, chaning new blank, rework Injector | (Pass/Fail) | N/A | N/A | N/A | Failed, either bad peak shape or high carryover. Details in the reports | Failed 0.002 spec in multiple runs (4 times) | Passed, after rework injector and compression test (Passed 3 times)
7 | Re-run TG3&4 and changed new Blank vial, Mfg final test | (Pass/Fail) | Pass | Pass | Failed, small high blank | ? | ? | ?
8 | Re-run TG3&4 and changed new Blank vial, Mfg final test (2x or above) | (Pass/Fail) |  |  | Pass
* Group A: Production Test run first on 6 instruments (followed by 2 times of SQT, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass

## Test Group A Reports
UUT 1 (F18SM4690A)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2 |  |  | SQT Test Run 3 |  |  | Post SQT Mfg Test (Passed)
UUT2(F18SM44699A) |  |  |  |  | Post SQT Mfg Test (Passed)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2
UUT3 (F18SM7084A)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2 |  | Post SQT Mfg Test run1 (Failed) |  |  | Post SQT Mfg Test Run2(Passed)
UUT4 (F18SM4632A)
PreSQT Mfg Test (Failed 5x times) | SQT Test Run 1 (Failed) |  | SQT Test Run 2 (Passed, but bad peak shape/incorrect integration) |  |  | SQT Test Multiple Runs (10x), failing carryover after injector rework |  |  |  | Post SQT Mfg Test Run ?
UUT5 (F18SM7125A)
PreSQT Mfg Test (Failed 2x times) | SQT Test Run 1 (Failed) |  | SQT Test Run 2 (Failed) |  |  | SQT Test Multiple Runs (4 times), failed high carryover |  |  |  | Post SQT Mfg Test Run ?
UUT6 (F18SM7140A)
PreSQT Mfg Test (Failed 2x times) | SQT Test Run 1 (Failed 0.002 spec) |  | SQT Test Run 2 (Failed 0.002spec) |  |  | SQT Test Multiple Runs (3x), failed high carryover |  |  | SQT Test Multiple Runs (3x), Passed |  | Post SQT Mfg Test Run ?

## Test Group B
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM7119A | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA
1 | Run SQT carryover test 1st run | (Pass/Fail) | Missing peak
2 | Run SQT carryover test 2nd run | (Pass/Fail) | Missing peak
3 | Rerun carryover using manufacturing final test | (Pass/Fail) | Failed 0.002 spec
3 | Rerun carryover using manufacturing final test after changing blank and TG3&4 | (Pass/Fail) | Passed
* Group B: The 2 times of SQT run first on 6 instruments (followed by Production Test, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass

## Test Group B Reports
UUT 1 (F18SM7119A)
SQT Test Run1 (Missing peak) | SQT Test Run 2 (Missing Peak) | Mfg Test Run 1 (Failed 0.002 spec) | Mfg Test Run 2 (Passed)

## troubleshooting log
General troubleshooting log for failing units
1 | changing new blank
2 | priming system includes needle wash
3 | run seal pack adjustment
4 | run compression test ( May give a better reading)
5 | rerun TG3_4,
6 | rework injector (most effective)
7 | repeating above steps if still failing