## Scope
Scope
1. The intent of this evaluation is to compare results of carryover test between Waters SQT carryover test and Waters Alliance Mfg carryover test (WAT270119)
Conditions:
A. Please execute test per test group A and test group B
B. All units must be fresh units without injector change. Do not include units with injectors change in this evaluation
C. All chemistry to be retain in the original vials and not move swap.

## Test Group A
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM4690A | F18SM44699A | F18SM7084A | F18SM7165A | F18SM7171A | F18SM4799A
1 | Perform carryover using manufacturing final test | (Pass/Fail) | Pass, but high blank | Pass, but high blank | Pass, but high blank | Pass, but high blank | Pass, but high blank | Pass, but high blank
3 | Rerun same units using SQT carryover test 1st run | (Pass/Fail) | Fail, high blank | Pass, but high blank | Missing, but high blank | Pass, but high blank | Pass, but high blank | Pass, but high blank
4 | Rerun same units using SQT carryover test 2nd run | (Pass/Fail) | Missing, but high blank | Pass, but high blank | Missing, but high blank | Pass, but high blank | Failed, but high blank | in-progress
7 | Perform carryover using manufacturing final test | (Pass/Fail) | Pass | Pass | Failed, small high blank | Pass, but high blank | Failed, high carryover/blank
8 | Perform carryover using manufacturing final test | (Pass/Fail) |  |  | Pass | Pass, but high blank | Failed, high carryover/blank
[6 empty] | New SQT Template from Mike/Josh | New SQT Template from Mike/Josh | New SQT template from Mike/Josh, Tester 6 port 2
* Group A: Production Test run first on 6 instruments (followed by 2 times of SQT, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass

## Test Group A Reports
UUT 1 (F18SM4690A)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2 |  |  | SQT Test Run 3 |  |  |  | Post SQT Mfg Test (Passed)
UUT2(F18SM44699A) |  |  |  |  | Post SQT Mfg Test (Passed)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2
UUT3 (F18SM7084A)
PreSQT Mfg Test | SQT Test Run 1 | SQT Test Run 2 |  | Post SQT Mfg Test run1 (Failed) |  |  | Post SQT Mfg Test Run2(Passed)
UUT4 (F18SM7165A)
PreSQT Mfg Test (High Blank/baseline) | SQT Test Run 1 (Pass but high blank/baseline) |  | SQT Test Run 2 (Pass, but High Blank/baseline) |  |  |  | PostSQT Mfg Final Test Run1 |  | PostSQT Mfg Final Test Run2
UUT5 (F18SM7171A)
PreSQT Mfg Test (Pass but high blank) | SQT Test Run 1 (Pass but high blank/baseline) |  | SQT Test Run 2 (Failed 0.002 and high blank/baseline) |  |  | PostSQT Mfg final test run1 (failed, high carryover/high blank) |  | PostSQT Mfg final test run2 (failed, high carryover/high blank)
UUT6 (F18SM4799A)
PreSQT Mfg Test (Pass but high blank) | SQT Test Run 1

## Test Group B
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM7119A | F18SM7163A | F18SM7164A | F18SM7186A | F18SM4803A | F18SM4XXXA
1 | Run SQT carryover test 1st run | (Pass/Fail) | Missing peak | Passed | Passed | Passed | in-progress
2 | Run SQT carryover test 2nd run | (Pass/Fail) | Missing peak | Passed | Passed | Passed
3 | Rerun carryover using manufacturing final test | (Pass/Fail) | Failed 0.002 spec | Failed 0.002 Spec | Passed | Pass, but high blank
3 | Rerun carryover using manufacturing final test | (Pass/Fail) | Passed | Passed | Failed 0.002 Spec | Passed
3 | Rerun carryover using manufacturing final test | (Pass/Fail) | Passed |  | Passed | Passed
|  |  |  | New SQT from Mike/Josh | New SQT from Mike/Josh | New SQT from Mike/Josh
* Group B: The 2 times of SQT run first on 6 instruments (followed by Production Test, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass

## Test Group B Reports
UUT 1 (F18SM7119A)
SQT Test Run1 (Missing peak) | SQT Test Run 2 (Missing Peak) |  | Mfg Test Run 1 (Failed 0.002 spec) |  | Mfg Test Run 2 (Passed)
UUT 2 (F18SM7163A)
SQT Test Run1 (passed) |  | SQT Test Run 2 (Passed) | Mfg Test Run 1 (Failed 0.002 spec) |  | Mfg Test Run 2 (Pass)
UUT 3 (F18SM7164A)
SQT Test Run1 (Passed) |  | SQT Test Run 2 (Passed) | Mfg Test Run 1 (Passed) | Mfg Test Run 2 (Failed 0.002 spec) |  | Mfg Test Run 3 (Passed)
UUT 4 (F18SM7186A)
SQT Test Run1 (Passed) |  | SQT Test Run 2 (Passed) | Mfg Test Run 1 (Failed 0.002 spec) | Mfg Test Run2(Passed) |  | Mfg Test Run3(Passed)