## Scope
Scope
1. The intent of this evaluation is to compare results of carryover test between Waters SQT carryover test and Waters Alliance Mfg carryover test (WAT270119)
Conditions:
A. Please execute test per test group A and test group B
B. All units must be fresh units without injector change. Do not include units with injectors change in this evaluation
C. All chemistry to be retain in the original vials and not move swap.

## Test Group A
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6 | S/N of UUT #7 | S/N of UUT #8 | S/N of UUT #9
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM4690A | F18SM44699A | F18SM7084A | F18SM4632A | F18SM7125A | F18SM7140A | F18SMxxxA | F18SMxxxxA | F18SMxxxxA
1 | Perform carryover using manufacturing final test PreSQT Run 1 | (Pass/Fail) | Pass, but high blank | Pass, but high blank | Pass, but high blank | Failed, high blank | Failed, high blank | Failed, high blank
2 | Perform carryover using manufacturing final test PreSQT Run 2 | (Pass/Fail) | N/A | N/A | N/A | Failed, high blank | Failed, high blank | Failed, high blank
3 | Rerun same units using SQT carryover test Run 1 | (Pass/Fail) | Fail, high blank | Pass, but high blank | Missing, but high blank | Failed 0.002 spec | Failed 0.002 spec | Failed 0.002 spec
4 | Rerun same units using SQT carryover test Run 2 | (Pass/Fail) | Missing, but high blank | Pass, but high blank | Missing, but high blank | Passed at 0.002 | Failed 0.002 spec | Passed
7 | Perform carryover using manufacturing final test PostSQT | (Pass/Fail) | Pass | Pass | Pass after retest | ? | ? | ?
* Group A: Production Test run first on 6 instruments (followed by 2 times of SQT, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass

## Test Group B
|  |  | S/N of UUT #1 | S/N of UUT #2 | S/N of UUT #3 | S/N of UUT #4 | S/N of UUT #5 | S/N of UUT #6
S/N | Test sequence | Carryover test (Pass/Fail) | F18SM7119A | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA | F18SM4XXXA
1 | Run SQT carryover test Run 1 | (Pass/Fail) | Missing peak
2 | Run SQT carryover test Run 2 | (Pass/Fail) | Missing peak
3 | Rerun carryover using manufacturing final test Run 1 | (Pass/Fail) | Failed 0.002 spec
3 | Rerun carryover using manufacturing final test Run 2 | (Pass/Fail) | Pass
* Group B: The 2 times of SQT run first on 6 instruments (followed by Production Test, on the same instruments)
Manufacuring Final Test
SQT carryover test
Pass