## Pass
Alliance CarryOver Test Monitoring
TSP WAT271129 Rev 33 | [5 empty] | Alliance Test Tesmplate : 1.8
UUT Firmware : v 3.04
Unit | Date | S/N | Test Station (station. Test port) | Carryover % (<0.002%) | Spec | Remarks | Evaluation
1 | 2018-04-06 | C18SM7917A | 4.4 | 0.0003 | 0.002 | Pass on first test |  | High Carryover
2 | 2018-04-06 | D18SM7918A | 5.1 | 0.0003 | 0.002 | Pass on first test |  | Abnormal Peak
3 | 2018-04-07 | C18SM7909A | 2.3 | 0.0004 | 0.002 | Pass on first test
4 | 2018-04-06 | D18SM4964A | 4.2 | 0.0003 | 0.002 | Pass on first test
5 | 2018-04-06 | D18SM4965A | 5.1 | 0.0003 | 0.002 | Pass on first test
6 | 2018-04-07 | D18SM4968A | 3.1 | 0.0004 | 0.002 | Pass on first test
7 | 2018-04-07 | D18SM4974A | 2.2 | 0.0002 | 0.002 | Pass on first test
8 | 2018-04-07 | C18SM4897A | 4.2 | 0.0003 | 0.002 | Pass on first test
9 | 2018-04-09 | D18SM4962A | 5.2 | 0.0012 | 0.002 | Resolve unit
10 | 2018-04-07 | D18SM4966A | 5.4 | 0.0002 | 0.002 | Resolve unit
11 | 2018-04-07 | D18SM4967A | 5.1 | 0.0002 | 0.002 | Pass on first test
12 | 2018-04-07 | D18SM4969A | 3.2 | 0.0005 | 0.002 | Pass on first test
13 | 2018-04-07 | D18SM4970A | 4.4 | 0.0008 | 0.002 | Pass on first test |  |  | Final Test Station | Total | Remarks
14 | 2018-04-09 | D18SM4972A | 1.3 | 0.0003 | 0.002 | Resolve unit |  |  | 1.1 | 13
15 | 2018-04-07 | D18SM4973A | 2.1 | 0.0003 | 0.002 | Pass on first test |  |  | 1.2 | 18
16 | 2018-04-07 | D18SM4975A | 4.1 | 0.0003 | 0.002 | Pass on first test |  |  | 1.3 | 27
17 | 2018-04-09 | D18SM4976A | 3.2 | 0.0003 | 0.002 | Pass on first test |  |  | 1.4 | 27
18 | 2018-04-09 | D18SM4977A | 2.3 | 0.0003 | 0.002 | Pass on first test |  |  | 2.1 | 19
19 | 2018-04-09 | D18SM4978A | 3.1 | 0.0003 | 0.002 | Pass on first test |  |  | 2.2 | 24
20 | 2018-04-09 | D18SM4979A | 4.2 | 0.0002 | 0.002 | Pass on first test |  |  | 2.3 | 21
21 | 2018-04-09 | D18SM4981A | 4.1 | 0.0003 | 0.002 | Pass on first test |  |  | 2.4 | 31
22 | 2018-04-09 | D18SM4983A | 3.4 | 0.0003 | 0.002 | Pass on first test |  |  | 3.1 | 30
23 | 2018-04-06 | C18SM7897A | 1.4 | 0.0004 | 0.002 | Pass on first test |  |  | 3.2 | 33
24 | 2018-04-06 | C18SM7898A | 3.1 | 0.0012 | 0.002 | Resolve unit |  |  | 3.3 | 17
25 | 2018-04-06 | C18SM7903A | 2.1 | 0.0003 | 0.002 | Resolve unit |  |  | 3.4 | 15
26 | 2018-04-06 | C18SM7904A | 1.1 | 0.0004 | 0.002 | Resolve unit |  |  | 4.1 | 22
27 | 2018-04-06 | C18SM7906A | 1.4 | 0.0002 | 0.002 | Resolve unit |  |  | 4.2 | 24
28 | 2018-04-06 | C18SM7908A | 3.2 | 0.0004 | 0.002 | Resolve unit |  |  | 4.3 | 25
29 | 2018-04-09 | C18SM7910A | 2.4 | 0.0003 | 0.002 | Pass on first test |  |  | 4.4 | 20
30 | 2018-04-06 | C18SM7911A | 3.2 | 0.0015 | 0.002 | Pass on first test |  |  | 5.1 | 22
31 | 2018-04-06 | C18SM7912A | 3.3 | 0.0005 | 0.002 | Resolve unit |  |  | 5.2 | 24
32 | 2018-04-06 | C18SM7914A | 4.2 | 0.0005 | 0.002 | Pass on first test |  |  | 5.3 | 19
33 | 2018-04-06 | C18SM7915A | 4.1 | 0.0006 | 0.002 | Pass on first test |  |  | 5.4 | 6
34 | 2018-04-06 | C18SM7916A | 4.3 | 0.0002 | 0.002 | Resolve unit |  |  | Total | 437
35 | 2018-04-09 | C18SM7913A | 5.1 | 0.0003 | 0.002 | Resolve unit
36 | 2018-04-09 | D18SM4982A | 4.4 | 0.0002 | 0.002 | Pass on first test
37 | 2018-04-10 | D18SM4984A | 2.2 | 0.0002 | 0.002 | Pass on first test
38 | 2018-04-10 | D18SM4985A | 4.1 | 0.0002 | 0.002 | Pass on first test
39 | 2018-04-09 | D18SM4986A | 3.1 | 0.0003 | 0.002 | Pass on first test
40 | 2018-04-10 | D18SM4988A | 1.4 | 0.001 | 0.002 | Pass on first test
41 | 2018-04-10 | D18SM4989A | 1.3 | 0.0002 | 0.002 | Pass on first test
42 | 2018-04-10 | D18SM4990A | 2.3 | 0.0002 | 0.002 | Pass on first test
43 | 2018-04-10 | D18SM4991A | 2.4 | 0.0003 | 0.002 | Pass on first test
44 | 2018-04-10 | D18SM4993A | 3.2 | 0.0001 | 0.002 | Pass on first test
45 | 2018-04-11 | C18SM7907A | 4.3 | 0.0012 | 0.002 | Resolve unit
46 | 2018-04-11 | D18SM4015A | 2.4 | 0.0004 | 0.002 | Pass on first test
47 | 2018-04-11 | D18SM4007A | 4.1 | 0.0006 | 0.002 | Pass on first test
48 | 2018-04-11 | D18SM4005A | 5.1 | 0.0002 | 0.002 | Pass on first test
49 | 2018-04-11 | D18SM4004A | 3.2 | 0.0002 | 0.002 | Resolve unit
50 | 2018-04-10 | D18SM4001A | 1.4 | 0.0002 | 0.002 | Pass on first test
51 | 2018-04-11 | D18SM4999A | 1.3 | 0.0005 | 0.002 | Resolve unit
52 | 2018-04-11 | D18SM4998A | 2.4 | 0.0003 | 0.002 | Pass on first test
53 | 2018-04-11 | D18SM4996A | 5.1 | 0.0004 | 0.002 | Pass on first test
54 | 2018-04-10 | D18SM4994A | 4.2 | 0.0002 | 0.002 | Pass on first test
55 | 2018-04-11 | D18SM4980A | 3.4 | 0.0001 | 0.002 | Resolve unit
56 | 2018-04-09 | D18SM4971A | 1.4 | 0.0002 | 0.002 | Resolve unit
57 | 2018-04-07 | D18SM4963A | 5.3 | 0.0008 | 0.002 | Resolve unit
58 | 2018-04-11 | D18SM4857A | 1.3 | 0.0003 | 0.002 | Pass on first test
59 | 2018-04-12 | D18SM4987A | 2.1 | 0.0003 | 0.002 | Pass on first test
60 | 2018-04-12 | D18SM4003A | 1.2 | 0.0004 | 0.002 | Resolve unit
61 | 2018-04-12 | D18SM4006A | 3.2 | 0.0001 | 0.002 | Pass on first test
62 | 2018-04-12 | D18SM4009A | 1.4 | 0.0009 | 0.002 | Pass on first test
63 | 2018-04-12 | D18SM4014A | 2.4 | 0.0004 | 0.002 | Pass on first test
64 | 2018-04-12 | D18SM4016A | 1.3 | 0.0001 | 0.002 | Pass on first test
65 | 2018-04-12 | D18SM4018A | 4.1 | 0.0008 | 0.002 | Pass on first test
66 | 2018-04-12 | D18SM4022A | 3.1 | 0.0001 | 0.002 | Pass on first test
67 | 2018-04-11 | D18SM4992A | 3.1 | NA | 0.002 | Resolve unit
68 | 2018-04-14 | D18SM4008A | 2.1 | 0.0003 | 0.002 | Pass on first test
69 | 2018-04-13 | D18SM4010A | 1.2 | 0.0003 | 0.002 | Pass on first test
70 | 2018-04-13 | D18SM4011A | 2.2 | NA | 0.002 | Pass on first test
71 | 2018-04-13 | D18SM4012A | 1.4 | 0.0019 | 0.002 | Pass on first test
72 | 2018-04-13 | D18SM4025A | 4.3 | NA | 0.002 | Pass on first test
73 | 2018-04-13 | D18SM4027A | 3.4 | 0.0003 | 0.002 | Pass on first test
74 | 2018-04-13 | D18SM4029A | 3.1 | NA | 0.002 | Pass on first test
75 | 2018-04-13 | D18SM4031A | 2.4 | 0.0005 | 0.002 | Pass on first test
76 | 2018-04-13 | D18SM4032A | 2.1 | 0.0002 | 0.002 | Pass on first test
77 | 2018-04-13 | D18SM4033A | 3.2 | 0.0004 | 0.002 | Pass on first test
78 | 2018-04-13 | D18SM4034A | 3.3 | 0 | 0.002 | Pass on first test
79 | 2018-04-13 | D18SM4037A | 4.2 | NA | 0.002 | Pass on first test
80 | 2018-04-13 | D18SM4038A | 4.4 | NA | 0.002 | Pass on first test
81 | 2018-04-14 | D18SM7920A | 3.1 | 0.0004 | 0.002 | Pass on first test
82 | 2018-04-14 | D18SM7921A | 3.2 | NA | 0.002 | Pass on first test
83 | 2018-04-14 | D18SM7922A | 4.3 | NA | 0.002 | Pass on first test
84 | 2018-04-14 | D18SM7925A | 4.4 | 0.0001 | 0.002 | Pass on first test
85 | 2018-04-14 | D18SM7926A | 1.4 | 0.0002 | 0.002 | Pass on first test
86 | 2018-04-14 | D18SM4021A | 2.3 | 0.0003 | 0.002 | Resolve unit
87 | 2018-04-13 | D18SM4036A | 4.1 | NA | 0.002 | Pass on first test
88 | 2018-04-14 | D18SM4039A | 2.2 | NA | 0.002 | Pass on first test
89 | 2018-04-13 | D18SM4040A | 2.1 | 0.0004 | 0.002 | Pass on first test
90 | 2018-04-14 | D18SM4041A | 2.4 | 0.0003 | 0.002 | Pass on first test
91 | 2018-04-14 | D18SM4043A | 2.4 | 0.0001 | 0.002 | Pass on first test
92 | 2018-04-14 | D18SM4044A | 1.4 | 0.0005 | 0.002 | Pass on first test
93 | 2018-04-12 | C18SM7905A | 3.3 | NA | 0.002 | Resolve unit
94 | 2018-04-16 | D18SM7928A | 1.3 | 0.0002 | 0.002 | Pass on first test
95 | 2018-04-16 | D18SM7930A | 3.2 | 0.0001 | 0.002 | Pass on first test
96 | 2018-04-16 | D18SM7933A | 4.1 | NA | 0.002 | Resolve unit
97 | 2018-04-16 | D18SM7936A | 2.2 | NA | 0.002 | Pass on first test
98 | 2018-04-16 | D18SM7937A | 1.4 | NA | 0.002 | Pass on first test
99 | 2018-04-16 | D18SM7938A | 4.3 | 0.0006 | 0.002 | Pass on first test
100 | 2018-04-16 | D18SM7939A | 3.2 | 0.0003 | 0.002 | Pass on first test
101 | 2018-04-16 | D18SM7941A | 1.3 | 0.0006 | 0.002 | Pass on first test
102 | 2018-04-16 | D18SM4995A | 5.2 | NA | 0.002 | Pass on first test
103 | 2018-04-16 | D18SM4997A | 5.2 | NA | 0.002 | Pass on first test
104 | 2018-04-10 | D18SM4000A | 5.2 | NA | 0.002 | Pass on first test
105 | 2018-04-16 | D18SM40002 | 5.3 | 0.0013 | 0.002 | Pass on first test
106 | 2018-04-16 | D18SM4013A | 5.3 | 0.0017 | 0.002 | Pass on first test
107 | 2018-04-16 | D18SM4017A | 5.2 | 0.0003 | 0.002 | Pass on first test
108 | 2018-04-16 | D18SM4019A | 5.2 | 0.0003 | 0.002 | Pass on first test
109 | 2018-04-16 | D18SM4020A | 5.3 | 0.0013 | 0.002 | Pass on first test
110 | 2018-04-16 | D18SM4023A | 5.3 | 0.0013 | 0.002 | Pass on first test
111 | 2018-04-16 | D18SM4026A | 5.2 | NA | 0.002 | Pass on first test
112 | 2018-04-16 | D18SM4028A | 5.1 | 0.0003 | 0.002 | Pass on first test
113 | 2018-04-16 | D18SM4030A | 5.3 | 0.0012 | 0.002 | Pass on first test
114 | 2018-04-14 | D18SM7923A | 4.2 | 0.0004 | 0.002 | Resolve unit
115 | 2018-04-16 | D18SM7924A | 1.2 | NA | 0.002 | Resolve unit
116 | 2018-04-14 | D18SM7927A | 2.1 | 0.0002 | 0.002 | Resolve unit
117 | 2018-04-16 | D18SM7929A | 2.4 | 0.0006 | 0.002 | Resolve unit
118 | 2018-04-16 | D18SM7931A | 4.4 | NA | 0.002 | Resolve unit
119 | 2018-04-16 | D18SM7935A | 3.4 | 0 | 0.002 | Resolve unit
120 | 2018-04-16 | D18SM7940A | 1.4 | NA | 0.002 | Resolve unit
121 | 2018-04-16 | D18SM7942A | 5.2 | 0.0017 | 0.002 | Pass on first test
122 | 2018-04-16 | D18SM7944A | 5.3 | 0.0014 | 0.002 | Resolve unit
123 | 2018-04-16 | D18SM7946A | 2.2 | NA | 0.002 | Pass on first test
124 | 2018-04-16 | D18SM7024A | 5.4 | 0.0007 | 0.002 | Pass on first test
125 | 2018-04-16 | D18SM4035A | 5.1 | 0.0003 | 0.002 | Pass on first test
126 | 2018-04-16 | D18SM4045A | 4.3 | 0.0004 | 0.002 | Pass on first test
127 | 2018-04-16 | D18SM4046A | 4.3 | 0.0002 | 0.002 | Pass on first test
128 | 2018-04-17 | D18SM4048A | 2.2 | 0.0013 | 0.002 | Resolve unit
129 | 2018-04-17 | D18SM4051A | 3.3 | 0.0011 | 0.002 | Pass on first test
130 | 2018-04-17 | D18SM4052A | 2.1 | 0.0002 | 0.002 | Pass on first test
131 | 2018-04-17 | D18SM4054A | 4.4 | 0.0008 | 0.002 | Pass on first test
132 | 2018-04-18 | D18SM7919A | 3.4 | 0 | 0.002 | Resolve unit
133 | 2018-04-17 | D18SM7934A | 2.3 | 0.0006 | 0.002 | Resolve unit
134 | 2018-04-18 | D18SM7943A | 3.1 | 0.0002 | 0.002 | Resolve unit
135 | 2018-04-18 | D18SM7945A | 1.4 | 0.0002 | 0.002 | Resolve unit
136 | 2018-04-19 | D18SM4047A | 5.1 | NA | 0.002 | Resolve unit
137 | 2018-04-18 | D18SM4049A | 4.1 | 0.0001 | 0.002 | Resolve unit
138 | 2018-04-17 | D18SM4050A | 3.2 | NA | 0.002 | Resolve unit
139 | 2018-04-18 | D18SM4055A | 5.2 | NA | 0.002 | Pass on first test
140 | 2018-04-18 | D18SM4056A | 2.4 | 0.0002 | 0.002 | Resolve unit
141 | 2018-04-18 | D18SM4060A | 4.3 | 0.0004 | 0.002 | Pass on first test
142 | 2018-04-18 | D18SM4061A | 2.2 | 0.0006 | 0.002 | Pass on first test
143 | 2018-04-18 | D18SM4063A | 3.2 | NA | 0.002 | Resolve unit
144 | 2018-04-18 | D18SM4064A | 1.1 | 0.0002 | 0.002 | Pass on first test
145 | 2018-04-18 | D18SM4065A | 2.3 | 0.0017 | 0.002 | Resolve unit
146 | 2018-04-18 | D18SM4066A | 1.2 | NA | 0.002 | Pass on first test
147 | 2018-04-18 | D18SM4067A | 4.1 | NA | 0.002 | Pass on first test
148 | 2018-04-18 | D18SM4068A | 3.3 | 0.001 | 0.002 | Pass on first test
149 | 2018-04-18 | D18SM4071A | 1.4 | 0.0015 | 0.002 | Pass on first test
150 | 2018-04-19 | D18SM4053A | 4.2 | NA | 0.002 | Resolve unit
151 | 2018-04-20 | D18SM4057A | 2.1 | 0.0015 | 0.002 | Resolve unit
152 | 2018-04-19 | D18SM4059A | 4.4 | 0.0013 | 0.002 | Resolve unit
153 | 2018-04-19 | D18SM4069A | 5.2 | NA | 0.002 | Resolve unit
154 | 2018-04-19 | D18SM4070A | 2.4 | NA | 0.002 | Pass on first test
155 | 2018-04-19 | D18SM4073A | 3.1 | 0.0003 | 0.002 | Pass on first test
156 | 2018-04-19 | D18SM4075A | 2.3 | NA | 0.002 | Resolve unit
157 | 2018-04-19 | D18SM4076A | 3.2 | 0.0001 | 0.002 | Resolve unit
158 | 2018-04-19 | D18SM4077A | 3.3 | 0.0011 | 0.002 | Pass on first test
159 | 2018-04-19 | D18SM4078A | 3.4 | NA | 0.002 | Resolve unit
160 | 2018-04-19 | D18SM4080A | 4.3 | NA | 0.002 | Pass on first test
161 | 2018-04-19 | D18SM4081A | 3.1 | 0.0007 | 0.002 | Resolve unit
162 | 2018-04-19 | D18SM4085A | 4.2 | NA | 0.002 | Pass on first test
163 | 2018-04-19 | D18SM4087A | 3.2 | 0.0007 | 0.002 | Pass on first test
164 | 2018-04-20 | D18SM4042A | 1.3 | 0.0003 | 0.002 | Resolve unit | ASP execution eval
165 | 2018-04-20 | D18SM4072A | 1.2 | NA | 0.002 | Pass on first test | ASP execution eval
166 | 2018-04-20 | D18SM4074A | 2.2 | NA | 0.002 | Resolve unit
167 | 2018-04-21 | D18SM4079A | 4.1 | 0.0001 | 0.002 | Resolve unit
168 | 2018-04-20 | D18SM4082A | 5.1 | 0.0004 | 0.002 | Pass on first test
169 | 2018-04-20 | D18SM4084A | 2.4 | 0.0002 | 0.002 | Resolve unit
170 | 2018-04-20 | D18SM4088A | 3.1 | 0.0004 | 0.002 | Pass on first test
171 | 2018-04-20 | D18SM4090A | 4.2 | 0.0015 | 0.002 | Resolve unit
172 | 2018-04-20 | D18SM4091A | 5.2 | 0.0003 | 0.002 | Pass on first test
173 | 2018-04-20 | D18SM4092A | 2.3 | 0 | 0.002 | Resolve unit
174 | 2018-04-20 | D18SM4093A | 3.2 | 0.0003 | 0.002 | Pass on first test
175 | 2018-04-21 | D18SM4094A | 3.4 | NA | 0.002 | Resolve unit
176 | 2018-04-20 | D18SM4098A | 2.1 | 0.0003 | 0.002 | Pass on first test
177 | 2018-04-20 | D18SM4099A | 1.3 | 0.0005 | 0.002 | Pass on first test | ASP execution eval
178 | 2018-04-20 | D18SM4100A | 3.1 | NA | 0.002 | Pass on first test | ASP execution eval
179 | 2018-04-20 | D18SM4102A | 3.2 | 0.0002 | 0.002 | Pass on first test | ASP execution eval
180 | 2018-04-21 | D18SM4104A | 5.1 | 0.0003 | 0.002 | Pass on first test | ASP execution eval
181 | 2018-04-21 | D18SM4106A | 2.1 | 0.001 | 0.002 | Pass on first test | ASP execution eval
182 | 2018-04-23 | D18SM4086A | 4.4 | NA | 0.002 | Pass on first test
183 | 2018-04-23 | D18SM4089A | 3.3 | 0.0003 | 0.002 | Resolve unit
184 | 2018-04-20 | D18SM4095A | 5.2 | NA | 0.002 | Pass on first test
185 | 2018-04-23 | D18SM4096A | 3.2 | 0.0005 | 0.002 | Pass on first test
186 | 2018-04-23 | D18SM4097A | 1.2 | NA | 0.002 | Pass on first test | ASP execution eval
187 | 2018-04-23 | D18SM4103A | 2.4 | NA | 0.002 | Pass on first test
188 | 2018-04-23 | D18SM4105A | 2.2 | NA | 0.002 | Pass on first test | ASP execution eval
189 | 2018-04-23 | D18SM4107A | 2.1 | 0.0008 | 0.002 | Pass on first test | ASP execution eval
190 | 2018-04-23 | D18SM4111A | 1.3 | 0.0002 | 0.002 | Pass on first test
191 | 2018-04-23 | D18SM4112A | 2.1 | 0.0002 | 0.002 | Pass on first test
192 | 2018-04-23 | D18SM4116A | 5.1 | 0.0008 | 0.002 | Pass on first test | ASP execution eval
193 | 2018-04-24 | D18SM7947A | 3.4 | 0.0006 | 0.002 | Resolve unit
194 | 2018-04-24 | D18SM7952A | 3.1 | 0.0002 | 0.002 | Pass on first test
195 | 2018-04-18 | C18SM7896A | 4.3 | NA | 0.002 | Pass on first test
196 | 2018-04-24 | C18SM7900A | 4.3 | NA | 0.002 | Resolve unit | ASP execution eval
197 | 2018-04-23 | C18SM7901A | 4.3 | 0.0011 | 0.002 | Pass on first test
198 | 2018-04-17 | C18SM7902A | 2.2 | NA | 0.002 | Pass on first test
199 | 2018-04-24 | D18SM4058A | 4.2 | 0.0018 | 0.002 | Pass on first test
200 | 2018-04-24 | D18SM4101A | 2.4 | 0.0001 | 0.002 | Resolve unit
201 | 2018-04-24 | D18SM4108A | 3.1 | 0.0007 | 0.002 | Resolve unit | ASP execution eval
202 | 2018-04-24 | D18SM4110A | 4.2 | 0.001 | 0.002 | Resolve unit
203 | 2018-04-24 | D18SM4115A | 2.2 | 0.0006 | 0.002 | Pass on first test | ASP execution eval
204 | 2018-04-24 | D18SM4118A | 4.1 | 0.0001 | 0.002 | Pass on first test | ASP execution eval
205 | 2018-04-24 | D18SM4119A | 1.3 | 0.0016 | 0.002 | Resolve unit | ASP execution eval
206 | 2018-04-24 | D18SM4120A | 2.3 | 0.0002 | 0.002 | Resolve unit | ASP execution eval
207 | 2018-04-24 | D18SM4121A | 3.4 | 0.0004 | 0.002 | Pass on first test | ASP execution eval
208 | 2018-04-24 | D18SM4122A | 3.2 | 0.0005 | 0.002 | Pass on first test | ASP execution eval
209 | 2018-04-25 | D18SM4083A | 1.2 | NA | 0.002 | Resolve unit | ASP execution eval
210 | 2018-04-25 | D18SM4113A | 4.4 | 0.0013 | 0.002 | Resolve unit | ASP execution eval
211 | 2018-04-25 | D18SM4117A | 5.2 | NA | 0.002 | Resolve unit | ASP execution eval
212 | 2018-04-25 | D18SM4123A | 1.4 | 0.0011 | 0.002 | Resolve unit
213 | 2018-04-25 | D18SM4125A | 2.2 | 0.0003 | 0.002 | Pass on first test
214 | 2018-04-26 | D18SM4126A | 4.1 | 0.0003 | 0.002 | Resolve unit
215 | 2018-04-25 | D18SM4127A | 5.3 | 0.0003 | 0.002 | Resolve unit
216 | 2018-04-25 | D18SM4128A | 1.4 | NA | 0.002 | Pass on first test
217 | 2018-04-25 | D18SM4129A | 5.4 | 0.0005 | 0.002 | Pass on first test
218 | 2018-04-25 | D18SM4131A | 1.3 | 0.0002 | 0.002 | Pass on first test
219 | 2018-04-25 | D18SM4132A | 2.2 | 0.0014 | 0.002 | Pass on first test
220 | 2018-04-26 | D18SM4133A | 3.1 | 0.0009 | 0.002 | Pass on first test
221 | 2018-04-26 | D18SM4134A | 1.4 | 0.0001 | 0.002 | Resolve unit
222 | 2018-04-26 | D18SM4136A | 4.3 | 0.0002 | 0.002 | Pass on first test
223 | 2018-04-26 | D18SM4137A | 1.2 | NA | 0.002 | Resolve unit
224 | 2018-04-26 | D18SM4140A | 1.3 | 0.0006 | 0.002 | Resolve unit
225 | 2018-04-25 | D18SM7948A | 3.2 | 0.0004 | 0.002 | Pass on first test
226 | 2018-04-25 | D18SM7949A | 3.1 | 0.0011 | 0.002 | Pass on first test
227 | 2018-04-25 | D18SM7950A | 2.3 | 0.0003 | 0.002 | Pass on first test
228 | 2018-04-25 | D18SM7951A | 1.3 | 0.001 | 0.002 | Pass on first test
229 | 2018-04-26 | D18SM7953A | 2.4 | 0.001 | 0.002 | Resolve unit
230 | 2018-04-25 | D18SM7955A | 4.3 | NA | 0.002 | Pass on first test
231 | 2018-04-26 | D18SM7956A | 3.4 | 0.001 | 0.002 | Resolve unit
232 | 2018-04-26 | D18SM4109A | 3.2 | 0.0018 | 0.002 | Resolve unit | ASP execution eval
233 | 2018-04-26 | D18SM4114A | 2.1 | NA | 0.002 | Resolve unit | ASP execution eval
234 | 2018-04-26 | D18SM4124A | 5.1 | 0.0008 | 0.002 | Resolve unit
235 | 2018-04-26 | D18SM4141A | 1.3 | 0.0001 | 0.002 | Pass on first test
236 | 2018-04-27 | D18SM4146A | 5.3 | 0.002 | 0.002 | Resolve unit
237 | 2018-04-26 | D18SM4147A | 3.4 | 0.0001 | 0.002 | Pass on first test
238 | 2018-04-26 | D18SM4149A | 3.1 | 0.0012 | 0.002 | Pass on first test
239 | 2018-04-26 | D18SM4150A | 1.2 | 0.0002 | 0.002 | Pass on first test
240 | 2018-04-27 | D18SM4155A | 4.2 | 0.0005 | 0.002 | Resolve unit
241 | 2018-04-26 | D18SM4142A | 2.2 | 0.0018 | 0.002 | Resolve unit
242 | 2018-04-26 | D18SM4143A | 1.4 | 0.001 | 0.002 | Resolve unit
243 | 2018-04-26 | D18SM4144A | 4.3 | 0.0003 | 0.002 | Pass on first test
244 | 2018-04-26 | D18SM4135A | 5.4 | 0.0004 | 0.002 | Resolve unit
245 | 2018-04-27 | D18SM4145A | 4.1 | 0.0009 | 0.002 | Resolve unit
246 | 2018-04-27 | D18SM4148A | 2.4 | 0.0009 | 0.002 | Resolve unit
247 | 2018-04-26 | D18SM4151A | 4.4 | 0.0019 | 0.002 | Pass on first test
248 | 2018-04-27 | D18SM4152A | 5.2 | NA | 0.002 | Resolve unit
249 | 2018-04-27 | D18SM4154A | 4.3 | NA | 0.002 | Resolve unit
250 | 2018-04-27 | D18SM4156A | 1.2 | 0.0004 | 0.002 | Pass on first test
251 | 2018-04-27 | D18SM4158A | 1.4 | 0.0003 | 0.002 | Pass on first test
252 | 2018-04-27 | D18SM4159A | 2.2 | 0.0013 | 0.002 | Pass on first test
253 | 2018-04-27 | D18SM4160A | 3.1 | 0.0002 | 0.002 | Resolve unit
254 | 2018-04-27 | D18SM4165A | 5.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
255 | 2018-04-28 | D18SM4166A | 5.2 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
256 | 2018-04-28 | D18SM4167A | 5.3 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
257 | 2018-04-28 | D18SM4169A | 1.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
258 | 2018-04-28 | D18SM4170A | 1.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
259 | 2018-04-28 | D18SM4171A | 2.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
260 | 2018-04-27 | D18SM4162A | 1.1 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
261 | 2018-04-27 | D18SM4138A | 3.3 | 0.0008 | 0.002 | Resolve unit
262 | 2018-04-26 | D18SM7954A | 4.2 | 0.0012 | 0.002 | Resolve unit
263 | 2018-04-28 | D18SM7957A | 2.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
264 | 2018-04-28 | D18SM7959A | 2.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
265 | 2018-04-28 | D18SM7960A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
266 | 2018-04-28 | D18SM7962A | 2.4 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
267 | 2018-04-28 | D18SM7963A | 5.1 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
268 | 2018-04-28 | D18SM7965A | 5.3 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
269 | 2018-04-28 | D18SM7966A | 1.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
270 | 2018-04-30 | D18SM7964A | 5.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
271 | 2018-04-30 | D18SM7967A | 2.2 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
272 | 2018-04-30 | D18SM7968A | 2.1 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
273 | 2018-04-30 | D18SM7969A | 1.2 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
274 | 2018-04-30 | D18SM7970A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
275 | 2018-04-30 | D18SM7971A | 3.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
276 | 2018-04-30 | D18SM7972A | 1.3 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
277 | 2018-04-30 | D18SM7974A | 5.1 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
278 | 2018-04-25 | D18SM4130A | 2.3 | 0.0003 | 0.002 | Pass on first test
279 | 2018-04-30 | D18SM4139A | 2.3 | 0.0012 | 0.002 | Resolve unit
280 | 2018-04-30 | D18SM4157A | 1.1 | 0.001 | 0.002 | Pass on first test
281 | 2018-04-30 | D18SM4161A | 3.2 | 0.0002 | 0.002 | Resolve unit
282 | 2018-04-30 | D18SM4163A | 4.2 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
283 | 2018-04-30 | D18SM4164A | 4.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
284 | 2018-04-30 | D18SM4172A | 5.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
285 | 2018-04-30 | D18SM4173A | 2.4 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
286 | 2018-04-30 | D18SM4174A | 3.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
287 | 2018-04-30 | D18SM4175A | 2.3 | 0.0018 | 0.002 | Resolve unit | Double Press Seal Eval
288 | 2018-04-30 | D18SM4177A | 3.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
289 | 2018-05-02 | D18SM4178A | 4.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
290 | 2018-04-30 | D18SM4182A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
291 | 2018-05-02 | D18SM4176A | 2.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
292 | 2018-05-02 | D18SM4179A | 3.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
293 | 2018-05-02 | D18SM4181A | 4.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
294 | 2018-05-02 | D18SM4183A | 2.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
295 | 2018-05-02 | D18SM4184A | 1.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
296 | 2018-05-02 | D18SM4185A | 1.2 | 0.0017 | 0.002 | Resolve unit | Double Press Seal Eval
297 | 2018-05-02 | D18SM4186A | 1.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
298 | 2018-05-02 | D18SM4187A | 3.3 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
299 | 2018-05-02 | D18SM4188A | 2.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
300 | 2018-05-03 | D18SM4189A | 2.2 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
301 | 2018-05-02 | D18SM4190A | 5.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
302 | 2018-05-02 | D18SM4191A | 4.4 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
303 | 2018-05-02 | D18SM4192A | 4.2 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
304 | 2018-05-02 | D18SM4194A | 4.3 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
305 | 2018-05-03 | D18SM4196A | 1.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
306 | 2018-05-03 | D18SM4197A | 4.1 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
307 | 2018-05-03 | D18SM4199A | 3.3 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
308 | 2018-05-02 | D18SM4201A | 3.2 | 0.0014 | 0.002 | Resolve unit | Double Press Seal Eval
309 | 2018-05-02 | D18SM4205A | 1.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
310 | 2018-05-03 | D18SM4168A | 5.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
311 | 2018-05-03 | D18SM4195A | 2.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
312 | 2018-05-03 | D18SM4198A | 1.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
313 | 2018-05-03 | D18SM4200A | 4.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
314 | 2018-05-03 | D18SM4202A | 3.2 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
315 | 2018-05-03 | D18SM4203A | 1.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
316 | 2018-05-03 | D18SM4204A | 5.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
317 | 2018-05-02 | D18SM4206A | 2.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
318 | 2018-05-03 | D18SM4208A | 4.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
319 | 2018-05-03 | D18SM4209A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
320 | 2018-05-03 | D18SM4210A | 5.1 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
321 | 2018-05-03 | D18SM4211A | 4.1 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
322 | 2018-05-03 | D18SM4212A | 1.1 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
323 | 2018-05-03 | D18SM4213A | 2.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
324 | 2018-05-03 | D18SM4214A | 1.3 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
325 | 2018-05-03 | D18SM4215A | 2.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
326 | 2018-05-03 | D18SM4217A | 2.4 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
327 | 2018-05-03 | D18SM4219A | 1.4 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
328 | 2018-05-03 | D18SM4220A | 3.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
329 | 2018-04-30 | D18SM7958A | 3.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
330 | 2018-04-28 | D18SM7961A | 4.1 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
331 | 2018-05-02 | D18SM7973A | 5.3 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
332 | 2018-05-04 | E18SM7975A | 5.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
333 | 2018-05-04 | E18SM7976A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
334 | 2018-05-04 | E18SM7977A | 2.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
335 | 2018-05-04 | E18SM7978A | 1.4 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
336 | 2018-05-04 | E18SM7979A | 1.3 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
337 | 2018-05-04 | E18SM7981A | 4.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
338 | 2018-05-04 | E18SM7982A | 1.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
339 | 2018-05-04 | E18SM7983A | 1.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
340 | 2018-05-04 | E18SM7984A | 4.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
341 | 2018-05-05 | E18SM7985A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
342 | 2018-05-04 | E18SM7986A | 2.4 | NA | 0.002 | Pass on first test | Double Press Seal Eval
343 | 2018-05-04 | E18SM7987A | 1.3 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
344 | 2018-05-05 | E18SM7989A | 2.4 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
345 | 2018-05-03 | E18SM4207A | 4.3 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
346 | 2018-05-03 | E18SM4209A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
347 | 2018-05-03 | E18SM4210A | 5.1 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
348 | 2018-05-04 | D18SM4162A | 3.4 | NA | 0.002 | Pass on first test | Double Press Seal Eval
349 | 2018-04-30 | D18SM4180A | 5.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
350 | 2018-05-04 | E18SM4216A | 3.3 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
351 | 2018-05-03 | E18SM4218A | 3.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
352 | 2018-05-04 | E18SM4222A | 4.1 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
353 | 2018-05-04 | E18SM4223A | 4.4 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
354 | 2018-05-05 | E18SM4225A | 5.1 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
355 | 2018-05-04 | E18SM4226A | 2.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
356 | 2018-05-05 | E18SM$227A | 4.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
357 | 2018-05-05 | E18SM4228A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
358 | 2018-05-05 | E18SM4229A | 4.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
359 | 2018-05-05 | E18SM4230A | 3.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
360 | 2018-05-05 | E18SM4234A | 3.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
361 | 2018-05-05 | E18SM$235A | 1.4 | NA | 0.002 | Resolve unit | Double Press Seal Eval
362 | 2018-05-05 | E18SM4238A | 4.1 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
363 | 2018-05-07 | E18SM4232A | 2.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
364 | 2018-05-07 | E18SM4233A | 5.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
365 | 2018-05-07 | E18SM4236A | 1.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
366 | 2018-05-07 | E18SM4237A | 1.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
367 | 2018-05-07 | E18SM4239A | 1.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
368 | 2018-05-07 | E18SM4240A | 5.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
369 | 2018-05-07 | E18SM4241A | 4.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
370 | 2018-05-07 | E18SM4242A | 4.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
371 | 2018-05-07 | E18SM4243A | 3.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
372 | 2018-05-07 | E18SM4244A | 1.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
373 | 2018-05-07 | E18SM4245A | 4.4 | NA | 0.002 | Pass on first test | Double Press Seal Eval
374 | 2018-05-07 | E18SM4246A | 3.2 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
375 | 2018-05-07 | E18SM4247A | 3.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
376 | 2018-05-07 | E18SM4248A | 3.4 | 0.0018 | 0.002 | Resolve unit | Double Press Seal Eval
377 | 2018-05-07 | E18SM4249A | 2.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
378 | 2018-05-07 | E18SM4250A | 1.4 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
379 | 2018-05-07 | E18SM4251A | 4.1 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
380 | 2018-05-07 | E18SM4252A | 5.2 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
381 | 2018-05-07 | E18SM4253A | 5.3 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
382 | 2018-05-07 | E18SM4254A | 4.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
383 | 2018-05-07 | E18SM4256A | 4.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
384 | 2018-05-07 | E18SM4257A | 3.1 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
385 | 2018-05-07 | E18SM4260A | 3.3 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
386 | 2018-05-07 | E18SM4261A | 1.2 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
387 | 2018-05-07 | E18SM4262A | 1.1 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
388 | 2018-05-08 | E18SM7980A | 3.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
389 | 2018-05-08 | E18SM4255A | 2.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
390 | 2018-05-08 | E18SM4258A | 3.2 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
391 | 2018-05-07 | E18SM4259A | 4.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
392 | 2018-05-07 | E18SM4263A | 1.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
393 | 2018-05-08 | E18SM4264A | 4.1 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
394 | 2018-05-08 | E18SM4265A | 5.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
395 | 2018-05-08 | E18SM4266A | 3.4 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
396 | 2018-05-08 | E18SM4267A | 2.4 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
397 | 2018-05-08 | E18SM4268A | 5.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
398 | 2018-05-08 | E18SM4269A | 5.3 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
399 | 2018-05-08 | E18SM4270A | 4.2 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
400 | 2018-05-08 | E18SM4271A | 4.4 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
401 | 2018-05-08 | E18SM4272A | 4.3 | 0.0015 | 0.002 | Pass on first test | Double Press Seal Eval
402 | 2018-05-08 | E18SM4273A | 3.3 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
403 | 2018-05-08 | E18SM4274A | 3.1 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
404 | 2018-05-08 | E18SM4278A | 1.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
405 | 2018-05-08 | E18SM4275A | 1.4 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
406 | 2018-05-08 | E18SM4280A | 2.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
407 | 2018-05-08 | E18SM4281A | 2.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
408 | 2018-05-09 | E18SM4283A | 5.1 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
409 | 2018-05-09 | E18SM4284A | 2.4 | 0.0015 | 0.002 | Resolve unit | Double Press Seal Eval
410 | 2018-05-09 | E18SM4285A | 5.3 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
411 | 2018-05-09 | E18SM4289A | 4.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
412 | 2018-05-09 | E18SM4290A | 3.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
413 | 2018-05-09 | E18SM4291A | 1.3 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
414 | 2018-05-09 | E18SM4292A | 2.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
415 | 2018-05-09 | E18SM4293A | 2.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
416 | 2018-05-09 | E18SM4294A | 2.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
417 | 2018-05-09 | E18SM4295A | 1.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
418 | 2018-05-09 | E18SM4297A | 3.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
419 | 2018-05-09 | E18SM4298A | 1.1 | NA | 0.002 | Resolve unit | Double Press Seal Eval
420 | 2018-05-10 | E18SM4276A | 1.2 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
421 | 2018-05-10 | E18SM4277A | 5.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
422 | 2018-05-10 | E18SM4288A | 3.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
423 | 2018-05-09 | E18SM4299A | 5.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
424 | 2018-05-09 | E18SM4300A | 5.3 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
425 | 2018-05-09 | E18SM4301A | 3.2 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
426 | 2018-05-10 | E18SM4302A | 2.4 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
427 | 2018-05-10 | E18SM4303A | 2.3 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
428 | 2018-05-10 | E18SM4304A | 1.4 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
429 | 2018-05-10 | E18SM4305A | 1.3 | 0.0018 | 0.002 | Pass on first test | Double Press Seal Eval
430 | 2018-05-10 | E18SM4306A | 3.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
431 | 2018-05-10 | E18SM4311A | 5.2 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
432 | 2018-05-10 | E18SM4312A | 1.1 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
433 | 2018-05-10 | E18SM4313A | 4.4 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
434 | 2018-05-10 | E18SM4315A | 5.3 | 0.0008 | 0.002 | Resolve unit | Double Press Seal Eval
435 | 2018-05-10 | E18SM4319A | 2.3 | 0.0017 | 0.002 | Pass on first test | Double Press Seal Eval
436 | 2018-05-10 | E18SM4320A | 2.4 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
437 | 2018-05-10 | E18SM4323A | 3.1 | 0.0019 | 0.002 | Pass on first test | Double Press Seal Eval

## Fail
Alliance CarryOver Test Monitoring
Unit | Date | S/N | Test Station (station. Test port) | Carryover % (<0.002%) | Current Status | Failure Mode | Action to Resolve | Failure Mode | Failure Summary | Total | Final Test Station | No. Of UUT's
Ex | 2018-04-01 |  | X.X | 0.0025 | Resolved | Abnormal Graph |  | High Blank | Abnormal Graph | 13 | 1.1 | 3
1 | 2018-04-07 | D18SM4962A | 5.2 | 0.0008 | Resolved | Abnormal Graph | REPLACE BLANK (9) | High Carryover % | Post Blank No Value | 7 | 1.2 | 6
2 | 2018-04-07 | D18SM4966A | 5.4 | 0.0002 | Resolved | Abnormal Graph | REPLACE BLANK (9) | Abnormal Graph | High Blank | 102 | 1.3 | 8
3 | 2018-04-09 | D18SM4972A | 1.3 | NA | Resolved | Abnormal Graph | REPLACE BLANK (9) | Post Blank no value | Retention Time Drifted | 1 | 1.4 | 12
4 | 2018-04-06 | C18SM7898A | 3.1 | 0.0021 | Resolved | Abnormal Graph | REPLACE BLANK (9) | Retention Time Drifted | High Carryover % | 43 | 2.1 | 7
5 | 2018-04-06 | C18SM7903A | 2.1 | NA | Resolved | Post Blank no value | REPLACE BLANK (9) | Status |  |  | 2.2 | 7
6 | 2018-04-06 | C18SM7904A | 1.1 | 0.0017 | Resolved | Abnormal Graph | REPLACE BLANK (9) | Resolved |  |  | 2.3 | 11
7 | 2018-04-06 | C18SM7906A | 1.4 | 0.0006 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9) 2. Change tester from 1.2 to tester 1.4 | Troubleshooting |  |  | 2.4 | 11
8 | 2018-04-06 | C18SM7908A | 3.2 | 0.001 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9) 2. Change tester from 2.2 to tester 3.2 |  |  |  | 3.1 | 10
9 | 2018-04-06 | C18SM7912A | 3.3 | 0.0032 | Resolved | High Blank | REPLACE BLANK (9) |  |  |  | 3.2 | 11
10 | 2018-04-06 | C18SM7916A | 4.3 | NA | Resolved | Post Blank no value | REPLACE BLANK (9) |  |  |  | 3.3 | 13
11 | 2018-04-07 | C18SM7913A | 3.3 | 0.0011 | Resolved | High Blank | 1. REPLACE BLANK (9), 2. Change tester from 3.4 to tester 3.3 3. Change tester from 3.3 to 5.1 4. Replace Seal (P/n: WAT 033276) |  |  |  | 3.4 | 9
12 | 2018-04-10 | C18SM7907A | 4.3 | 0.0032 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replaced Blank (9) 5. Replaced Blank (9) 6. Replaced Blank (9) 7. Replaced Seal (P/n: WAT 033276) |  |  |  | 4.1 | 10
13 | 2018-04-11 | D18SM4004A | 3.2 | NA | Resolved | Abnormal Graph | REPLACE BLANK (9) |  |  |  | 4.2 | 11
14 | 2018-04-11 | D18SM4999A | 1.3 | NA | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Retest 4. Retest 5. Retest |  |  |  | 4.3 | 7
15 | 2018-04-10 | D18SM4980A | 3.4 | 0.0008 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9), 2. Replace Seal (P/n: WAT 033276) |  |  |  | 4.4 | 6
16 | 2018-04-09 | D18SM4971A | 1.4 | NA | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  |  |  | 5.1 | 6
17 | 2018-04-07 | D18SM4963A | 5.3 | 0.0006 | Resolved | Abnormal Graph | 1. Replaced Blank (9) 2. Replaced Blank (9) |  |  |  | 5.2 | 8
18 | 2018-04-12 | D18SM4003A | 1.2 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  |  |  | 5.3 | 7
19 | 2018-04-11 | D18SM4992A | 3.1 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  |  |  | 5.4 | 3
20 | 2018-04-13 | D18SM4021A | 2.3 | NA | Resolved | Retention Time Drifted | 1. Replaced Blank (9) 2. Replaced GPV p/n: WAT270524 |  |  |  | Total | 166
21 | 2018-04-09 | C18SM7905A | 3.3 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
22 | 2018-04-16 | D18SM7933A | 4.1 | 0.0067 | Resolved | High Blank | 1. Replace Seal (P/n: WAT 033276)
23 | 2018-04-14 | D18SM7923A | 4.2 | 0.0013 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Seal (P/n WAT033276)
24 | 2018-04-16 | D18SM7924A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276) 3. Replaced Blank (9)
25 | 2018-04-14 | D18SM7927A | 2.1 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276)
26 | 2018-04-16 | D18SM7929A | 2.4 | 0.003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276)
27 | 2018-04-16 | D18SM7931A | 4.4 | 0.0003 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276)
28 | 2018-04-16 | D18SM7932A | 3.1 | 0.0007 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276)
29 | 2018-04-16 | D18SM7935A | 3.4 | 0.0005 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276)
30 | 2018-04-16 | D18SM7940A | 1.4 | NA | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276)
31 | 2018-04-16 | D18SM7944A | 5.3 | 0.005 | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Retest
32 | 2018-04-16 | D18SM4045A | 4.3 | 0.0147 | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Retest
33 | 2018-04-17 | D18SM4048A | 2.2 | 0.0035 | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Replaced Blank (9)
34 | 2018-04-18 | D18SM7919A | 3.4 | 0.002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
35 | 2018-04-17 | D18SM7934A | 2.3 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector seal (P/n:WAT033276) 4. Replaced injector (WAT279003915) from s/n:15J9D055 to 15J9D087
36 | 2018-04-17 | D18SM7943A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
37 | 2018-04-18 | D18SM7945A | 1.4 | 0.0009 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
38 | 2018-04-18 | D18SM4047A | 5.1 | NA | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
39 | 2018-04-17 | D18SM4049A | 4.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector seal (P/n:WAT033276)
40 | 2018-04-17 | D18SM4050A | 3.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Realigned and retest Injector (P/n WAT279003915)
41 | 2018-04-17 | D18SM4056A | 2.4 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
42 | 2018-04-18 | D18SM4063A | 3.2 | 0.0002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
43 | 2018-04-18 | D18SM4065A | 2.3 | 0.0424 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
44 | 2018-04-18 | D18SM4053A | 4.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
45 | 2018-04-19 | D18SM4057A | 2.1 | 0.0051 | Resolved | High Carryover % | 1. Replaced Blank (9)
46 | 2018-04-18 | D18SM4059A | 4.4 | 0.012 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
47 | 2018-04-19 | D18SM4069A | 5.2 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
48 | 2018-04-19 | D18SM4075A | 2.3 | 0.0036 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
49 | 2018-04-19 | D18SM4076A | 3.2 | 0.008 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
50 | 2018-04-19 | D18SM4078A | 3.4 | 0.0002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
51 | 2018-04-19 | D18SM4081A | 3.1 | 0.0047 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
52 | 2018-04-16 | D18SM4042A | 3.3 | 0.0013 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Seal (P/n WAT03276) 4. Change tester from tester 3.3 to tester 1.3
53 | 2018-04-19 | D18SM4074A | 2.2 | 0.0063 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3. Replaced tubing p/n 430001336 as tubing deformed during realignment of injector
54 | 2018-04-20 | D18SM4079A | 4.1 | 0.0063 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3. Replaced tubing p/n 430001336 as tubing deformed during realignment of injector
55 | 2018-04-19 | D18SM4084A | 2.4 | 0.0481 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
56 | 2018-04-20 | D18SM4090A | 4.2 | 0.0088 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
57 | 2018-04-20 | D18SM4092A | 2.3 | 0.005 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
58 | 2018-04-20 | D18SM4094A | 3.4 | 0.0009 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
59 | 2018-04-21 | D18SM4089A | 3.3 | 0.0056 | Resolved | High Carryover % | 1. Retest 2. Realigned and retest Injector (P/n WAT279003915)
60 | 2018-04-24 | D18SM7947A | 3.4 | 0.0006 | Resolved | High Blank | 1. Replaced Blank (9)
61 | 2018-04-24 | C18SM7900A | 4.3 | NA | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) 2. Change tester from tester 4.1 to tester 4.3 3. Realigned and retest Injector (P/n WAT279003915) 4. Realigned and retest Injector (P/n WAT279003915)
62 | 2018-04-23 | D18SM4101A | 2.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
63 | 2018-04-23 | D18SM4108A | 3.1 | 0.0014 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
64 | 2018-04-24 | D18SM4110A | 4.2 | 0.001 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
65 | 2018-04-24 | D18SM4119A | 1.3 | 0.0059 | Resolved | High Carryover % | 1. Retest
66 | 2018-04-24 | D18SM4120A | 2.3 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
67 | 2018-04-25 | D18SM4083A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915)
68 | 2018-04-25 | D18SM4113A | 4.4 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915)
69 | 2018-04-24 | D18SM4117A | 5.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
70 | 2018-04-25 | D18SM4123A | 1.4 | 0.0011 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
71 | 2018-04-24 | D18SM4126A | 4.1 | 0.0003 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915)
72 | 2018-04-25 | D18SM4127A | 5.3 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
73 | 2018-04-25 | D18SM4134A | 1.4 | NA | Resolved | High Blank | 1. Replaced Blank (9)
74 | 2018-04-26 | D18SM4137A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9)
75 | 2018-04-26 | D18SM4140A | 1.3 | 0.0007 | Resolved | High Blank | 1. Replaced Blank (9)
76 | 2018-04-25 | D18SM7953A | 2.4 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915)
77 | 2018-04-25 | D18SM7956A | 3.4 | 0.0006 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
78 | 2018-04-26 | D18SM4109A | 3.2 | 0.0027 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915)
79 | 2018-04-26 | D18SM4114A | 2.1 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915)
80 | 2018-04-25 | D18SM4124A | 5.1 | 0.0003 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) 2. Realigned and retest Injector (P/n WAT279003915)
81 | 2018-04-26 | D18SM4146A | 5.3 | 0.0073 | Resolved | High Carryover % | 1. Retest
82 | 2018-04-26 | D18SM4155A | 4.2 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9)
83 | 2018-04-26 | D18SM4142A | 2.2 | 0.0021 | Resolved | High Carryover % | 1. Replaced Blank (9)
84 | 2018-04-26 | D18SM4143A | 1.4 | 0.0013 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
85 | 2018-04-26 | D18SM4135A | 5.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and Retest Injector (279003915)
86 | 2018-04-27 | D18SM4145A | 4.1 | 0.0009 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and Retest Injector (279003915)
87 | 2018-04-27 | D18SM4148A | 2.4 | 0.0009 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
88 | 2018-04-27 | D18SM4152A | 5.2 | 0.0021 | Resolved | High Carryover % | 1. Replaced Blank (9)
89 | 2018-04-27 | D18SM4154A | 4.3 | NA | Resolved | High Blank | 1. Replaced Blank (9)
90 | 2018-04-27 | D18SM4160A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest
91 | 2018-04-27 | D18SM4138A | 3.3 | 0.0055 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
92 | 2018-04-25 | D18SM7954A | 4.2 | 0.0012 | Resolved | High Blank | 1. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
93 | 2018-04-28 | D18SM7957A | 2.1 | 0.0084 | Resolved | High Carryover % | 1. Replaced Blank (9)
94 | 2018-04-30 | D18SM7964A | 5.2 | 0.0027 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D075 to 16J9D105
95 | 2018-04-30 | D18SM7967A | 2.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
96 | 2018-04-28 | D18SM7968A | 2.1 | 0.0005 | Resolved | High Blank | 1. Realigned and Retest Injector (279003915)
97 | 2018-04-30 | D18SM7974A | 5.1 | 0.0003 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
98 | 2018-04-28 | D18SM4139A | 2.3 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
99 | 2018-04-28 | D18SM4161A | 3.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Realigned and Retest Injector (279003915)
100 | 2018-04-30 | D18SM4163A | 4.2 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915)
101 | 2018-04-30 | D18SM4164A | 4.4 | 0.0046 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D053 to 16J9D104
102 | 2018-04-30 | D18SM4175A | 2.3 | 0.0041 | Resolved | High Carryover % | 1. Replaced Blank (9)
103 | 2018-04-30 | D18SM4178A | 4.3 | 0.0174 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
104 | 2018-05-02 | D18SM4176A | 2.1 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced injector (WAT279003915) from s/n:16J9D093 to 17J9D006
105 | 2018-04-30 | D18SM4181A | 4.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
106 | 2018-05-02 | D18SM4183A | 2.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D093 to 17J9D006
107 | 2018-05-02 | D18SM4185A | 1.2 | 0.0006 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
108 | 2018-05-03 | D18SM4189A | 2.2 | 0.0018 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9)
109 | 2018-05-02 | D18SM4192A | 4.2 | 0.0046 | Resolved | High Carryover % | 1. Retest
110 | 2018-05-02 | D18SM4196A | 1.3 | 0.0006 | Resolved | High Blank | 1. Retest
111 | 2018-05-02 | D18SM4197A | 4.1 | 0.0025 | Resolved | High Carryover % | 1. Replaced Blank (9)
112 | 2018-05-02 | D18SM4199A | 3.3 | 0.001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
113 | 2018-05-02 | D18SM4201A | 3.2 | 0.0016 | Resolved | High Blank | 1. Retest
114 | 2018-04-30 | D18SM4168A | 5.4 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D060 to 17J9D028
115 | 2018-05-03 | D18SM4195A | 2.4 | 0.0004 | Resolved | High Blank | 1. Retest
116 | 2018-05-03 | D18SM4211A | 4.1 | 0.0018 | Resolved | Abnormal Graph | 1. Retest 2. Replaced Blank (9)
117 | 2018-05-03 | D18SM4214A | 1.3 | 0.0057 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
118 | 2018-05-03 | D18SM4217A | 2.2 | 0.0002 | Resolved | High Blank | 1. Retest
119 | 2018-04-28 | D18SM7961A | 4.1 | 0.001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
120 | 2018-05-04 | E18SM7978A | 1.4 | 0.0033 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (()
121 | 2018-05-04 | E18SM7984A | 4.4 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
122 | 2018-05-04 | E18SM7987A | 1.3 | 0.0009 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
123 | 2018-05-05 | E18SM7989A | 2.4 | 0.0005 | Resolved | High Blank | 1. Retest
124 | 2018-05-03 | E18SM4207A | 4.3 | 0.0012 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
125 | 2018-04-30 | D18SM4180A | 5.3 | 0.0109 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 16J9D086 to J17J9D055
126 | 2018-05-03 | E18SM4216A | 3.3 | 0.0048 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 17J9D049 to J17J9D026
127 | 2018-05-03 | E18SM4218A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replace Injector (279003915) from SN 17J9D027 to J17J9D034
128 | 2018-05-04 | E18SM4222A | 4.1 | 0.002 | Resolved | High Blank | 1. Retest
129 | 2018-05-04 | E18SM4223A | 4.4 | 0.0027 | Resolved | High Carryover % | 1. Retest
130 | 2018-05-04 | E18SM4225A | 5.1 | 0.0057 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
131 | 2018-05-05 | E18SM4230A | 3.2 | 0.0001 | Resolved | High Blank | 1. Retest
132 | 2018-05-05 | E18SM4235A | 1.4 | 0.0021 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
133 | 2018-05-05 | E18SM4238A | 4.1 | NA | Resolved | Abnormal Graph | 1. Retest 2. Replaced Blank (9)
134 | 2018-05-05 | E18SM4232A | 2.3 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replaced Blank (9)
135 | 2018-05-05 | E18SM4233A | 5.2 | 0.0005 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9)
136 | 2018-05-05 | E18SM4236A | 3.3 | NA | Resolved | Abnormal Graph | 1. Retest
137 | 2018-05-07 | E18SM4240A | 5.1 | 0.0003 | Resolved | High Blank | 1. Retest
138 | 2018-05-07 | E18SM4247A | 3.3 | NA | Resolved | High Blank | 1. Retest
139 | 2018-05-07 | E18SM4248A | 3.4 | 0.0032 | Resolved | High Carryover % | 1. Retest
140 | 2018-05-07 | E18SM4249A | 2.4 | 0.0034 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9)
141 | 2018-05-07 | E18SM4250A | 1.4 | 0.006 | Resolved | High Blank | 1. Retest
142 | 2018-05-07 | E18SM4252A | 5.2 | 0.0009 | Resolved | High Blank | 1. Retest
143 | 2018-05-07 | E18SM4257A | 3.3 | 0.0049 | Resolved | High Carryover % | 1. Retest
144 | 2018-05-07 | E18SM4255A | 2.3 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replace Injector (279003915) from SN 17J9D083 to 17J9D028
145 | 2018-05-08 | E18SM4258A | 3.2 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9)
146 | 2018-05-08 | E18SM4263A | 1.3 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 17J0D008 to 17JCD008
147 | 2018-05-08 | E18SM4266A | 3.4 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
148 | 2018-05-08 | E18SM4269A | 5.3 | 0.0026 | Resolved | High Carryover % | 1. Replaced Blank (9)
149 | 2018-05-08 | E18SM4270A | 4.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
150 | 2018-05-08 | E18SM4273A | 3.3 | 0.0036 | Resolved | High Carryover % | 1. Replaced Blank (9)
151 | 2018-05-08 | E18SM4274A | 3.1 | 0.0042 | Resolved | High Carryover % | 1. Replaced Blank (9)
152 | 2018-05-08 | E18SM4275A | 1.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
153 | 2018-05-09 | E18SM4283A | 5.1 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9)
154 | 2018-05-09 | E18SM4284A | 2.4 | 0.0052 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9)
155 | 2018-05-09 | E18SM4289A | 4.2 | 0.0044 | Resolved | High Carryover % | 1. Replaced Blank (9)
156 | 2018-05-09 | E18SM4295A | 1.4 | 0.0023 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9)
157 | 2018-05-09 | E18SM4298A | 1.1 | 0.0039 | Resolved | High Carryover % | 1. Replaced Blank (9)
158 | 2018-05-09 | E18SM4276A | 1.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
159 | 2018-05-10 | E18SM4288A | 3.3 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
160 | 2018-05-10 | E18SM4301A | 3.2 | 0.0016 | Resolved | High Blank | 1. Replaced Blank (9)
161 | 2018-05-09 | E18SM4302A | 2.4 | 0.0023 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9)
162 | 2018-05-10 | E18SM4303A | 2.3 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)
163 | 2018-05-10 | E18SM4306A | 3.1 | 0.0007 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9)
164 | 2018-05-10 | E18SM4311A | 5.2 | 0.0033 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9)
165 | 2018-05-10 | E18SM4312A | 1.1 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9)
166 | 2018-05-10 | E18SM4315A | 5.3 | 0.0095 | Resolved | High Carryover % | 1. Replaced Blank (9)
Note on D3: sgycha: Test station 1; Port 3
Note on G3: sgycha: drop downlist; state if otherwise

## FM Mon
Count of S/N | Column Labels
Row Labels | Abnormal Graph | High Blank | Post Blank no value | Retention Time Drifted | High Carryover % | Grand Total
2018-04-01
2018-04-06 | 4 | 1 | 2 |  |  | 7
2018-04-07 | 3 | 1 |  |  |  | 4
2018-04-09 | 1 | 1 | 1 |  |  | 3
2018-04-10 | 1 | 1 |  |  |  | 2
2018-04-11 | 1 | 1 | 1 |  |  | 3
2018-04-12 |  | 1 |  |  |  | 1
2018-04-13 |  |  |  | 1 |  | 1
2018-04-14 |  | 2 |  |  |  | 2
2018-04-16 |  | 8 | 2 |  |  | 10
2018-04-17 |  | 5 | 1 |  |  | 6
2018-04-18 |  | 7 |  |  |  | 7
2018-04-19 |  | 2 |  |  | 6 | 8
2018-04-20 |  | 1 |  |  | 3 | 4
2018-04-21 |  |  |  |  | 1 | 1
2018-04-23 |  | 2 |  |  |  | 2
2018-04-24 |  | 6 |  |  | 1 | 7
2018-04-25 |  | 9 |  |  |  | 9
2018-04-26 |  | 6 |  |  | 3 | 9
2018-04-27 |  | 4 |  |  | 2 | 6
2018-04-28 |  | 4 |  |  | 1 | 5
2018-04-30 |  | 5 |  |  | 5 | 10
2018-05-02 |  | 6 |  |  | 2 | 8
2018-05-03 | 1 | 5 |  |  | 2 | 8
2018-05-04 |  | 3 |  |  | 3 | 6
2018-05-05 | 2 | 4 |  |  | 1 | 7
2018-05-07 |  | 5 |  |  | 3 | 8
2018-05-08 |  | 5 |  |  | 3 | 8
2018-05-09 |  | 2 |  |  | 5 | 7
2018-05-10 |  | 5 |  |  | 2 | 7
Grand Total | 13 | 102 | 7 | 1 | 43 | 166

## Failure
# | Failure Code | Example | Trouble shooting guide
1 | High Carryover % |  | RETEST WITHOUT CHANGING BLANKS If Pass,Ship If fail, replace Seal PN WAT033276 (Saint Gobain) If after replace and fail again, on hold for review
2 | High Blank |  | RETEST WITHOUT CHANGING BLANKS If Pass,Ship If fail, replace Seal PN WAT033276 (Saint Gobain) If after replace seal, fail again on hold for review
3 | Abnormal Peak |  | RETEST WITHOUT CHANGING BLANKS If Pass,Ship If fail, replace blank fail again, on hold for review
4 | Post Blank no data |  | RETEST WITHOUT CHANGING BLANKS If Pass,Ship If fail, replace blank fail again, on hold for review

## Double Press of Seal
Double Press Evaluation Result
1 | D18SM4162A | Static Leak | Result | Qty
2 | D18SM4163A | High Blank | High Blank | 5
3 | D18SM4164A | High Carryover | High Carryover | 2
4 | D18SM4165A | Pass | Pass | 13
5 | D18SM4166A | Pass
6 | D18SM4167A | Pass
7 | D18SM4168A | High Blank
8 | D18SM4169A | Pass
9 | D18SM4170A | Pass
10 | D18SM4171A | Pass
11 | D18SM7959A | Pass
12 | D18SM7958A | Static Leak
13 | D18SM7957A | High Blank
14 | D18SM7960A | Pass
15 | D18SM7961A | High Blank
16 | D18SM7962A | Pass
17 | D18SM7963A | pass
18 | D18SM7964A | High Blank
19 | D18SM7965A | Pass
20 | D18SM7966A | Pass

## Eval Summary
SN | Evaluation | Cut In Serial/ Control Lot Size | Total Tested | First Pass | Yield
1 | ASP Execution Eval_Hand Tighthen | 24 | 24 | 14 | 58.33%
2 | Doule Press Seal | D18SM4162A D18SM7957A | 178 | 105 | 58.99%

## FPY
Count of Remarks | Column Labels
Row Labels | Pass on first test | Resolve unit | Grand Total | Date | Pass on first test | Resolve unit | Grand Total | FPY | Target
2018-04-06 | 8 | 7 | 15 | 2018-04-06 | 8 | 7 | 15 | 53.33% | 80%
2018-04-07 | 9 | 2 | 11 | 2018-04-07 | 9 | 2 | 11 | 81.82% | 80%
2018-04-09 | 9 | 4 | 13 | 2018-04-09 | 9 | 4 | 13 | 69.23% | 80%
2018-04-10 | 10 |  | 10 | 2018-04-10 | 10 |  | 10 | 100% | 80%
2018-04-11 | 6 | 5 | 11 | 2018-04-11 | 6 | 5 | 11 | 54.55% | 80%
2018-04-12 | 7 | 2 | 9 | 2018-04-12 | 7 | 2 | 9 | 77.78% | 80%
2018-04-14 | 10 | 3 | 13 | 2018-04-14 | 10 | 3 | 13 | 76.92% | 80%
2018-04-13 | 14 |  | 14 | 2018-04-13 | 14 |  | 14 | 100% | 80%
2018-04-16 | 24 | 7 | 31 | 2018-04-16 | 24 | 7 | 31 | 77.42% | 80%
2018-04-17 | 4 | 3 | 7 | 2018-04-17 | 4 | 3 | 7 | 57.14% | 80%
2018-04-18 | 9 | 7 | 16 | 2018-04-18 | 9 | 7 | 16 | 56.25% | 80%
2018-04-19 | 6 | 8 | 14 | 2018-04-19 | 6 | 8 | 14 | 42.86% | 80%
2018-04-20 | 10 | 6 | 16 | 2018-04-20 | 10 | 6 | 16 | 62.5% | 80%
2018-04-21 | 2 | 2 | 4 | 2018-04-21 | 2 | 2 | 4 | 50% | 80%
2018-04-23 | 10 | 1 | 11 | 2018-04-23 | 10 | 1 | 11 | 90.91% | 80%
2018-04-24 | 6 | 7 | 13 | 2018-04-24 | 6 | 7 | 13 | 46.15% | 80%
2018-04-25 | 11 | 5 | 16 | 2018-04-25 | 11 | 5 | 16 | 68.75% | 80%
2018-04-26 | 8 | 13 | 21 | 2018-04-26 | 8 | 13 | 21 | 38.1% | 80%
2018-04-27 | 5 | 8 | 13 | 2018-04-27 | 5 | 8 | 13 | 38.46% | 80%
2018-04-28 | 11 | 2 | 13 | 2018-04-28 | 11 | 2 | 13 | 84.62% | 80%
2018-04-30 | 10 | 11 | 21 | 2018-04-30 | 10 | 11 | 21 | 47.62% | 80%
2018-05-02 | 11 | 7 | 18 | 2018-05-02 | 11 | 7 | 18 | 61.11% | 80%
2018-05-03 | 15 | 11 | 26 | 2018-05-03 | 15 | 11 | 26 | 57.69% | 80%
2018-05-04 | 10 | 6 | 16 | 2018-05-04 | 10 | 6 | 16 | 62.5% | 80%
2018-05-05 | 5 | 5 | 10 | 2018-05-05 | 5 | 5 | 10 | 50% | 80%
2018-05-07 | 16 | 11 | 27 | 2018-05-07 | 16 | 11 | 27 | 59.26% | 80%
2018-05-08 | 10 | 8 | 18 | 2018-05-08 | 8 | 7 | 15 | 53.33% | 80%
2018-05-09 | 9 | 6 | 15 | 2018-05-09 | 7 | 5 | 12 | 58.33% | 80%
2018-05-10 | 7 | 8 | 15 | 2018-05-10 | 4 | 6 | 10 | 40% | 80%
Grand Total | 272 | 165 | 437