## consolidattion
eAlliance Carry Over Test Yield Improvement Actions Summary
SN | Implementation Date | Evaluation | Week | 2018-04-06 | 2018-04-13 | 2018-04-20 | 2018-04-27 | 2018-05-04 | 2018-05-11 | 2018-05-18 | 2018-05-25 | 2018-06-01 | 2018-06-08 | 2018-06-15 | 2018-06-22 | 2018-06-29 | 2018-07-06 | Remarks
|  |  | Yield | 53% | 80% | 77% | 55% | 60% | 53% | 55% | 66% | 60% | 64% | 63% | 69% | 88% | TBD
1 | 2018-04-04 | 04 Apr 18 1. Fill both 'Blank" vials (white cap) with new solution. 2. Placement of both vials on correct position (Carousel A - position 3 & 6) 3. After completion of Alliance Final test, discard both the 'Blank' solution. 4. Repeat step 1 for new Alliance unit. | Closed | bar: 2018-04-06
2 | 2018-04-06 | All 3x press and holding fixtures bolted to station to avoid movement | Closed | bar: 2018-04-06
3 | 2018-04-06 | Replace Blank (9) and change tester position from 2.2 to 3.2 to check if failure follows. Replace Blank(9) and change tester position from 3.4 to 3.3 and to 5.1. Conclusion : Failure followed. Tester no issue. | Closed | bar: 2018-04-06
4 | 2018-04-07 | Replace Seal (P/N WAT033276) for 7x units. 5 passed but takes longer cycle time to replace seal on the defect units and re-use. | Closed | bar: 2018-04-13
5 | 2018-04-11 | Used modified test result with Specification : Post Blank <0.002 (or Blank Peak(s) mising/blank) | Closed | bar: 2018-04-13
6 | 2018-04-18 | Replace Injector instead of Seal | bar: 2018-04-20
7 | 2018-04-20 | 4M Analysis for Injector Assy & Test - Evaluated and executed tightening sequence( 4 inlb and then 10 in lb) | Closed | bar: 2018-04-20
8 | 2018-04-27 | Implemented Double press of Seal | Closed | bar: 2018-04-27
9 | 2018-05-07 | Change preparation of Blanks 30:70 Mix from Weekly to Daily preparation | Closed | bar: 2018-05-11
10 | 2018-05-08 | Removal of Re-test after prime test fail without doing any thing | Closed | bar: 2018-05-11
11 | 2018-05-15 | Change of Blank from 30 :70 to 20:80 Mix | Closed | bar: 2018-05-18
12 | 2018-05-16 | Use of Milford 20:80 Mix (1st bottle) | Closed | bar: 2018-05-18
13 | 2018-05-16 | Production to prepare post blank + new vial + 1 new cap and perform Carryover test. If Carry over test fail, Production to replace post blank + new vial + 1 new cap and 1st Re-test with TG 3 & 4 If 1st re-test fail, Production to replace Injector and post blank + new vial + 1 new cap and perform 2nd Re-test with TG 3 & 4 If 2nd re-test fail, Contact WPPL Engineering | Closed | bar: 2018-05-18
14 | 2018-05-16 | Use of New Post Blank + New Vial + New Cap for every re-test | Closed | bar: 2018-05-18
15 | 2018-05-18 | Stop using TG 3 & 4 | Closed | bar: 2018-05-25
16 | 2018-05-19 | Re-start using TG 3 & 4 | Closed | bar: 2018-05-25
17 | 2018-05-22 | Use of WPPL 20:80 Mix (22-25 May 2018) | Closed | bar: 2018-05-25
18 | 2018-05-23 | Production to prepare post blank + re-use vial + 1 new cap and perform Carryover test. If carryover test fail, Production to replace post blank + re-use vial + 1 new cap and 1st re-test with TG 3 & 4 If 1st re-test fail, Production to replace post blank + re-use vial + 1 new cap and 2nd re-test with TG 3 & 4 If 2nd re-test fail, Production to replace Injector and post blank + re-use vial + 1 new cap and 3rd re-test with TG 3 & 4 If 3rd re-test fail, Contact WPPL Engineering | Closed | bar: 2018-05-25
19 | 2018-05-23 | Re-train on chemical Handling(Close dispenser cap after use,Re-fill vial, not touch dispenser tip | Closed | bar: 2018-05-25
20 | 2018-05-23 | Re-use of vials after washing with Ultrapure water 3x times | Closed | bar: 2018-05-25
21 | 2018-05-23 | Set up and Qualification of 6th final tester | In-progress | 2018-05-25: 3 ports Qualified and In -use
22 | 2018-05-25 | To ship 2x injector that failed carry over test for Milford evaluation | In-progress | 2018-06-01: shipped on 05 Jun 18. Pending Waters update | 1x High blank becomes High carry over and 1x passed High blank. Unable to replicate the failures.
23 | 2018-05-26 | Use of WPPL 20:80 Mix (2nd bottle) (26-31 May 2018) | Closed | bar: 2018-06-01
24 | 2018-06-01 | Use of WPPL 20:80 Mix (3rd bottle) (01 Jun onwards ) | Closed | bar: 2018-06-01
25 | 2018-06-01 | Re-visit Injector process and ensure process is ok | Closed | bar: 2018-06-01
26 | 2018-06-01 | Flushing with methanol on seal body | Closed | bar: 2018-06-01
27 | 2018-06-06 | WPPL 20:80 Mix (4th bottle) Flushing with methanol on seal body | Closed | bar: 2018-06-08
28 | 2018-06-09 | WPPL 20:80 Mix (5th bottle) Flushing with methanol on seal body | Closed | bar: 2018-06-15
29 | 2018-06-13 | WPPL 20:80 Mix (6th bottle) Flushing with methanol on seal body | Closed | bar: 2018-06-15
30 | 2018-06-19 | WPPL 20:80 Mix (7th bottle) Flushing with methanol on seal body | Closed | bar: 2018-06-22
31 | 2018-06-19 | Seal Inner hole control at 0.050 inches with Injector Stream/Seal/ Down (Qt 50 pcs) | Closed | 2018-06-15: 56% Yield | Total tested 49 units with 34 pass and 15 fail
32 | 2018-06-20 | High blank elimination evaluation | In-progress | 2018-06-22: Pending Hua zhen update | 6x fresh units and 3x high blank units to run using QA template and review the results
33 | 2018-06-23 | WPPL 20:80 Mix (8th bottle) Flushing with methanol on seal body | Closed | bar: 2018-06-29 to 2018-07-06
34 | 2018-06-25 | Record down Injector adjustment readings when re-aligned | Closed
35 | 2018-07-02 | Measure the needle wash volume for high blank failure units | In-progress | [14 empty] | Seeing some "high blank" like behaviors associated with low needle wash. Perform Needle Wash Pump effectivity test - Run 10 quick 0.5 minute injections and record values. | bar: 2018-07-06

## 50pcs Seal Pack
Eval on Seal Pack (WAT033276)
| Upper Seal (WAT033276) | Lower Seal (WAT033276) | Injector Serial Number | Instrument Serial Number | Stream count | Seal Count | Down Count | Injector adjustment at Final Test (Y/N) | Initial Down Count at Final Tester | Top (38-50) | Width (60-90) | Final Test Result Carryover (Pass / Fail ) | Actual Reading | Failure Code | Remark
1 | 0.05 | 0.05 | 24J9D047 | F18SM4746A | 101 | 76 | 42 | Y | 32 | 46 | 73 | PASS | 0.0002 | NA | NA
2 | 0.05 | 0.05 | 24J9D048 | F18SM4747A | 99 | 76 | 43 | Y | 33 | 46 | 75 | PASS | 0.0014 | NA | NA
3 | 0.05 | 0.05 | 24J9D049 | F18SM4748A | 101 | 74 | 41 | Y | 32 | 48 | 72 | PASS | 0.0008 | NA | NA
4 | 0.05 | 0.05 | 24J9D050 | F18SM4749A | 105 | 76 | 43 | Y | 33 | 48 | 77 | PASS | 0.0004 | NA | NA
5 | 0.05 | 0.05 | 24J9D051 | F18SM4750A | 101 | 76 | 43 | Y | 33 | 49 | 76 | PASS | 0.0005 | NA | NA
6 | 0.05 | 0.05 | 24J9D052 | F18SM4751A | 101 | 76 | 42 | Y | 37 | 43 | 76 | PASS | 0.0003 | NA | NA
7 | 0.05 | 0.05 | 24J9D053 | F18SM4752A | 101 | 78 | 43 | Y | 36 | 42 | 76 | FAIL | 0.0007 | High Blank | PASS 0.0004
8 | 0.05 | 0.05 | 24J9D054 | F18SM4753A | 101 | 76 | 41 | Y | 33 | 42 | 74 | FAIL | 0.0005 | High Blank | PASS 0.0008
9 | 0.05 | 0.05 | 24J9D055 | F18SM4754A | 105 | 74 | 42 | Y | Seal Pack Geometry Fault (1) | 42 | 74 | PASS | 0.0015 | NA | NA
10 | 0.05 | 0.05 | 24J9D056 | F18SM4755A | 97 | 78 | 43 | Y | 33 | 42 | 77 | FAIL | 0.001 | High Blank | PASS 0.0018
11 | 0.05 | 0.05 | 24J9D057 | F18SM4756A | 97 | 76 | 43 | Y | 36 | 43 | 75 | PASS | 0.0005 | NA | NA
12 | 0.05 | 0.05 | 24J9D058 | F18SM4757A | 103 | 74 | 41 | N | 39 | 39 | 74 | PASS | 0.0005 | NA | NA
13 | 0.05 | 0.05 | 24J9D059 | F18SM4758A | 99 | 76 | 42 | Y | 36 | 46 | 73 | PASS | 0.001 | NA | NA
14 | 0.05 | 0.05 | 24J9D060 | F18SM4759A | 103 | 72 | 41 | Y | 36 | 43 | 75 | PASS | 0.0012 | NA | NA
15 | 0.05 | 0.05 | 24J9D061 | F18SM4760A | 103 | 80 | 41 | Y | 33 | 46 | 78 | FAIL | 0.0016 | High Blank | PASS 0.0013
16 | 0.05 | 0.05 | 24J9D062 | F18SM4761A | 99 | 76 | 44 | Y | Seal Pack Geometry Fault (1) | 42 | 72 | PASS | 0.0008 | NA | NA
17 | 0.05 | 0.05 | 24J9D063 | F18SM4762A | 105 | 78 | 43 | Y | Seal Pack Geometry Fault (1) | 49 | 76 | PASS | 0.0006 | NA | NA
18 | 0.05 | 0.05 | 24J9D064 | F18SM4763A | 99 | 76 | 40 | Y | 33 | 43 | 76 | FAIL | 0.0015 0.0048 0.0042 | High Blank HRPB % HRPB % | PASS 0.0012 | *
19 | 0.05 | 0.05 | 24J9D065 | F18SM4764A | 107 | 72 | 42 | Y | 32 | 46 | 72 | PASS | 0.0013 | NA | NA
20 | 0.05 | 0.05 | 24J9D066 | F18SM4765A | 99 | 76 | 42 | Y | 33 | 43 | 76 | PASS | 0.0007 | NA | NA
21 | 0.05 | 0.05 | 24J9D067 | F18SM4766A | 103 | 74 | 43 | Y | 33 | 46 | 73 | PASS | 0.0018 | NA | NA
22 | 0.05 | 0.05 | 24J9D068 | F18SM4767A | 105 | 74 | 42 | Y | 37 | 43 | 75 | FAIL | 0.0017 | High Blank | PASS 0.0011
23 | 0.05 | 0.05 | 24J9D069 | F18SM4768A | 103 | 76 | 40 | N | 43 | 43 | 76 | FAIL | 0.0009 | High Blank | PASS 0.0016
24 | 0.05 | 0.05 | 24J9D070 | F18SM4769A | 101 | 76 | 40 | Y | 32 | 42 | 77 | FAIL | 0.0022 | High Blank | PASS 0.0017
25 | 0.05 | 0.05 | 24J9D071 | F18SM4770A | 101 | 74 | 40 | Y | 32 | 48 | 72 | PASS | 0.0008 | NA | NA
26 | 0.05 | 0.05 | 24J9D072 | F18SM4771A | 105 | 70 | 43 | Y | 37 | 43 | 70 | FAIL | 0.0048 0.0042 | HRPB % HRPB % | PASS 0.0007 | *
27 | 0.05 | 0.05 | 24J9D073 | F18SM4772A | 101 | 76 | 40 | Y | Seal Pack Geometry Fault (1) | 42 | 76 | PASS | 0.0008 | NA | NA
28 | 0.05 | 0.05 | 24J9D074 | F18SM4773A | 99 | 78 | 40 | Y | 32 | 43 | 78 | PASS | 0.001 | NA | NA
29 | 0.05 | 0.05 | 24J9D075 | F18SM4774A | 101 | 76 | 44 | N | 39 | 39 | 77 | FAIL | 0.0042 | HRPB % | PASS 0.0010
30 | 0.05 | 0.05 | 24J9D076 | F18SM4775A | 99 | 76 | 42 | Y | 32 | 43 | 76 | FAIL | 0.0016 0.0027 | High Blank , HRPB % | PASS 0.0013
31 | 0.05 | 0.05 | 24J9D077 | F18SM4776A | 101 | 76 | 38 | Y | 32 | 49 | 76 | PASS | 0.0008 | NA | NA
32 | 0.05 | 0.05 | 24J9D078 | F18SM4777A | 105 | 72 | 44 | Y | 37 | 39 | 74 | PASS | 0.0002 | NA | NA
33 | 0.05 | 0.05 | 24J9D079 | F18SM4778A | 101 | 76 | 44 | Y | 36 | 43 | 75 | PASS | 0.0012 | NA | NA
34 | 0.05 | 0.05 | 24J9D080 | F18SM4779A | 105 | 76 | 44 | Y | 36 | 46 | 75 | PASS | 0.0019 | NA | NA
35 | 0.05 | 0.05 | 24J9D081 | F18SM4780A | 103 | 78 | 41 | Y | Seal Pack Geometry Fault (1) | 43 | 76 | FAIL | 0.0006 0.0010 | High Blank High Blank | PASS 0.0005 | *
36 | 0.05 | 0.05 | 25J9D002 | F18SM4781A | 101 | 72 | 40 | N | 48 | 48 | 75 | PASS | 0.0018 | NA | NA
37 | 0.05 | 0.05 | 25J9D003 | F18SM4782A | 97 | 78 | 43 | Y | 33 | 43 | 78 | PASS | 0.0003 | NA | NA
38 | 0.05 | 0.05 | 25J9D004 | F18SM4783A | 93 | 78 | 44 | Y | 33 | 43 | 78 | PASS | 0.0004 | NA | NA
39 | 0.05 | 0.05 | 25J9D005 | F18SM4784A | 101 | 78 | 40 | Y | 33 | 43 | 76 | PASS | 0.0008 | NA | NA
40 | 0.05 | 0.05 | 25J9D006 | F18SM4785A | 99 | 76 | 41 | Y | Seal Pack Geometry Fault (1) | 39 | 77 | PASS | 0.0003 | NA | NA
41 | 0.05 | 0.05 | 25J9D007 | F18SM4786A | 97 | 78 | 41 | Y | 36 | 42 | 78 | PASS | 0.0006 | NA | NA
42 | 0.05 | 0.05 | 25J9D008 | F18SM4787A | 101 | 74 | 43 | Y | 33 | 42 | 74 | FAIL | 0.0002 | High Blank | PASS 0.0000 | *
43 | 0.05 | 0.05 | 25J9D009 | F18SM4788A | 99 | 76 | 43 | Y | 36 | 48 | 75 | PASS | 0.0016 | NA | NA
44 | 0.05 | 0.05 | 25J9D010 | F18SM4789A | 103 | 74 | 42 | Y | 32 | 39 | 74 | FAIL | 0.0002 | High Blank | PASS 0.0005 | *
45 | 0.05 | 0.05 | 25J9D011 | F18SM4790A | 101 | 76 | 43 | N | 42 | 42 | 76 | PASS | 0.0011 | NA | NA
46 | 0.05 | 0.05 | 25J9D012 | F18SM4791A | 99 | 74 | 44 | Y | 36 | 43 | 76 | PASS | 0.001 | NA | NA
47 | 0.05 | 0.05 | 25J9D013 | F18SM4792A | 103 | 74 | 43 | Y | 33 | 43 | 76 | FAIL | 0.0008 | High Blank | PASS 0.0008 | *
48 | 0.05 | 0.05 | 25J9D014 | F18SM4793A | 99 | 76 | 42 | Y | Seal Pack Geometry Fault (1) | 48 | 75 | PASS | 0.0004 | NA | NA
49 | 0.05 | 0.05 | 25J9D015 | F18SM7145A | 95 | 78 | 41 | Y | 33 | 43 | 74 | PASS | 0.0015 | NA | NA
50 | 0.05 | 0.05 | 25J9D016 | F18SM4795A

## needle wash volume measurement
High Blank Unit needle wash volume measure
1st sample set
UUT(ml) | Inj1 | Inj2 | Inj3 | Inj4 | Inj5 | Inj6 | Inj7 | Inj8 | Inj9 | Inj10 | Total Volume
SN: F18SM7165A | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2 | 3.2
SN: F18SM4869A | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5
2nd sample set ( After needle wash prime)
UUT(ml) | Inj1 | Inj2 | Inj3 | Inj4 | Inj5 | Inj6 | Inj7 | Inj8 | Inj9 | Inj10 | Total Volume
SN: F18SM7165A | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1 | 3.1
SN: F18SM4869A | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5