## Summary
Summary | Amount (US$)
Calibration Cost- Box Build | 1009.1235
Calibration Cost-PCBA | 0
At hoc-Spare and Misc | 3602.3581
Total | 4611.4816

## Internal
Internal Calibration for Equipment used in Waters@Kallang Place
April 2020 Internal Calibration
[7 empty] | Sum |  |  | 873.7
[7 empty] | SG&A |  |  | 43.685
[7 empty] | Profit |  |  | 91.7385
[7 empty] | Subtotal Internal Calibration |  |  | 1009.1235
S/N | EQID | OWNER | Equipment Description | MANUFACTURER | MODEL | SERIAL NO. | CAL INTERVAL | CALIBRATION DATE | DUE DATE | COST (US$) | Remarks
1 | 0207 | Joel Mamalateo Deduyo | Column Heater | Waters | WAT091727 | M03CHM015R | 12 Months | 2020-04-03 | 2021-04-03 | 33.34
2 | 0208 | Joel Mamalateo Deduyo | Column Heater | Waters | WAT091727 | E01CHM203R | 12 Months | 2020-04-03 | 2021-04-03 | 33.34
3 | 0238 | Joel Mamalateo Deduyo | Column Heater | Waters | WAT038040 | MX6KM9444M | 12 Months | 2020-04-02 | 2021-04-02 | 33.34
4 | 0458 | Joel Mamalateo Deduyo | Auto Pressure (Flowcell Tester) | Waters | NA | TE-82475-2 | 12 Months | 2020-04-09 | 2021-04-09 | 141.94
5 | 1429 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | D1387E759A | 12 Months | 2020-04-09 | 2021-04-09 | 105.29
6 | 1430 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | D1387E763A | 12 Months | 2020-04-09 | 2021-04-09 | 105.29
7 | 1431 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | D1387E765A | 12 Months | 2020-04-09 | 2021-04-09 | 105.29
8 | 1432 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | D1387E766A | 12 Months | 2020-04-03 | 2021-04-03 | 105.29
9 | 1433 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | C1387E135M | 12 Months | 2020-04-08 | 2021-04-08 | 105.29
10 | 2351 | Joel Mamalateo Deduyo | 2489 UV / Visible Detector | Waters | 186248900 | D1887E275A | 12 Months | 2020-04-08 | 2021-04-08 | 105.29

## External
External Calibration for Equipment used in Waters@Kallang Place
April 2020 External Calibration
[8 empty] | SUM |  | 0
[8 empty] | 10% Support Cost |  | 0
[8 empty] | Subtotal |  | 0
[7 empty] | Convertion rate |  |  | 1.4338
S/N | EQID | OWNER | Equipment Description | MANUFACTURER | MODEL | SERIAL NO. | CAL INTERVAL | CAL DATE | PO | Cost USD | Cost SGD | VENDOR | Remarks
1 | 0500 | Joel Mamalateo Deduyo | Electronic Weigh Balance | AND | HR200 | 12326184 | 06 Months | 2020-04-15 | NA | 0 | 0 | Micro Precision | Customer Paid
2 | 2035 | Joel Mamalateo Deduyo | Electronic Weigh Balance | AND | HR200 | 12336741 | 06 Months | 2020-04-15 | NA | 0 | 0 | Micro Precision | Customer Paid
3 | 2092 | Joel Mamalateo Deduyo | Electronic Weigh Balance | AND | HR200 | 12337042 | 06 Months | 2020-04-15 | NA | 0 | 0 | Micro Precision | Customer Paid

## PM
Preventive Maintenance for Equipment used in Waters@Kallang Place
April 2020 Internal Calibration
[7 empty] | Sum |  |  | 1209.72
[7 empty] | SG&A |  |  | 60.486
[7 empty] | Profit |  |  | 127.0206
[7 empty] | Subtotal Internal Calibration |  |  | 1397.2266
S/N | PM EQID | OWNER | Equipment Description | MANUFACTURER | MODEL | SERIAL NO. | PM INTERVAL | PM DATE | DUE DATE | COST (US$) | Remarks
1 | PM 03319 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT091728 | J99TCM867R | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
2 | PM 03320 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT038000 | G01TCM633M | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
3 | PM 03321 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT038000 | MX4MM6882M | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
4 | PM 03322 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT091728 | MX5DM2250R | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
5 | PM 03323 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT038000 | H01TCM643M | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
6 | PM 03324 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT038000 | MX5MM4884M | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
7 | PM 03325 | Joel Mamalateo Deduyo | Temperature Control Module | Waters | WAT091728 | E97TCM153R | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
8 | PM 03326 | Joel Mamalateo Deduyo | Temperature Control Module 2 | Waters | 186002441 | A18TC2481G | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
9 | PM 03327 | Joel Mamalateo Deduyo | Temperature Control Module 2 | Waters | 186002441 | B18TC2495G | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
10 | PM 03328 | Joel Mamalateo Deduyo | Temperature Control Module 2 | Waters | 186002441 | F13TC2204G | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
11 | PM 03329 | Joel Mamalateo Deduyo | Temperature Control Module 2 | Waters | 186002441 | F13TC2180G | 12 Months | 2020-04-02 | 2021-04-02 | 33.4
12 | PM 03372 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT364A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
13 | PM 03373 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT361A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
14 | PM 03374 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | K09UPT215A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
15 | PM 03375 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT362A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
16 | PM 03371 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT366A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
17 | PM 03376 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT363A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
18 | PM 03377 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT360A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29
19 | PM 03378 | Joel Mamalateo Deduyo | TUV Detector | Waters | 186015028 | M09UPT359A | 12 Months | 2020-04-17 | 2021-04-17 | 105.29

## SENAI_
WATERS CONSIGN EQUIPMENT TRACKING RECORD | [7 empty] | Total Cost | 0
April 2020 SENAI Calibration | [6 empty] | 4.2086 |  | 1.3973
2020
S/N | EQID | DESCRIPTION | EQUIPMENT SERIAL NO. | CUSTOMER PROPERTY TAG | Cal Due Date | Schedule for 2020 | Price in MYR | Price in SGD | Price in USD | Invoice | Calibration Report | Cert
1 | NA | NA | NA | NA | NA | NA | NA | NA | 0 | NA | NA | NA
[6 empty] | TOTAL | 0 | 0 | 0
[7 empty] | . | . | .

## Parts
Spare Parts used in Waters@Kallang Place
April 2020 Calibration
[5 empty] | SUM | 1532.705
[5 empty] | COA | 153.2705
[5 empty] | Total sum | 1685.9755
S/N | Date | DESCRIPTION | Part no | BAAN No. | Quantity | Cost (US$)/EA | Remarks
1 | 2020-04-03 | MASK,M1,2998 | WAT-405002546 | WAT-405002546 | 1 | 2.66 | For 2998 Box build flowcell noise tester s/n: G08998943N (TE186002698) failed on noise test.
2 | 2020-04-03 | WEDGE, REFERENCE DIODE | WAT-405002559 | WAT-405002559 | 1 | 6.9 | For 2998 Box build flowcell noise tester s/n: G08998943N (TE186002698) failed on noise test.
3 | 2020-04-03 | PLATE,BEAMSPLITTER LC/ | WAT-402000201 | WAT-402000201 | 1 | 22 | For 2998 Box build flowcell noise tester s/n: G08998943N (TE186002698) failed on noise test.
4 | 2020-04-09 | PHOTODIODE S-1227-66 BQ MATCH | WAT-WAT081100 | WAT-WAT081100 | 1 | 45.65 | For 2489 box build 186248900 s/n: D1387E765A @ Final test 1 Port 3 due to noise failure during calibration.
5 | 2020-04-14 | ASSY, HB DEUTERIUM LAMP | WAT-402000351 | WAT-402000351 | 6 | 153.626 | For 2489 box build 186248900 s/n: H1387E735A, H1387E730A,H1387E737A,H1387E734A, C1387E135M,H1387E727A due to lamp expiry hours.(Exceed 2000 hrs)
6 | 2020-04-15 | INLINE FILTER ASSY | WAT-WAT035190 | WAT-WAT035190 | 5 | 13.4 | For CHM s/n: L97CHM711M, MX6JM7867M, MX6KM9444M, M03CHM015R and E01CHM203R as part of maintenance.
7 | 2020-04-16 | FAN ASSY, 4.62 IN, 85 CFM | WATH-WAT270361 | WATH-WAT270361 | 2 | 14.95 | For TE82526 and TE82526-2 due to fail to home.
8 | 2020-04-24 | PWR SUP DC 80W MULTI-OUTPUT RM | WATH-279003674 | WATH-279003675 | 1 | 40 | For 2489 box build (to test spare bench) 186002487 s/n:L0787E095M due to filter initialization failure.
9 | 2020-04-24 | POWER SUPPLY, HB LAMP HAMAMATS | WAT-279001105 | WAT-279001105 | 2 | 101.46 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: M09UPT361A due to noise failure and K09UPT215A due to filter initialization failure.
10 | 2020-04-28 | MIRROR, 3X ELLIPSOIDAL DT | WAT-402000170 | WAT-402000170 | 1 | 92 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to corrosion on the side of the body.
11 | 2020-04-28 | HOLDER WINDOW | WAT-WAT081336 | WAT-WAT081336 | 1 | 2.219 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to oxides found on mirror surface.
12 | 2020-04-28 | WINDOW, OPTICS COVER | WAT-WAT080244 | WAT-WAT080244 | 1 | 4.2 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to oxides found on mirror surface.
13 | 2020-04-28 | FILTER, 2ND ORDER | WAT-WAT081066 | WAT-WAT081066 | 1 | 3.52 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to oxides found on mirror surface.
14 | 2020-04-28 | FILTER, ERBIUM CALIBRATION | WAT-WAT081065 | WAT-WAT081065 | 1 | 32.96 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to oxides found on mirror surface.
15 | 2020-04-28 | FILTER WHEEL, 2487 | WAT-415000231 | WAT-415000231 | 1 | 1.02 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to broken holder.(wear and tear)
16 | 2020-04-28 | GRATING, 1800 G/MM,W HOLDER | WAT-WAT081308 | WAT-WAT081308 | 1 | 58 | For TUV box build p/n: 186005028 @ BSM / QSP Final Tester s/n: K09UPT215A due to oxides and dirt found on mirror surface.

## Others
Others (Fixture, repairs etc)
April 2020 Calibration
|  |  | SUM |  | 471.96
|  |  | 10% Support Cost |  | 47.196
|  |  | Subtotal Others |  | 519.156
|  |  | Convertion rate |  | 1.4338
S/N | Date | Element / Avtivity | Quantity | Cost SGD | Cost (US$) | Remarks | Attachment
1 | 2020-04-27 | PM for GES units (S/no: C1525P993A) | 1 | 338.3481 | 235.98 | 1. Man-hours – from collection/domestic cleaning/ testing process) – 3.5 hours - Replaced 8 CHECK VALVE CARTRIDGE p/n: WAT-289000144 - $10.95 each ($87.6 - Sum of 8 check valves) - Replaced 4 Assy PLUNGER, p/n: WAT-WAT207069 - $15.14 each ($60.56 - Sum of 4 Assy plunger) - Replaced 4 WASHER,.375X.250X.005,TFE p/n: WAT-WAT026572 - $0.08 each ($0.32 - Sum of 4 washer)
2 | 2020-04-30 | PM for GES units (S/no: C0925P516A) | 1 | 338.3481 | 235.98 | 1. Man-hours – from collection/domestic cleaning/ testing process) – 3.5 hours - Replaced 8 CHECK VALVE CARTRIDGE p/n: WAT-289000144 - $10.95 each ($87.6 - Sum of 8 check valves) - Replaced 4 Assy PLUNGER, p/n: WAT-WAT207069 - $15.14 each ($60.56 - Sum of 4 Assy plunger) - Replaced 4 WASHER,.375X.250X.005,TFE p/n: WAT-WAT026572 - $0.08 each ($0.32 - Sum of 4 washer)