| Audit Nonconformity (ANC) Number XNC-SGYYMMNN |  | INC-SG200702 |  | Assignee, Company (Supplier designated point of contact or WPPL Auditee) |  |  |  | Luo Hua Zhen, WPPL (Test Engineering Operations) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Section 1: Introduction (To be completed by Waters Supplier Quality / Quality Engineer) |  |  |  |  |  |  |  |  |  |
| Category |  | CNC MNC NC |  |  | Failure Type | Facility & Equipment Controls |  |  |  |
| Statement |  | There are multiple inaccuracies observed in the calibration records provided by Micro Precision Calibration. |  |  |  |  |  |  |  |
| Requirement Standard, regulation or document that is not met |  | ISO9001:2015 Cl. 8.4 Control of externally provided processes, products and services. 8.4.1 The organization shall ensure that externally provided processes, products and services conform to requirements. |  |  |  |  |  |  |  |
| Objective Evidence |  | Evidence 1: It was observed the calibration cert no. 551220083592482 dated 02 May 2020 contains multiples inaccuracies in the Notes section. Evidence 2: It was observed on calibration cert no. 551220083320957 dated 11 Nov 2019 for equipment ID# WPPL0018 S/N# 122994G, the calibration date is 11 Nov 2019 and calibration due date is 11 Nov 2020. However, on the physical calibration sticker pasted on equipment by the service provider, the calibration date is 05 Nov 2019 and calibration due date is 05 Nov 2020 |  |  |  |  |  |  |  |
| Initiated By (Waters Supplier Quality / Quality Engineer) |  |  | Name/ Title |  |  |  | Sign |  | Date |
|  |  |  | Bayu Darmawan/ ISO Consultant |  |  |  |  |  |  |
| Issuance Date | Due Date for Sections 2 and 3 (Should be within 90 days from issuance.) |  | Due Date for Section 2 (Should be within 21 days from issuance for internal NC and MNC; 3 days for Internal CNC.) |  |  |  | Remarks (if any): |  |  |
|  |  |  | N/A, Supplier Audit Nonconformity |  |  |  |  |  |  |

| Section 2: Root Cause Analysis and Interim Containment Actions |  |  |
| --- | --- | --- |
| Section 2A: Form a Team List Team Members’ Names/ Function (minimum 2 personnel) |  |  |
| Section 2B: Implemented Interim Containment Actions (Temporary fixes to protect customer, including Correction Actions. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included. If containment is documented via Waters NCR, Waters NCR number should be included). |  | Completion Date (DD.MM.YY) |

| Section 2C: Root Cause Analysis (RCA) and Risk Assessment (Objective evidence should be attached and referenced with running number e.g. Refer to Attachment 1). |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| RCA Category |  | Select from Dropdown |  |  | Sub-Category |  | Select from Dropdown |  |
| RCA Category 2 (if any) |  | Select from Dropdown |  |  | Sub-Category 2 (if any) |  | Select from Dropdown |  |
| Section 2d: Plan of Permanent Corrective Action(s) (Each root cause in section 2 shall be addressed. Documentation shall be updated as appropriate. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included). |  |  |  |  |  |  |  | Plan Date (DD.MM.YY) |
| Section 2E: Plan Action(s) to Prevent Recurrence (Shall prevent problems on similar products, processes, locations or families. Documentation shall be updated as appropriate. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included). |  |  |  |  |  |  |  | Plan Date (DD.MM.YY) |
| Submission By (by Audit Nonconformity Assignee) |  |  | Req? | Name/ Title |  | Sign |  | Date |
|  |  |  | X |  |  |  |  |  |
| RCA Approval | Quality Manager/ Quality Supervisor (of Assignee) |  | X |  |  |  |  |  |
|  | Waters Supplier Quality / Quality Engineer |  | X |  |  |  |  |  |
|  | Other |  |  |  |  |  |  |  |
| Remarks (if any): |  |  |  |  |  |  |  |  |

| Section 3A: Implementation of Permanent Corrective Action(s) (Attach evidence) |  |  | Completion Date (DD.MM.YY) |
| --- | --- | --- | --- |
| Section 3B: Action(s) to Prevent Recurrence (Attach evidence) |  |  | Completion Date (DD.MM.YY) |
| Submission By (by Assignee) | Name/ Title | Sign | Date |

| Section 4: Verification and Validation (Shall confirm that actions do not adversely affect the finished device. Examples: Refer to: ECO Number, Validation Protocol Number, Risk Management File. Where not applicable, document rationale). |  |  |  |  |  |  | Completion Date (DD.MM.YY) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Verification of Effectiveness: (Evidence reviewed to confirm effectiveness of actions; such as number of records monitored). |  |  |  |  | Acceptance Criteria: (Results that needs to be observed for actions to be deemed effective). |  | Result (Pass/ Fail) |
|  |  |  |  |  |  |  | Choose an item |
| Remarks (if any): Attached extension request WPPL-FM-25. |  |  |  |  |  |  |  |
| Submission By (by Assignee) |  |  | Req? | Name/ Title |  | Sign | Date |
|  |  |  | X |  |  |  |  |
| Approval |  | Quality Manager/ Quality Supervisor (of Assignee) | X |  |  |  |  |
|  |  | Management Rep/ QMS In-Charge (of Assignee) | X |  |  |  |  |
| CAPA Determination (by Waters SQE/ QE below. Ineffective actions with no CAPA issued should be justified) |  |  | No CAPA needed, Actions Effective. No CAPA needed, justification in remarks. CAPA number: CAPA- . |  |  |  |  |
| Approval | Waters Supplier Quality / Quality Engineer |  | Req? | Name/ Title |  | Sign | Date |
|  |  |  | X |  |  |  |  |
| Approval | Other |  |  |  |  |  |  |
| Final Approval | Waters Quality Manager / Designee) |  | X |  |  |  |  |
| Remarks (if any): Isolated occurrence of effectiveness fail. ANC shall be revised and submitted before due date. |  |  |  |  |  |  |  |