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Date: Jul 13, 2020 Cert No. 551220083507606
Customer:
WATERS PACIFIC PTE LTD
1 SCIENCE PARK ROAD,#02-08 THE CAPRICORN
SINGAPORE SCIENCE PARK I
SINGAPORE 117528 Work Order #: SG-11013648
MPC Control #: 14354681 Serial Number: 14354681
Asset ID: 2095 Department: N/A
Gage Type: LOAD BOX Performed By: SENGTHYE CHEW
Manufacturer: CHROMA Received Condition: OPERATIONAL
Model Number:  S14-03 Returned Condition: REPORT OF VALUE
Size: N/A Cal. Date: March 06, 2020
Temp/RH: 22.0°C/52.0% Cal. Interval: 12 MONTHS
Location: Calibration performed at MPC facility Cal. Due Date: March 06, 2021

Calibration Notes:

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of
Value.See attached data. (1 page)

Calibration performed for Flextronics Pte Ltd.

This supersedes Certificate Number 551220083573991

Standards Used to Calibrate Equipment

I.D. Description. Model Serial Manufacturer Cal. Due Date  Traceability #
CA8094 MULTIMETER 3458A 2823A25888 HEWLETT PACKARD  Jan 10,2021  1-12015829011-1

Procedures Used in this Event

Procedure Name Description
CHROMA 190XX PRODUCT Using a Load Box to Verify Hipot Tester
Calibrating Technician: QC Approval: &/&/VL/
(L~ .
NGTHYE CHEW JUN CALIMAG
STATEMENTS OF PASS OR FAIL CONFORMANCE: The uncertainty of measurement has been taken into account when determining compliance with specification, as per ILAC-G8:03/2009. All measurements and test results guard banded to ensure the

probability of false-accept does not exceed 2% in compliance with ANSI/NCSL Z540.3-2006 and in case without guard banded the p of fals pt on test L inty ratio.

THE CALIBRATION REPORT STATUS:

PASS- Term used when compliance statement is given, and the measurement result is PASS.

PASS?- Term used when compliance statement is given, and the measurement result is conditional passed or PASSZ.

FAIL- Term used when compliance statement is given, and the measurement result is FAIL.

FAILZ- Term used when compliance statement is given, and the measurement result is conditional failed or FAILZ.

REPORT OF VALUE - Term used when reported measurement is not requiring compliance statement in report.

ADJUSTED- When adjustments are made to an instrument which changes the value of measurement from what was measured as found to new value as left.
LIMITED - When an instrument fails calibration but is still functional in a limited manner.

The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%, unless otherwise stated. This
calibration report complies with ISO/IEC 17025:2017 and ANSI/NCSL Z540.3. Calibration cycles and resulting due dates were submitted/approved by the customer. Any number of factors may cause an instrument to drift out of tolerance before the next
scheduled calibration. Recalibration cycles should be based on frequency of use, environmental conditions and customer's established systematic accuracy. All standards are traceable to SI through the National Institute of Standards and Technology (NIST)
and/or recognized national or international standards laboratories. Services rendered include proper service ion and are for no less than thirty (30) days. The information on this report pertains only to the instrument identified,
this may not be reproduced in part or in a whole without the prior written approval of the issuing MP Calibration Laboratory.
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Calibration Report of Chroma S14-03 Load Box

MPC Control #: 14354681 Serial Number: 14354681
Asset ID: NA Calibration Date: March 06, 2020

Resistance Check

Function Tested Nominal Measured Value Correction Uncertainty ()
PASS 50 mQ 50.66 mQ -0.66 mQ 0.00070 mQ
PASS 450 kQ 448.02 kQ 1.98 kQ 0.0081 kQ
FAIL 150 mQ 149.97 mQ 0.03mQ 0.0021 mQ

FAIL 400 kQ 399.40 kQ 0.60 kQ 0.0072 kQ

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of Value.
The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds

to a coverage probability of approximately 95% , unless otherwise stated.

- End of Calibration Report -
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Revised Cal Report (EQID2096 - Load Box).pdf
MICRO PRECISION CALIBRATION PTE LTD g
/_I\ BLOCK 5008 ANG MO KIO AVENUE 5 #07-02/03 & \\\_/// 7

/,
\\

TECHPLACE Il SINGAPORE 569874 d

—
D PRECISION 65 67484384 il;\/,i ANMAB

"/
Ip R

e g . . ;//\\:ACCREDITED
Certificate of Calibration fr\\\\ P -
i\ AC-1969.15
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Customer:
WATERS PACIFIC PTE LTD
1 SCIENCE PARK ROAD.,#02-08 THE CAPRICORN
SINGAPORE SCIENCE PARK I
SINGAPORE 117528 Work Order #: SG-11014146
MPC Control #: 14354691 Serial Number: 14354691
Asset ID: 2096 Department: N/A
Gage Type: LOAD BOX Performed By: SENGTHYE CHEW
Manufacturer: CHROMA Received Condition: OPERATIONAL
Model Number:  S14-03 Returned Condition: REPORT OF VALUE
Size: N/A Cal. Date: May 02, 2020
Temp/RH: 22.0°C/52.0% Cal. Interval: 12 MONTHS
Location: Calibration performed at MPC facility Cal. Due Date: May 02, 2021

Calibration Notes:

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of
Value.See attached data. (1 page)

Calibration performed for Flextronics Pte Ltd.

This supersedes Certificate Number 551220083592482

Standards Used to Calibrate Equipment

I.D. Description. Model Serial Manufacturer Cal. Due Date  Traceability #
CA8094 MULTIMETER 3458A 2823A25888 HEWLETT PACKARD  Jan 10,2021  1-12015829011-1

Procedures Used in this Event

Procedure Name Description
CHROMA 190XX PRODUCT Using a Load Box to Verify Hipot Tester, Chroma 190XX Product, Jun-16-2016

Calibrating Technician: QC Approval: ?&/&/ﬂb
' NGT

STATEMENTS OF PASS OR FAIL CONFORMANCE: The uncertainty of measurement has been taken into account when determining compllance with specification, as per ILAC-G8:03/2009. All measurements and test results guard banded to ensure the
probability of false-accept does not exceed 2% in compliance with ANSI/NCSL Z540.3-2006 and in case without guard banded the p of fals pt on test L inty ratio.

THE CALIBRATION REPORT STATUS:

PASS- Term used when compliance statement is given, and the measurement result is PASS.

PASS?- Term used when compliance statement is given, and the measurement result is conditional passed or PASSZ.

FAIL- Term used when compliance statement is given, and the measurement result is FAIL.

FAILZ- Term used when compliance statement is given, and the measurement result is conditional failed or FAILZ.

REPORT OF VALUE - Term used when reported measurement is not requiring compliance statement in report.

ADJUSTED- When adjustments are made to an instrument which changes the value of measurement from what was measured as found to new value as left.
LIMITED - When an instrument fails calibration but is still functional in a limited manner.

The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%, unless otherwise stated. This
calibration report complies with ISO/IEC 17025:2017 and ANSI/NCSL Z540.3. Calibration cycles and resulting due dates were submitted/approved by the customer. Any number of factors may cause an instrument to drift out of tolerance before the next
scheduled calibration. Recalibration cycles should be based on frequency of use, environmental conditions and customer's established systematic accuracy. All standards are traceable to SI through the National Institute of Standards and Technology (NIST)
and/or recognized national or international standards laboratories. Services rendered include proper 's service ion and are for no less than thirty (30) days. The information on this report pertains only to the instrument identified,
this may not be reproduced in part or in a whole without the prior written approval of the issuing MP Calibration Laboratory.
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Calibration Report of Chroma S$14-03 Load Box

MPC Control #: 14354691 Serial Number: 14354691

Asset ID: 2096 Calibration Date: May 02, 2020

Resistance Check

Function Tested Nominal As Found As Left Uncertainty ()
PASS 50 mQ 50.00 mQ 50.00 mQ 0.00070 mQ
PASS 450 kQ 448.34kQ 448.34kQ 0.0081 kQ
FAIL 150 mQ 150.00 mQ 150.00 mQ 0.0021 mQ
FAIL 400 kQ 399.29kQ 399.29kQ 0.0072kQ

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of Value.
The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds
to a coverage probability of approximately 95% , unless otherwise stated.

- End of Calibration Report -

Confidential: Controlled Document. Page 1of 1
Cannot be reproduced without the approval of the Quality Department. TQDG-FM-002 rev04






image4.emf
WATERS (PCBA)_2020_FEB_MY53002751-Revised.pdf


WATERS (PCBA)_2020_FEB_MY53002751-Revised.pdf
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o I’jlll ‘1“\ CALIBRATION LABORATORY
! AC-1969.15
Date: February 18, 2020 Cert No. 551220083477955
Customer:
WATERS PACIFIC PTE LTD
1 SCIENCE PARK ROAD #02-08 THE CAPRICORN
SINGAPORE SCIENCE PARK Il
SINGAPORE 117528 Work Order #: SG-11013266
MPC Control #  MY53002751/95155 Serial Number: MY53002751/95155
Asset ID: N/A Department: N/A
Gage Type: 6 1/2 DIGIT MULTIMETER Performed By: YIOK TEE CHUA
Manufacturer: AGILENT Received Condition: IN TOLERANCE
Model Number:  34411A Returned Condition: IN TOLERANCE
Size: REFER DATA Cal. Date: February 18, 2020
Temp/RH: 224°C149% Cal. Interval: 12 MONTHS
Location: Calibration performed at Customer's facility Cal. Due Date: February 18, 2021
Calibration Notes:
Performance Tests Passed. See attached data (4 Pages).
Calibration performed at GES SINGAPORE PTE LTD.
Standards Used to Calibrate Equipment
1.D. Description. Model Serial Manufacturer Cal. Due Date  Traceability #
FA0003 CALIBRATOR 5500A 6940015 FLUKE May 28, 2020  551220083035835
DH3876 MULTI-FUNCTION CALIBRATOR 5700A/NAF 6020913 FLUKE Apr 9, 2020 551220082911745
Procedures Used in this Event
Procedure Name Description
AGILENT 34410A/11A Rev. FIFTH EDITION Digital Multimeter, Agilent 34410A/11A, Fifth Edition, Mar-01-2012

Calibrating Technician: QC Approval: %/’ﬂ/ /
( ? P

YIOK TEE CHUA JUN CALIMAG

of Pass or Fail Ci The of has been taken inlo account when i with as per ILAC-G8:03/2009. All measurements and lest results guard banded Lo ensure the
prabability of false-accept daes not exceed 2% in compliance with ANSIINCSL Z540.3-2008
The status of compliance with the acceptance criteria is reported as:
PASS - Compliant with specification:
FAIL - Not compliant with specification,
FAIL? - The measured value is not within the acceptance fimits. However, a portion of the expanded uncertainty of measurement at 85% is within the specified tolerance

PASS* - The measured value is within acceptance limits. However, a portion of the expanded uncerlainty of measurement at 95% exceeds the specified lolerance.

The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%. unless otherwise
stated. This calibration report complies with ISQYIEC 170252017 and ANSINCSL Z540.3 Method 6-Guard Bands based on Test Uncertainly Ratio. Calibration cycles and resulling due dates were submitied/approved by the cuslomer. Any number of
factors may cause an instrument ta drift out of tolerance before the next scheduled calibration. Recalibralion cycles should be based on frequency of use. environmental conditions and customer's y accuracy. All

are traceable to Sl through the National Institute of Standards and Technology (NIST) and/or national or standards Services rendered include proper manufacturer's service instruction and are warranted

for no less than thirty (30) days. The information on this report pertains anly to the instrument identified, this may netbe reproduced in part or in a whole without the prier written approval of the issuing MP Calibration Laboratory.

Page 1 of 1 (CERT, Rev 6)
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Measurement Report Agilent 34411A Digital Multimeter

1 year specifications

Procedure:

1D Number:

34411A AGILENT

MY53002751/95155

Order Number:

Calibration Date:

SG-11013266

18-02-20

Section 1 - Zero Offset Verification
Front panel input terminals

Function/Range Nominal pisicioi As Found As Left Upper Result] Uncertainty (%)
Limit Limit
DCV 7 100 mvV 0.0000 MV ~0.0035 mV 0.0004 mV 0.0004 mV 0.0035 mV__ | PASS 0.000058 mV
DCV/1V 0.000000 V 20.000007 V 0.000001 V 0.000001 V 0.000007 V| PASS| __ 0.00000058 V
DCV/10V 0.00000 V 20.00005 V 0.00000 V 0.00000 V 0.00005V | PASS 0.0000058 V
DCV /100 V 0.0000 V ~0.0006 V 0.0000 V 0.0000 V 0.0006 V PASS 0.000058 V
DCV / 1000 V 0.000 V ~0.006 V 20,001 V 20,001V 0.006 V PASS 0.00058 V
DCI/ 100 pA 0.0000 pA 0.0250 yA 0.0000 pA 0.0000 pA 0.0250 yA | PASS 0.000058 pA
DCI/ 1 mA 0.000000 mA | _-0.060000 mA | 0.000002 mA | 0.000002 mA | 0.060000 mA | PASS| _ 0.00000058 mA
DCI7 10 mA 0.00000 mA | -0.00200 mA | 0.00000 mA | _0.00000 mA 0.00200 mA__| PASS| __ 0.0000058 mA
DCI/ 100 mA 0.0000 mA -0.0050 mA 0,0000 mA 0.0000 mA 0.0050 mA | PASS 0.000058 mA
DCI/1 A 0.000000 A 20.000100 A 0.000000 A 0.000000 A 0.000100 A | PASS| _ 0.00000058 A
DCI/3 A 0.00000 A 20.00060 A 0.00000 A 0.00000 A 0.00060 A | PASS 0.0000058 A
4-Wire / 100 Q 0.0000 Q -0.0040 0 0.0006 Q 0.0006 Q 0.0040 Q PASS 0.000058 Q
4-Wire / 1 kQ 0.000000 kQ | -0.000010 kQ | 0.000001 k@ | 0.000001 kQ 0.000010 kO | PASS| _ 0.00000058 kQ
3-Wire / 10 kQ 0.00000 kQ 0.00010 kQ 0.00000 kQ 0.00000 kO 0.00010kQ | PASS| __ 0.0000058 kQ
4-Wire / 100 kQ 0.0000 kQ 20.0010 kQ 0.0000 kQ 0.0000 kQ 0.0010kQ | PASS 0.000058 kQ
4-Wire /1MQ | _0.000000 MQ | -0.000010 MQ | _0.000000 MQ | 0.000000 MQ | _0.000010 MQ | PASS| _ 0.00000058 MQ
2-Wire /10 MQ_|__0.00000 MQ_|_-0.00010 MQ_|__0.00000 MQ | _0.00000 MQ 0.00010 MQ | PASS| __ 0.0000058 MQ
4-Wire / 100 MO | 0.0000 MQ -0.0010 MQ 0.0000 MQ 0.0000 MQ 0.0010 MQ | PASS 0.000058 MQ
2-Wire / 100 Q 0.0000 Q 20,0040 O 0.0000 O 0.0000 Q 0.0040 Q PASS 0.000058 Q
2-Wire / 1 kQ 0.000000 kO | -0.000010 kQ | _0.000000 kG | _0.000000 kQ 0.000010 kO | PASS| __ 0.00000058 kQ
2-Wire / 10 kQ 0.00000 kQ -0.00010 kQ 0.00000 kQ 0.00000 kQ 0.00010 kQ | PASS| __ 0.0000058 kQ
2-Wire /100 kQ 0.0000 kQ 20,0010 kO 0.0000 kQ 0.0000 kQ 0.0010kQ | PASS 0.000058 kQ
2.Wire/1MQ | 0.000000 MQ | -0.000070 MQ | 0.000000 MO | _0.000000 MQ | 0.000010 MQ__ | PASS| _ 0.00000058 MQ
2-Wire / 10 MQ_|__0.00000 MO | _-0.00070 MQ_|_0.00000 MQ | _0.00000 MQ 0.00010 MQ | PASS| __ 0.0000058 MQ
2-Wire / 100 MQ | 0.0000 MQ -0.0010 MQ 0.0000 MQ 0.0000 MQ 0.0010MQ | PASS 0.000058 MQ
Page 1 of 4
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Section 1 - Zero Offset Verification ( continued)
Rear panel input terminals

Function/Range Nominal LL?:;r As Found As Left Li[:a:ir Result Uncertainty (%)
DCV /100 mV 0.0000 mV -0.0035 mV 0.0002 mV 0.0002 mV 0.0035 mV PASS 0.000058 mV
DCV/1V 0.000000 V -0.000007 V 0.000001 V 0.000001 Vv 0.000007 V PASS 0.00000058 V
DCV/10V 0.00000 V -0.00005 V 0.00000 V 0.00000 V 0.00005 V PASS 0.0000058 V
DCV /100 V 0.0000 V -0.0006 V 0.0000 V 0.0000 V 0.0006 V PASS 0.000058 V
DCV /1000 V 0.000 V -0.006 V 0.000 V 0.000 V 0.006 V PASS 0.00058 V
DCI/ 100 pA 0.0000 pA -0.0250 pA 0.0006 pA 0.0006 pA 0.0250 pA PASS 0.000058 pA
DCI /1 mA 0.000000 mA -0.060000 mA 0.000002 mA 0.000000 mA 0.060000 mA PASS 0.00000058 mA
DCI/ 10 mA 0.00000 mA -0.00200 mA 0.00002 mA 0.00000 mA 0.00200 mA PASS 0.0000058 mA
DCI / 100 mA 0.0000 mA -0.0050 mA -0.0001 mA -0.0001 mA 0.0050 mA PASS 0.000058 mA
DCII1A 0.000000 A -0.000100 A 0.000000 A 0.000000 A 0.000100 A PASS 0.00000058 A
DCI/3A 0.00000 A -0.00060 A 0.00000 A 0.00000 A 0.00060 A PASS 0.0000058 A
4-Wire /100 Q 0.0000 Q -0.0040 O 0.0005 Q 0.0005 Q 0.0040 Q PASS 0.000058 Q
4-Wire [ 1 kQ 0.000000 kQ -0.000010 kQ 0.000000 kQ 0.000000 kO 0.000010 kQ PASS 0.00000058 kQ
4-Wire / 10 kQ 0.00000 kQ -0.00010 kQ 0.00000 kQ 0.00000 kQ 0.00010 kQ PASS 0.0000058 kQ
4-Wire / 100 kQ 0.0000 kQ -0.0010 kQ 0.0000 kQ 0.0000 kQ 0.0010 kQ PASS 0.000058 kQ
4-Wire / 1 MQ 0.000000 MQ -0.000010 MQ 0.000000 MQ 0.000000 MQ 0.000010 MQ PASS 0.00000058 MQ
4-Wire / 10 MQ 0.00000 MQ -0.00010 MQ 0.00000 MQ 0.00000 MQ 0.00010 MQ PASS 0.0000058 MQ
4-Wire / 100 MQ 0.0000 MQ -0.0010 MQ 0.0000 MQ 0.0000 MQ 0.0010 MQ PASS 0.000058 MQ
2-Wire / 100 Q 0.0000 Q -0.0040 O 0.0000 Q 0.0000 Q 0.0040 Q PASS 0.000058 Q
2-Wire / 1 kQ 0.000000 kQ -0.000010 kQ 0.000000 kQ 0.000000 kQ 0.000010 kQ PASS 0.00000058 kQ
2-Wire / 10 kQ 0.00000 kQ -0.00010 kQ 0.00000 kQ 0.00000 kQ 0.00010 kQ PASS 0.0000058 kQ
2-Wire / 100 kQ 0.0000 kQ -0.0010 kQ 0.0000 kQ 0.0000 kQ 0.0010 kQ PASS 0.000058 kQ
2-Wire / 1 MQ 0.000000 MQ -0.000010 MQ 0.000000 MQ 0.000000 MQ 0.000010 MQ PASS 0.00000058 MQ
2-Wire / 10 MQ 0.00000 MQ -0.00010 MQ 0.00000 MQ 0.00000 MQ 0.00010 MQ PASS 0.0000058 MQ
2-Wire / 100 MQ 0.0000 MQ -0.0010 MQ 0.0000 MQ 0.0000 MQ 0.0010 MQ PASS 0.000058 MQ
Section 2 - DC Voltage Gain Verification Test
. Lower Upper :
Range Nominal Limit As Found As Left Limit Result| Uncertainty ( £ )
100 mV 100.0000 mV 99.9915 mV 99.9990 mV 99.9990 mV 100.0085 mV [ PASS 0.0017 mV
100 mV -100.0000 mV -100.0085 mV -99.9990 mV -99.9990 mV -99.9915 mV PASS 0.0017 mV
1V 1.000000 V 0.999958 V 0.999980 V 0.999980 V 1.000042 V PASS 0.0000098 V
10V 10.00000 V 9.99965 V 9.99986 V 9.99986 V 10.00035 V PASS 0.000043 V
10V -10.00000 V -10.00035 V -9.99989 V -9.99989 V -9.99965 V PASS 0.000043 V
100 V 100.0000 V 99.9954 V 99.9973V 99.8973 V 100.0046 V PASS 0.00065 V
1000 V 1000.000 V 999.954 V 999.980 V 999.980 V 1000.046 V PASS 0.0061 V
Section 3 - DC Current Gain Verification Test
i Lower Upper .
Range Nominal Limit As Found As Left Limit Result Uncertainty ( )
100 pA 100.0000 pA 99.9250 pA 100.0000 pA 100.0000 pA 100.0750 pA PASS 0.010 pA
1 mA 1.000000 mA 0.999440 mA 0.999900 mA 0.999900 mA 1.000560 mA PASS 0.000043 mA
10 mA 10.00000 mA 9.99300 mA 10.00000 mA 10.00000 mA 10.00700 mA PASS 0.00041 mA
100 mA 100.0000 mA 99.9450 mA 100.0001 mA 100.0001 mA 100.0550 mA PASS 0.0053 mA
1A 1.000000 A 0.998900 A 0.999900 A 0.999900 A 1.001100 A PASS 0.000094 A
3A 2.00000 A 1.99640 A 1.99980 A 1.99980 A 2.00360 A PASS 0.00018 A
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Section 4 - AC Voltage Gain Verification Test

Range Nominal Lo_wc'er As Found As Left U?p?r Result] Uncertainty (%)
Frequency Limit Limit
100 mV / 1 kHz 100.0000 mV 99.9100 mV 99.9750 mV 99.9750 mV 100.0900 mV PASS 0.015 mV
100 mV / 50 kHz 100.0000 mV 99.8500 mV 99.9994 mV 99.9994 mV 100.1500 mV PASS 0.027 mV
100 mV /300 kHz | 100.0000 mV 98.3000 mV 99.9955 mV 99.9955 mV 101.7000 mV PASS 0.1 mV
1V/1kHz 1.000000 V 0.999100 V 0.999801 V 0.999801 V 1.000900 V PASS 0.000053 V
1V /50 kHz 1.000000 V 0.998500 V 1.000433 V 1.000433 V 1.001500 V PASS 0.000085 V
1V /300 kHz 1.000000 V 0.983000 V 0.894578 V 0.994578 V 1.017000 V PASS 0.00050 V
10V /10 Hz 10.00000 V 9.99097 V 9.99712 V 9.99712 V 10.00903 V PASS 0.0028 V
10V /1kHz 10.00000 V 9.99100 V 10.00326 V 10.00326 V 10.00900 V PASS 0.00050 V
10 V / 50 kHz 10.00000 V 9.98500 V 9.99384 V 9.99384 V 10.01500 V PASS 0.00085 V
10 V / 300 kHz 10.00000 V 9.83000 V 10.00344 V 10.00344 V 10.17000 V PASS 0.0034 V
100V /1 kHz 100.0000 V 99.9100 V 100.0120 V 100.0120 V 100.0900 V PASS 0.0046 V
100 V / 50 kHz 100.0000 V 99.8500 V 100.0211 V 100.0211 V 100.1500 V PASS 0.0076 V
100 V / 300 kHz 50.0000 V 48.9000 V 49.9805V 49.9805 V 51.1000 V PASS 0.015V
750V /1 kHz 750.000 V 749.010 V 750.371 V 750.371 V 750.990 V PASS 0.056 V
750 V / 50 kHz 210.000 V 209415V 210122V 210122V 210.585V PASS 0.016 V
750 V / 300 kHz 70.000 V 65.410V 71.365V 71.365V 74.590 V PASS 0.020 V
Section 5 - AC Current Gain Verification Test
Range Nominal Lowee As Found As Left Uppar Result| Uncertainty (%)
Frequency Limit Limit
100 pA / 1 kHz 100.0000 A 99.8600 pA 100.0064 pA 100.0064 LA 100.1400 pA | PASS 0.020 pA
100 pA / 5 kHz 100.0000 pA 99.8600 pA 100.0620 pA 100.0620 pA 100.1400 pA PASS 0.040 pA
1mA/1kHz 1.000000 mA 0.998600 mA 1.000810 mA 1.000810 mA 1.001400 mA PASS 0.00016 mA
1 mA/5kHz 1.000000 mA 0.998600 mA 1.000766 mA 1.000766 mA 1.001400 mA PASS 0.00033 mA
10 mA / 1 kHz 10.00000 mA 9.98600 mA 10.00811 mA 10.00811 mA 10.01400 mA PASS 0.0016 mA
10 mA/5kHz 10.00000 mA 9.97600 mA 9.99730 mA 9.99730 mA 10.02400 mA PASS 0.0028 mA
100 mA /1 kHz 100.0000 mA 99.8600 mA 100.0759 mA 100.0759 mA 100.1400 mA PASS 0.015 mA
100 mA /5 kHz 100.0000 mA 99.7600 mA 99.9280 mA 99.9280 mA 100.2400 mA PASS 0.032 mA
1A/ 1kHz 1.000000 A 0.998600 A 0.999962 A 0.999962 A 1.001400 A PASS 0.00030 A
1A/5kHz 1.000000 A 0.998600 A 1.000270 A 1.000270 A 1,001400 A PASS 0.00052 A
3A/1kHz 2.00000 A 1.99520 A 2.00000 A 2.00000 A 2.00480 A PASS 0.00056 A
3A/5kHz 2.00000 A 1.99520 A 1.99921 A 1.99921 A 2.00480 A PASS 0.00096 A
Section 6 - Frequency Gain Verification Test
Input Voltage Nominal lﬁ:?‘r As Found As Left llj_?:?tr Result] Uncertainty ( )
10 mV 40.00000 Hz 39.97200 Hz 39.98922 Hz 39.99922 Hz 40.02800 Hz PASS 0.0000058 Hz
10V 10.00000 kHz 9.99930 kHz 9.99991 kHz 9.99991 kHz 10.00070 kHz PASS 0.0000058 kHz
Section 7 - Ohms Gain Verification Test
4-Wire Ohms, @OOMohms Verify only. 2-Wire Ohms.
Range Nominal lﬁ:‘;r As Found As Left l'll__’i’rf‘?tr Result] Uncertainty (£)
100 O 100.0000 Q 99.9860 O 100.0000 Q 100.0000 Q 100.0140 OQ PASS 0.0011Q
1kQ 1.000000 kQ 0.999890 kQ 1.000002 kQ 1.000002 kQ 1.000110 kQ PASS 0.0000090 kQ
10 kQ 10.00000 kQ 9.99890 kQ 9.99997 kQ 9.99997 kQ 10.00110 kQ PASS 0.000090 kQ
100 kQ 100.0000 kQ 99.9890 kQ 100.0000 kQ 100.0000 kQ 100.0110 kQ PASS 0.0014 kQ
1MQ 1.000000 MQ 0.999870 MQ 0.999952 MQ 0.999952 MQ 1.000130 MQ PASS 0.000024 MQ
10 MQ 10.00000 MQ 9.99590 MQ 9.99947 MQ 9.99947 MQ 10.00410 MQ PASS 0.00043 MQ
100 MQ 100.00000 MQ | 99.999199 MQ | 99.999820 MQ | 99.999820 MQ 100.000801 MQ [ PASS 0.00084 MQ

\0% \2/’3/9‘0
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ID Number;

MY53002751/95155

Section 8 - Capacitance Measurement

. Lower Upper "
Range Nominal Limit As Found As Left Limit Result Uncertainty ( £ )
1.0000 pyF 1.0000 pF 0.9950 pF 1.0010 pF 1.0010 pF 1.0050 pF PASS 0.0040 pF
10.000 pF 1.000 pF 0.986 pF 1.001 pF 1.001 pF 1.014 pF PASS 0.0040 pF
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MICRO PRECISION CALIBRATION PTE LTD g
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Certificate of Calibration fr\\\\ P -
i\ AC-1969.15
Date: Jul 13, 2020 Cert No. 551220083694803
Customer:
WATERS PACIFIC PTE LTD
1 SCIENCE PARK ROAD.,#02-08 THE CAPRICORN
SINGAPORE SCIENCE PARK I
SINGAPORE 117528 Work Order #: SG-11014146
MPC Control #: 7322788 Serial Number: 7322788
Asset ID: 2097 Department: N/A
Gage Type: LOAD BOX Performed By: SENGTHYE CHEW
Manufacturer: CHROMA Received Condition: OPERATIONAL
Model Number:  S14-03 Returned Condition: REPORT OF VALUE
Size: N/A Cal. Date: May 02, 2020
Temp/RH: 22.0°C/52.0% Cal. Interval: 12 MONTHS
Location: Calibration performed at MPC facility Cal. Due Date: May 02, 2021

Calibration Notes:

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of Value.
See attached data. (1 page)

Calibration performed for Flextronics Pte Ltd.

This supersedes Certificate Number 551220083592483

Standards Used to Calibrate Equipment

I.D. Description. Model Serial Manufacturer Cal. Due Date  Traceability #
CA8094 MULTIMETER 3458A 2823A25888 HEWLETT PACKARD  Jan 10,2021  1-12015829011-1
Procedures Used in this Event

Procedure Name Description
CHROMA 190XX PRODUCT Using a Load Box to Verify Hipot Tester, Chroma 190XX Product, Jun-16-2016

Calibrating Technician: QC Approval: ?&/&/ﬂb
' NGT

STATEMENTS OF PASS OR FAIL CONFORMANCE: The uncertainty of measurement has been taken into account when determining compllance with specification, as per ILAC-G8:03/2009. All measurements and test results guard banded to ensure the
probability of false-accept does not exceed 2% in compliance with ANSI/NCSL Z540.3-2006 and in case without guard banded the p of fals pt on test L inty ratio.

THE CALIBRATION REPORT STATUS:

PASS- Term used when compliance statement is given, and the measurement result is PASS.

PASS?- Term used when compliance statement is given, and the measurement result is conditional passed or PASSZ.

FAIL- Term used when compliance statement is given, and the measurement result is FAIL.

FAILZ- Term used when compliance statement is given, and the measurement result is conditional failed or FAILZ.

REPORT OF VALUE - Term used when reported measurement is not requiring compliance statement in report.

ADJUSTED- When adjustments are made to an instrument which changes the value of measurement from what was measured as found to new value as left.
LIMITED - When an instrument fails calibration but is still functional in a limited manner.

The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%, unless otherwise stated. This
calibration report complies with ISO/IEC 17025:2017 and ANSI/NCSL Z540.3. Calibration cycles and resulting due dates were submitted/approved by the customer. Any number of factors may cause an instrument to drift out of tolerance before the next
scheduled calibration. Recalibration cycles should be based on frequency of use, environmental conditions and customer's established systematic accuracy. All standards are traceable to SI through the National Institute of Standards and Technology (NIST)
and/or recognized national or international standards laboratories. Services rendered include proper 's service ion and are for no less than thirty (30) days. The information on this report pertains only to the instrument identified,
this may not be reproduced in part or in a whole without the prior written approval of the issuing MP Calibration Laboratory.

Page 1 of 1 (CERT, Rev 7)
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Calibration Report of Chroma S$14-03 Load Box

MPC Control #: 7322788 Serial Number: 7322788

Asset ID: 2097 Calibration Date: May 02, 2020

Resistance Check

Function Tested Nominal As Found As Left Uncertainty ()
PASS 50 mQ 51.00 mQ 51.00 mQ 0.00070 mQ
PASS 450 kQ 448.26 kQ 448.26 kQ 0.0081 kQ
FAIL 150 mQ 149.00 mQ 149.00 mQ 0.0021 mQ
FAIL 400 kQ 399.61kQ 399.61kQ 0.0072kQ

The user shall determine the suitability of the equipment for it's intended use. The calibration status is defined as Report of Value.
The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds
to a coverage probability of approximately 95% , unless otherwise stated.

- End of Calibration Report -

Confidential: Controlled Document. Page 1of 1
Cannot be reproduced without the approval of the Quality Department. TQDG-FM-002 rev04
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