| Audit Nonconformity (ANC) Number XNC-SGYYMMNN |  | SNC-SG200701 |  | Assignee, Company (Supplier designated point of contact or WPPL Auditee) |  |  |  | Amy Li, Micro Precision Calibration Pte Ltd |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Section 1: Introduction (To be completed by Waters Supplier Quality / Quality Engineer) |  |  |  |  |  |  |  |  |  |
| Category |  | CNC MNC NC |  |  | Failure Type | Production & Process Controls |  |  |  |
| Statement |  | Inconsistency of calibration procedure used for C011477 |  |  |  |  |  |  |  |
| Requirement Standard, regulation or document that is not met |  | ISO 17025: 2017 Cl 7.2.1.3 The laboratory shall ensure that it uses the latest valid version of a method unless it is not appropriate or possible to do so. When necessary, the application of the method shall be supplemented with additional details to ensure consistent application. |  |  |  |  |  |  |  |
| Objective Evidence |  | Tektronic Equipment ID C011477 was calibrated using 2 different set of standards in 2 cycle. Technician SengThye Chew used standard CT2147 in Year 2020 calibration that was not called out in Procedure Tektronix TDS3000C. The calibration cert id affected are 551220082926506 and 551200831518677. |  |  |  |  |  |  |  |
| Initiated By (Waters Supplier Quality / Quality Engineer) |  |  | Name/ Title |  |  |  | Sign |  | Date |
|  |  |  | Lim Shaynee / Supplier Quality Engineer |  |  |  |  |  | 03.Aug.2020 |
| Issuance Date | Due Date for Sections 2 and 3 (Should be within 90 days from issuance.) |  | Due Date for Section 2 (Should be within 21 days from issuance for internal NC and MNC; 3 days for Internal CNC.) |  |  |  | Remarks (if any): N/A |  |  |
| 03.Aug.2020 | 01.Nov.2020 |  | N/A, Supplier Audit Nonconformity |  |  |  |  |  |  |

| Section 2: Root Cause Analysis and Interim Containment Actions |  |  |
| --- | --- | --- |
| Section 2A: Form a Team List Team Members’ Names/ Function (minimum 2 personnel) | Rheabelle Acuin - Regional Quality Manager-Asia/ Corporate Quality Administrator Joel de Vera - Laboratory Manager Jun Calimag - Corporate Technical and Quality Development Senior Officer |  |
| Section 2B: Implemented Interim Containment Actions (Temporary fixes to protect customer, including Correction Actions. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included. If containment is documented via Waters NCR, Waters NCR number should be included). |  | Completion Date (DD.MM.YY) |
| Re-issue revised Calibration Certificate to reflect the correct standard equipment used during calibration based on the calibration procedure. Cert No. 551220083733159 Completed by: Jun Calimag/ Corporate Technical and Quality Development Senior Officer Completed By: Jun Calimag |  | 27.Jul.2020 |

| Section 2C: Root Cause Analysis (RCA) and Risk Assessment (Objective evidence should be attached and referenced with running number e.g. Refer to Attachment 1). |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| The calibration technician inadvertently added the standard (CT2147) for the calibration of Tektronix TDS 3054C Digital Oscilloscope that is not declared in the procedure used Tektronix TDS3000C. Likewise, this standard declared was overlooked by the QC personnel during the Quality checking process such as checking the procedure use, the declared standard and comparing the calibration certificate from the previous certificate. Refer to Attachment 1 |  |  |  |  |  |  |  |  |
| RCA Category |  | Corrective & Preventive Actions |  |  | Sub-Category |  | Man |  |
| RCA Category 2 (if any) |  | Select from Dropdown |  |  | Sub-Category 2 (if any) |  | Select from Dropdown |  |
| Section 2d: Plan of Permanent Corrective Action(s) (Each root cause in section 2 shall be addressed. Documentation shall be updated as appropriate. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included). |  |  |  |  |  |  |  | Plan Date (DD.MM.YY) |
| Conducted a refresher training for all technician about the process of declaring the standard to be used in CMS database during calibration and the QC personnel about the QC process to be particular in the recommended standard declared in the procedure used. Completed By: Joel De Vera/Laboratory Manager Rheabelle Acuin/ Regional Quality Manager |  |  |  |  |  |  |  | 27-28.July. 2020 |
| Section 2E: Plan Action(s) to Prevent Recurrence (Shall prevent problems on similar products, processes, locations or families. Documentation shall be updated as appropriate. Where applicable, Part Number(s), Part Description(s), Quantities, Action Owner should be included). |  |  |  |  |  |  |  | Plan Date (DD.MM.YY) |
| Not Applicable |  |  |  |  |  |  |  | Not Applicable |
| Submission By (by Audit Nonconformity Assignee) |  |  | Req? | Name/ Title |  | Sign |  | Date |
|  |  |  | X | Joel de Vera - Laboratory Manager |  |  |  | 04.Aug.2020 |
| RCA Approval | Quality Manager/ Quality Supervisor (of Assignee) |  | X | Rheabelle Acuin - Regional Quality Manager-Asia/ Corporate Quality Administrator |  |  |  | 04.Aug.2020 |
|  | Waters Supplier Quality / Quality Engineer |  | X |  |  |  |  |  |
|  | Other |  |  |  |  |  |  |  |
| Remarks (if any): |  |  |  |  |  |  |  |  |

| Section 3A: Implementation of Permanent Corrective Action(s) (Attach evidence) |  |  | Completion Date (DD.MM.YY) |
| --- | --- | --- | --- |
| Refer to the attached Training Records (Attachment 2) |  |  | 27-28.July.2020 |
| Section 3B: Action(s) to Prevent Recurrence (Attach evidence) |  |  | Completion Date (DD.MM.YY) |
| Not Applicable |  |  |  |
| Submission By (by Assignee) | Name/ Title | Sign | Date |
|  | Joel de Vera - Laboratory Manager |  | 04.Aug.2020 |

| Section 4: Verification and Validation (Shall confirm that actions do not adversely affect the finished device. Examples: Refer to: ECO Number, Validation Protocol Number, Risk Management File. Where not applicable, document rationale). |  |  |  |  |  |  | Completion Date (DD.MM.YY) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Verification of Effectiveness: (Evidence reviewed to confirm effectiveness of actions; such as number of records monitored). |  |  |  |  | Acceptance Criteria: (Results that needs to be observed for actions to be deemed effective). |  | Result (Pass/ Fail) |
|  |  |  |  |  |  |  | Choose an item |
| Remarks (if any): Attached extension request WPPL-FM-25. |  |  |  |  |  |  |  |
| Submission By (by Assignee) |  |  | Req? | Name/ Title |  | Sign | Date |
|  |  |  | X |  |  |  |  |
| Approval |  | Quality Manager/ Quality Supervisor (of Assignee) | X |  |  |  |  |
|  |  | Management Rep/ QMS In-Charge (of Assignee) | X |  |  |  |  |
| CAPA Determination (by Waters SQE/ QE below. Ineffective actions with no CAPA issued should be justified) |  |  | No CAPA needed, Actions Effective. No CAPA needed, justification in remarks. CAPA number: CAPA- . |  |  |  |  |
| Approval | Waters Supplier Quality / Quality Engineer |  | Req? | Name/ Title |  | Sign | Date |
|  |  |  | X |  |  |  |  |
| Approval | Other |  |  |  |  |  |  |
| Final Approval | Waters Quality Manager / Designee) |  | X |  |  |  |  |
| Remarks (if any): Isolated occurrence of effectiveness fail. ANC shall be revised and submitted before due date. |  |  |  |  |  |  |  |