## Waters Equipment Loc
Person In Charge : Joel
Locations | STATIONS | Customer ID | Equipment | Maker | Model | Serial No. | Calibration Procedure | Calibration Date | Calibration Cycle (Months/ Days) | Calibration Due Date | Calibration By Internal | Calibration By External | Calibration Report No | Reason of Failure and Length of downtime | Replacement Item / Spare Parts | Replacement Cost (US$) | STATUS | TEST/PROCESS | Cycle Time (mins) | % YEILD @ FINAL TEST STATION ONLY | Number of common testers | Criticalilty | Risk
eAlliance Sub Assembly Stations | 2690 eAlliance Front Panel Cell | TE82546 | 2690 eAlliance Front Panel Test Station | Waters | eAlliance | TE82546 | NA | NA | NA | NA | NA | NA | NA | 06 Mar 18 - 06 Mar 18 Failed to detect LCD panel | Replaced the RS type micro switch | consigned | In-Use | TEST | 15 | 98% | 1 | 3 | 3
| Pump Head Tester | TE 82519 | 2690 Pump Head Tester | WATERS | eAlliance | TE 82519 | 069004 | 2017-12-02 | 183 | 2018-06-03 | Shamsul | NA | CR/530 |  |  |  | In-Use | TEST | 10 | 90% | 1 | 3 | 3
| eAlliance Pump Tray 1 | TE82526-1 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 3
|  | 90891 | Electronic Weighing Balance | WATERS | HR200 | 12305757 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | MDF120074-3 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance Pump Tray 2 | TE82526 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 2
|  | 95206 | Electronic Weighing Balance | AND | HR200 | 12316922 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF120074-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance Pump Tray 3 | TE82526-4 | 2690 Pump Tray Tester |  |  | TE82526-4 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 2
|  | 90038 | Electronic Weighing Balance |  |  | 12200404 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF140265-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance HPMV Cell | TE 82625-1 | HPMV Tester | WATERS | NA | TE 82625-1 | 735000288 | 2018-01-19 | 184 | 2018-07-22 | Shamsul | NA | CR/561 | 13 Apr 18 - 13 Apr 18 Failed on ramp and decay test | Replaced interface tubing(external and internal tubing) Replaced inline filter p/n:WAT035190 | consigned $13.20 | In-Use | TEST | 10 | 90% | 1 | 3 | 3
| 2690/2790 Syringe Drive Cell | TE82656 | 2690/2790 Syringe Drive Tester | Waters | eAlliance | TE82656 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 10 | 95% | 1 | 3 | 2
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-1 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-2 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-2 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3
| Fluidics Tray Tester | TE82523 | Fluidics Tray Tester |  |  | TE82523 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 5 | 90% | 1 | 3 | 3
| Injector Seal Pack Cell/Engineering Lab | TE 82503-04 | Injector Seal Pack Tester | Waters | NA | TE 82503-04 | WAT043816 | 2018-01-18 | 184 | 2018-07-21 | Shamsul | NA | CR/562 |  |  |  | In-Use | TEST | 15 | 85% | 2 | 3 | 3
|  | TE 82503-05 | Injector Seal Pack Tester |  | NA | TE 82503-05 | WAT043816 | 2018-03-02 | 184 | 2018-09-02 | Shamsul | NA | CR/562 | 02 Mar 18 - 02 Mar 18 Failed on decay test | Replaced flow valve (FV2) p/n:ASFVO | consigned | In-Use | TEST | 15 | 85% | 2 | 3 | 3
| eAlliance Heater Chiller Cell | TE82620 | Heater Chiller Tester | Waters | NA | TE 82620 | 735000206 | 2013-03-15 | 365 | 2014-03-15 | Shamsul | NA | CR/454 | Inactive since 14 Mar 14 (ref no: RE/040/14) |  |  | Not-In-Use | TEST | 45 | 95% | 1 | 3 | 3
|  | NA | Multimeter | FLUKE | 110 | 85730619 | NA | 2018-01-05 | 365 | 2019-01-05 | NA | SGS | NA | [7 empty] | 1 | 3 | 3
|  | NA | RS 232 C Precision DC Module | Keithley | KNM-RTD32 | 000693213 | NA | 2010-11-19 | 60 Mths | 2015-11-19 | NA | ESSCO | CR/422086 | Inactive since 14 Mar 14 (ref no: RE/130/15) | [6 empty] | 1 | 3 | 3
eAlliance Final Test | Alliance Programming Station | NA | Autoloader 2 |  |  | Lace 32 B08L32122N | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 25 | 95% |  | 1
| Rear Panel Station | TE82573-1 | NA |  |  | Lace 32 C08L32093W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 7 | 95% |  | 1
| eAlliance Sample Station | 90913 | Electronic Weighing Balance | WATERS | HR200 | 12315717 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 5 | 95% | 6 | 2 | 2
eAlliance Final Test Station 1 | eAlliance Final Test 1 CPU | IBM | CPU | KKPRA0M | CPU | C12L32255W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% | 5 | 1 | 1
| eAlliance Final Test 1 | 81029 | Temperature Control Module | WATERS | NA | H01TCM643M | 735001271 | 2018-01-23 | 365 | 2019-01-23 | Shamsul | NA | CR/572 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 1 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E726A | WAT-RW-SKI01717 | 2018-04-06 | 365 | 2019-04-06 | Shamsul | NA | CR/514 | 23 Mar 18 - 23 Mar 18 Communication Failure 06 Apr 18 to 06 Apr 18 Lamp expiry hours | Replaced PCB p/n:210000355 Replaced lamp p/n: 402000351 | $216.43 $155.94 | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | D15CHM641R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/510 | 20 Mar 18 - 20 Mar 18 Failed on chromatography test | Replaced inline filter p/n:WAT035190 | 13.2 | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 1 port 2 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E759A | WAT-RW-SKI01717 | 2018-04-10 | 365 | 2019-04-10 | Shamsul | NA | CR/515 | 20 Mar 18- 20 Mar 18 Lamp Hours expiry hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81014 | Column Heater | WATERS | WAT091727 | F00CHM531R | NA | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/558 | 20 Mar 18 - 20 Mar 18 Failed on chromatography test | Replaced inline filter p/n:WAT035190 | 13.2 | In-Use | TEST |  |  |  | 1
| eAlliance Final Test 1 Port 3 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E765A | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/518 | 20 Mar 18- 20 Mar 18 Lamp Hours expiry hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  |  | 1
|  | 81021 | Column Heater | WATERS | WAT038040 | M00CHM910M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/563 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 1 Port 4 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | H1387E729A | WAT-RW-SKI01717 | 2017-07-22 | 365 | 2018-07-22 | Shamsul | NA | CR/515 | 20 Mar 18- 20 Mar 18 Lamp Hours expiry hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  |  | 1
|  | 81020 | Column Heater | WATERS | WAT038040 | H01CHM352M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/564 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 2 | eAlliance Final Test 2 CPU | M50 | CPU | WATERS | 2489 | KKPRA3T | NA | NA | 365 | NA |  | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Test 2 | 81024 | Temperature Control Module | WATERS | 2489 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | C1387E135M | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/516 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81008 | Column Heater | WATERS |  | M03CHM015R | WAT38172 | 2018-04-04 | 365 | 2019-04-04 | Shamsul | NA | CR/509 | 20 Mar 18 - 20 Mar 18 Failed on chromatography test | Replaced inline filter p/n:WAT035190 | 13.2 | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E710A | WAT-RW-SKI01717 | 2017-07-11 | 365 | 2018-07-11 | Shamsul | NA | CR/545 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81018 | Column Heater | WATERS |  | MX6JM7863M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/563 | 20 Mar 18 - 20 Mar 18 Failed on chromatography test | Replaced inline filter p/n:WAT035190 | 13.2 | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E731A | WAT-RW-SKI01717 | 2017-06-19 | 365 | 2018-06-19 | Shamsul | NA | CR/545 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81009 | Column Heater | WATERS |  | E01CHM203R | WAT38172 | 2018-04-03 | 365 | 2019-04-03 | Shamsul | NA | CR/508 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E727A | WAT-RW-SKI01717 | 2017-07-11 | 365 | 2018-07-11 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81016 | Column Heater |  |  | MX6JM7867M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/569 | 20 Mar 18 - 20 Mar 18 Failed on chromatography test | Replaced inline filter p/n:WAT035190 | 13.2 | In-Use | TEST |  |  |  | 1
eAlliance Final Test Station 3 | eAlliance Final Tester 3 | Lace 32 | CPU | KKTCN1W | CPU | M11L32097W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Tester 3 | 81030 | Temperature Control Module | WATERS | WAT091728 | MX5MM4884M | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/573 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Tester 3 port 1 | 81013 | Column Heater | WATERS | WAT091727 | F00CHM528R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/568 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E735A | WAT-RW-SKI01717 | 2017-07-11 | 365 | 2018-07-11 | Shamsul | NA | CR/544 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Tester 3 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E728A | WAT-RW-SKI01717 | 2017-06-28 | 365 | 2018-06-28 | Shamsul | NA | CR/541 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81012 | Column Heater | WATERS | WAT091727 | B03CHM551R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/570 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Tester 3 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E737A | WAT-RW-SKI01717 | 2017-07-12 | 365 | 2018-07-12 | Shamsul | NA | CR/542 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81010 | Column Heater | WATERS | WAT091727 | L97CHM711M | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/559 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Tester 3 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E734A | WAT-RW-SKI01717 | 2017-07-21 | 365 | 2018-07-21 | Shamsul | NA | CR/536 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81019 | Column Heater | Waters | WAT091727 | D01CHM128R | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/565 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 3 | 81026 | Temperature Control Module | WATERS | 2489 | G01TCM633M | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/574 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
eAlliance Final Test Station 4 | eAlliance Final Test 4 | NA | CPU | Waters | NA | H10L32039W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Test 4 | 81033 | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 4 Port 1 | 81022 | Column Heater | WATERS | WAT091727 | H01CHM335M | WAT038172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | D1387E766A | WAT-RW-SKI01717 | 2017-06-20 | 365 | 2018-06-20 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 4 Port 2 | 81017 | Column Heater | WATERS | WAT038040 | MX6KM9444M | WAT38172 | 2018-04-02 | 365 | 2019-04-02 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E733A | WAT-RW-SKI01717 | 2017-06-20 | 365 | 2018-06-20 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 4 | 81024 | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 4 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E732A | WAT-RW-SKI01717 | 2017-06-20 | 365 | 2018-06-20 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81011 | Column Heater | WATERS | WAT038040 | K03CHM938R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 4 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E736A | WAT-RW-SKI01717 | 2017-06-21 | 365 | 2018-06-21 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81023 | Column Heater | WATERS | WAT038040 | MX5MM5993M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 5 | eAlliance Final Test 5 | NA | CPU | Waters | NA | C12L32205W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80%
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 5 Port 1 | NA | Column Heater | WATERS | WAT091727 | C13CHM795R | WAT038172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | B1587E610A | WAT-RW-SKI01717 | 2018-02-20 | 365 | 2019-02-20 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 Port 2 | NA | Column Heater | WATERS | WAT038040 | D12CHM299R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E607R | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 5 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E608A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | K14CHM398R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 5 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E609A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | L14CHM438R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| Hipot | 95037 | Electrical Safety Analyzer |  |  | 5037655 | NA | 2017-11-09 | 365 | 2018-11-09 | [6 empty] | In-Use | TEST | 7 |  | 4 | 2 | 2
Not-In-Use
QA WORKSTATION | QA WORKSTATION | NA | CPU |  |  | C12L32205W | NA | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST | 270 | 80% | 7 | 1 | 1
| QA WORKSTATION | NA | Column Heater |  |  | H05SMH075E | WAT38172 | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector |  |  | D1387E763A | WAT-RW-SKI01717 | 2018-04-11 | 365 | 2019-04-11 | Shamsul | NA | CR/541 |  |  |  |  | TEST |  |  | 18 | 1 | 1
PUMP CELLS | PUMP Cell | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | CR/10336649-2.1 |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| 1500 SERIES CELL STATION | 90273 | Electronic Weighing Balance | AND | HR200 | 12308385 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 80 | 90% | 6 | 2 | 2
|  | TE 82705-2 | 1500 Series Final Test Station | WATERS | NA | TE 82705-2 | 735000477 | 2017-11-24 | 181 | 2018-05-24 | Shamsul | NA | CR/546 |  |  |  | In-Use | TEST |  |  | 1 | 3 | 3
| 515 PUMP CELL STATION | 90018 | Electronic Weighing Balance | AND | HR200 | 12303279 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 35 | 90% | 6 | 2 | 2
|  | TE 82534 | 515 Final Tester | WATERS | NA | TE 82534 | 070801 | 2017-11-24 | 181 | 2018-05-24 | Shamsul | NA | CR/547 | 19 Mar 18 - 19 Mar 18 Failed on Flow test | Replaced flow valve (FV3) p/n:ASFVO | Consigned | In-Use | TEST |  |  | 1 | 3 | 3
BSM/QSM FINAL TESTER | BSM/QSM Cell | 98206 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | NA |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| BSM/QSM TE82943-11 | 90606 | Electronic Weighing Balance | AND | HR-200 | 90606 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-11 | BSM Final Function Tester | WATERS | TE82943-11 | TE82943-11 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | Waters | CR/549 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84010 | QSM_TUV DETECTOR | WATERS | TE 82943-11 | M09UPT 364A | 735001828 rev B | 2018-03-02 | 365 | 2019-03-02 | Shamsul | NA | CR/506 | 02 Mar 18- 02 Mar 18 Lamp Hours expiry hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-12 | 90605 | Electronic Weighing Balance | AND | HR-200 | 90605 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-12 | BSM Final Function Tester | WATERS | TE82943-12 | TE82943-12 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | na | CR/550 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84007 | QSM_TUV DETECTOR | WATERS | TE 82943-12- | M09UPT 361A | 735001828 rev B | 2017-05-04 | 365 | 2018-05-04 | Shamsul | NA | CR/523 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-13 | 90608 | Electronic Weighing Balance | AND | HR-200 | 90608 | NA | 2018-09-18 | 181 | 2019-03-18 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-13 | BSM Final Function Tester | WATERS | TE82943-13 | TE82943-13 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/551 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84011 | QSM_TUV DETECTOR | WATERS | TE 82943-13- | K09UPT 215A | 735001828 rev B | 2017-05-03 | 365 | 2018-05-03 | Shamsul | NA | CR/525 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-14 | 90604 | Electronic Weighing Balance | AND | HR-200 | 90604 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-14 | BSM Final Function Tester | WATERS | TE82943-14 | TE82943-14 | NA | 2018-01-15 | 365 | 2019-01-15 | Shamsul | NA | CR/552 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84008 | QSM_TUV DETECTOR | WATERS | TE 82943-14- | M09UPT 362A | 735001828 rev B | 2017-04-25 | 365 | 2018-04-25 | Shamsul | NA | CR/522 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-15 | 90603 | Electronic Weighing Balance | AND | HR-200 | 90603 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-15 | BSM Final Function Tester | WATERS | TE82943-15 | TE82943-15 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/553 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84009 | QSM_TUV DETECTOR | WATERS | TE 82943-15- | M09UPT 363A | 735001828 rev B | 2017-04-26 | 365 | 2018-04-26 | Shamsul | NA | CR/521 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-16 | 90600 | Electronic Weighing Balance | AND | HR-200 | 90600 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-16 | BSM Final Function Tester | WATERS | TE82943-16 | TE82943-16 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/554 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84006 | QSM_TUV DETECTOR | WATERS | TE 82943-16- | M09UPT 360A | 735001828 rev B | 2017-05-04 | 365 | 2018-05-04 | Shamsul | NA | CR/524 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-17 | 90607 | Electronic Weighing Balance | AND | HR-200 | 90607 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-17 | BSM Final Function Tester | WATERS | TE82943-17 | TE82943-17 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/555 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84005 | QSM_TUV DETECTOR | WATERS | TE 82943-17 | M09UPT 359A | 735001828 rev B | 2017-04-22 | 365 | 2018-04-22 | Shamsul | NA | CR/520 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-18 | 90610 | Electronic Weighing Balance | AND | HR-200 | 90610 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-18 | BSM Final Function Tester | WATERS | TE82943-18 | TE82943-18 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/556 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | NA | QSM_TUV DETECTOR | WATERS | TE 82943-18 | M09UPT 366A | 735001828 rev B | 2017-04-25 | 365 | 2018-04-25 | Shamsul | NA | CR/519 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| Actuator Tester TE 82800-2 | TE 82800-2 | Aquity Actuator Test | WATERS | TE 82800-2 | TE 82800-2 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 2 | 2 | 2
| Actuator Tester TE 82800-3 | TE 82800-3 | Aquity Actuator Test | WATERS | TE 82800-3 | TE 82800-3 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 2 | 2 | 2
Cleanroom | Clean Room | NA | Picoammeter | Keithley | 485 | 435110 | NA | 2018-03-28 | 365 | 2019-03-28 | NA | Keithley | NA |  |  |  | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | NA | Light Meter |  |  | NA | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | TE 82475-2 | Flowcell Tester | WATER | NA | TE 82475-2 | TE82475-2 | 2018-04-11 | 365 | 2019-04-10 | Shamsul | NA | CR/513 |  |  |  | In-Use | TEST | NA | 90% | 2 | 1 | 1
| Clean Room | TE82881-2 | Optic Bench fine alignment Tester (Box Build) | Waters | 2998 | K15UPD782R | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82881-2 | PDA Fine Alignment CPU |  |  | D08L32121W / C08BLM058C | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82880-2 | PDA ROUGH ALIGNMENT CPU | Waters |  | L07L32070W | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE# 82874 | TUV lamp supply | Waters | 2998 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE # 82965 | Filter Wheel Alignment | Waters | TE 82965 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE9990293134 | Mercury lamp power Supply | Waters |  | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | Test | NA | 90% |  | 1
| Clean Room | TE829002998 | Flowcell Tester CPU | Waters |  | NA | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 TE# 999029981 | 2998 fine alignment(box build) |  |  | E08998580N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 | 2998 fine alignment CPU |  |  | B08L32065W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82917-1 | 2998 rough alignment CPU | Waters |  | B08L32068W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE186002998 | Flowcell tester(2998) | Waters |  | G08998943N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82951-1 | PDA Assembly CPU |  |  | J06L32014W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82868-2 | PDA Aperture Alignment System | Waters |  | K07L32122W PN: 668000351 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | PROCESS | NA | 90% | 1 | 3 | 3
| Clean Room | 95515 | Starlite 150 |  |  | RSL1002609 | NA | NA | NA | NA |  |  |  | 30 Mar 18 - Present Have issue on X axis | Need repair and calibration |  | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE83096-2 | PDA,2998 Coarse and fine Align |  |  | C17L32063W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
Eng's Area | Engineering Lab | Pressure calibrator | Pressure Calibrator | AMETEK | APC10KGINDG | 9093205 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | CR/10311178-1.1 |  |  |  | In-Use | TEST | As per usage on evaluation only | NA | 1 | 1 | 1
| Engineering Lab | FLUKE179 | Multimeter | FLUKE | 179 | 90240092 | NA | 2018-01-31 | 365 | 2019-01-30 | NA | External | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | High Volrage Probe | High Voltage Probe | FLUKE | 80K-6 | 89200011 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Picometer | Picometer |  |  | 310788 | NA | Calibrate only 1 month before the other one due date ( Cost Management) |  |  |  | Keithly | NA |  |  |  | In-Use |  |  | NA
| Engineering Lab | Clamp Meter | Clamp Meter | KYORITSU | 2033 | 66386 | NA | 2017-08-05 | 365 | 2018-08-04 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Fluke 51 II | Digital Thermometer with probe | Fluke | 51 II | 89280141 | NA | 2017-08-10 | 365 | 2018-08-10 | NA | SGS | FMS/323 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | NA | PDA detector |  |  | B08UPD563M | NA | NA | NA | NA | [6 empty] | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | FLUKE110 | Multimeter |  |  | 85730619 | NA | 2018-01-05 | 365 | 2019-01-05 | [11 empty] | 1
| Engineering Lab | 90406 | Liquilaz | Particle Measuring Systems | E20P | 51053 | NA | 2017-08-17 | 365 | 2018-08-16 | NA | Particle Measuring Systems | 110130306 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | 95037 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 5037655 | NA | 2018-01-09 | 365 | 2019-01-09 | [11 empty] | 1
| Engineering Lab | 92207 | Electronic Weighing Balance |  |  | 92207 | NA | 2018-03-08 | 182 | 2018-09-06
| Engineering Lab | 90282 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 1140660 | NA | 2013-06-21 | 365 | 2014-06-21 | NA | SGS | NA | Inactive-Remove from line (RE/094/13) Date 03/12/13. |  |  | In-Use | TEST | 7 | NA | 4 | 2 | 2
Detector | DOWNLOADING AND PROGRAMMING | TE 82964-1 | Download and programme(2489 and TUV) | Waters | NA | TE 82964-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 3 | 2 | 3
In-Use
|  | TE82972-1 | Download and programme(2998, PDA and elamda) | Waters | NA | TE82972-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 |  |  | 2 | 3
In-Use
|  | TE 82969-1 | Calibration and Linearity test for PDA family |  |  | TE 82969-1 | NA | NA | NA | NA | [11 empty] | 1
| 2489/TUV | TE 82903-10 | Final Tester | Waters | NA | TE 82903-10 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-8 | Final Tester | Waters | NA | TE 82903-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2489/TUV | TE 82903-9 | Final Tester | Waters | NA | TE 82903-9 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-6 | Final Tester | Waters | NA | TE 82903-6 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2998, PDA eLambda | TE 82882-8 | Final Tester | Waters | NA | TE 82882-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 16 | 2 | 1
|  | TE 82882-9 | Final Tester | Waters | NA | TE 82882-9 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 1 | 1
|  | TE 82882-10 | Final Tester | Waters | NA | TE 82882-10 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82882-11 | Final Tester | Waters | NA | TE 82882-11 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82943 - 13 | 2998 Final Tester |  |  | TE 82943 - 13 | NA | NA | NA | NA | [11 empty] | 1
|  | TE 82943 - 14 | 2998 Final Tester |  |  | TE 82943 - 14 |  | NA | NA | NA | [11 empty] | 1
| Printer | NA | HP Laserjet P3005n |  |  | SN: JPKGJ01792 | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | NA | NA | NA | NA
|  | NA | Brother |  |  | E69530C4N142942 |  | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | NA | NA | NA | NA
| Hipot | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 08280865 | NA | 2018-10-13 | 365 | 2019-10-13 | NA | SGS | CR/10339085/1.1 |  |  |  |  | TEST | 7 | NA | 4 | 2 | 2
Production Floor | Label Printing Area | WATERS | WATERSISNL1 |  |  | 54177579 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
Production Floor | Label Printing Area | WATERS | WATERSPNL1 |  |  | 54177587 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
| TE82943-14 | NA | Electronic Weighing Balance | AND | HR-200 | 12336741 | NA | 2017-12-08 | 182 | 2018-06-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST
Formulas:
Calibration Due Date = [Calibration Date]+[Calibration Cycle (Months/ Days)]
Calibration Due Date = [Calibration Date]+365
Note on P4: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on R4: skwjdedu: Green - Fixture working Red - Fixture not working
Note on W4: user: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on X4: user: 3 = High Risk, production volume high, will be affected once tester down 2 = Medium Risk, tester down will not affect output for max 2 weeks 1 = Low Risk, tester down will not affect output when down for more than 2 weeks.
Note on V9: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V11: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on G13: SGSJDEDU: EQID 0216
Note on V13: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V19: skwjdedu: 1 down
Note on V26: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V29: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V30: skwjdedu: 12 working, 4 spare
Note on V31: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V34: skwjdedu: 12 working, 4 spare
Note on V36: skwjdedu: 12 working, 4 spare
Note on G47: SGSMSBAf: original CPU KKTCN1W
Note on V49: skwjdedu: 12 working, 4 spare
Note on V50: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V51: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V52: skwjdedu: 12 working, 4 spare
Note on V53: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V54: skwjdedu: 12 working, 4 spare
Note on V55: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V56: skwjdedu: 12 working, 4 spare
Note on V60: skwjdedu: 12 working, 4 spare
Note on V61: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V62: skwjdedu: 12 working, 4 spare
Note on V63: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V64: skwjdedu: 4 working, 3 spare
Note on V65: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V66: skwjdedu: 12 working, 4 spare
Note on V67: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V68: skwjdedu: 12 working, 4 spare
Note on V71: skwjdedu: 12 working, 4 spare
Note on V72: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V73: skwjdedu: 12 working, 4 spare
Note on V74: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V75: skwjdedu: 4 working, 3 spare
Note on V76: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V77: skwjdedu: 12 working, 4 spare
Note on V78: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V79: skwjdedu: 12 working, 4 spare
Note on V80: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V83: skwjdedu: 12 working, 4 spare
Note on V84: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V86: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V87: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V89: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V92: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V95: skwjdedu: BSM no need detectors
Note on V98: skwjdedu: BSM no need detectors
Note on V101: skwjdedu: BSM no need detectors
Note on V104: skwjdedu: BSM no need detectors
Note on V107: skwjdedu: BSM no need detectors
Note on V110: skwjdedu: BSM no need detectors
Note on V113: skwjdedu: BSM no need detectors
Note on V116: skwjdedu: BSM no need detectors
Note on V120: skwjdedu: 1 - clean room, 1 - eng area but not calibrated. Express calibration - 2 weeks, normal - 1 mth.
Note on V122: skwjdedu: 1 - clean room, 1 spare
Note on C141: SGSJDEDU: EQID 0279
Note on C144: SGSJDEDU: EQID 0196
Note on V176: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors

## Column Inventory
P/N 1860002699 ( Column )
Date | Qty In | Qty Out | Balance | Remarks
2016-06-15 | 20 |  |  | Received on 15 Jun 16
2016-06-28 |  | 1 | 19 | Replaced column from S/N:010932104111 to 01093507912602 on FT 4 Port 2 (CHM 81017)
2016-09-23 |  | 1 | 18 | Replaced column from s/n:010932104128 to S/n:01103524612501 on FT 1 Port 3 (CHM 81020)
2016-11-23 |  | 1 | 17 | Replaced column from s/n:010932104115 to s/n: 01093524015005 on FT 2 Port 2 (CHM 81018)
2016-12-15 |  | 1 | 16 | Replaced column from s/n:010933275130 to s/n: 01093524014901 on FT 1 Port 1 (CHM S/n:D15CHM641R)
2016-12-16 |  | 1 | 15 | Replaced column from s/n:010933275126 to s/n: 01093524015004 on FT 5 Port 4 (CHM S/n:L14CHM438R)
2017-04-10 |  | 1 | 14 | Replaced Column from s/n: 01073013010321 to s/n:01093524014910 FT 1 port 3 (CHM 81021)
2017-07-07 |  | 1 | 13 | Replaced from s/n:010135074109 to s/n: 01093524015001 FT 4 port 3 (CHM 81011)
2017-10-19 |  | 1 | 12 | Replaced column from s/n:010933275126 to s/n:01093524014909 FT 2 port 3 (CHM 81009)
2017-10-25 |  | 1 | 11 | Replaced column from s/n:01093507912601 to s/n:01093524015003 FT 2 port 1 (CHM 81008)
2017-11-21 |  | 1 | 10 | Replaced column p/n:186002699 from s/n:010932104120 to s/n:01103524612507 FT 2 port 4 (CHM 81016)
2018-01-10 |  | 1 | 9 | Replaced column p/n:186002699 from s/n:010730188123 to s/n:01093524015006 FT 1 port 2 (CHM 81014)