## FORM
Section 1: Introduction
Requester | Joel Mamalateo Deduyo | Date |  | 2020-02-28
Evaluation# |  | Version #
Category | Others | TMD/ NCR # |  | N/A
Affected P/N(s) | 186269506 | P/N Description(s) |  | e2695 Separations Module W/ H/C
Section 2: Evaluation Assessment
Problem Statement (Describe briefly of the issue. Any impacts to the production capacity, output, resources, or product designs. Include serial numbers and document numbers with revision where applicable. List down potential causes.)
1. Waters request to evaluate new firmware ver 3.05 prior to release new instrument firmware. 2. To install Alliance new firmware ver. 3.05 onto (Donwloading station: Aluto loader 3) and Alliance PN 186269506 and perform Final Test. 3. After completion, Alliance new, Flex to provide test record to Waters. 4. Remove the new firmware ver 3.05 from Aluto loader 3 and reverse back the Alliance PN 186269506 and perform Final Test.
Evaluation Plan (Actions must be listed in a sequential manner.)
Action Owner | Action Description |  |  |  | Expected Completion Date
Mark Jussaume | 1. Mark Jussaume to install Alliance new firmware ver 3.05 to Autoloader 3 Station . |  |  |  | 2020-02-28
Joel Mamalateo Deduyo | 1. Install the Alliance new Firmware ver 3.05 and perform Alliance Final Test as per WAT271129. |  |  |  | 2020-03-02
2. Perform Alliance Final Test as per WAT271129 using Alliance Final Tester 6. If UUT Pass, submit report to Waters. If UUT Fail, submit report to Waters. 3. Delete the Alliance new Firmware ver 3.05 from Autoloder3 Re-intall back the current Alliance firmware ver 3.04 and perform Alliance Final Test as per WAT271129. Release back to Production.
Section 3: Evaluation Approvals
| Name / Title |  | Signature |  | Date
CMO Mfg Lead
Waters Mfg Engineer
Waters Quality Engineer

## Sheet1
TMD
First Article
NCR
Failure Analysis
Others