## Slide 1: QSM – Test Improvement
Apr 2021

## Slide 2: Data Comparison 6 months data in between test software version 1.3 & 1.5
FPY Improvement over 6 months - ~10%
Time saved based on 10% yield improvement over 6 mths and 6 units per day - 0.1hr per unit
Methanol saved – 116mL or USD 1.62 per unit
Software Version | Date | Models | Total Tested Qty | Priming Failure Qty | % of failure
1.3(Before improvement) | 01 Apr'20 - 19 Pct'20 | 186017000 | 298 | 36 | 12.08%
1.5 (After improvement) | 20 Oct'20 - 30Mar'21 | 186017000 & 186017016 | 747 | 15 | 2.01%
Software Version | Avg. Time taken for troubleshooting (labour & machine) per unit (hr) | Total troubleshooting (labour & machine) hr for 747 instruments
1.3(Before improvement) | 0.4 | 298.8
1.5 (After improvement) | 0.3 | 224.1
Time saved for 10% improvement for 747 instruments | 0.1 | 74.7
Software Version | MeOH usage for priming per unit (L) | Cost per unit (USD) | Total MeOH usage for 747 instruments | Total cost for solvent for 747 instruments(USD)
1.3(Before improvement) | 0.5 | $7.00 | 373.5 | $5229.00
1.5 (After improvement) | 0.384 | $5.38 | 286.848 | $4015.87
Saving MeOH for 10% improvement for 747 instruments | 0.116 | $1.62 | 86.652 | $1213.13